## Summary
259 | Total Vulnerabilities
0 | Total Vulnerabilities Remediated
0 | Total Vulnerabilities Submitted to SDCD
0 | Total Devices
Data is based on ICAMP scan and NMDD inventory. Repeat vulnerabilities are highlighted in yellow. CISA Top 10\BOD vulnerabilities are highlighted in red.

## Reference Sheet
Reference Sheet
Remediation Action Responses:
Please note the System Owner remediation actions (column A on the Vulnerability Data tab) have been modified to add the status of how the action will be documented in REEF by SDCD (in parentheses).
System Owner remediation actions on Vulnerability Data tab
Above Baseline-(PVP entry required)
Confirmed but not remediated due to lack of resources (Pending)
False Positive (False Positive)
Non-Medical/SPS Device (Reassigned)
Patch Installation Not Approved-(PVP entry required)
Patched Application or OS (Completed)
Removed 3rd Party Application (Completed)
Replaced/Retired (Completed)
Unsupported Operating System-(PVP entry required)
Vendor Remediation Scheduled (Pending)
For Pending, PVP entry required, and Reassigned options, the System Owner will include the following in the “Additional Notes” (column B on the Vulnerability Data tab) for each vulnerability, as applicable:
Pending: expected date of vendor remediation, or date when resources will be available to remediate.
POAM Vulnerability Portal (PVP) Entry Required: these are devices that cannot be patched for some reason and will require a PVP Entry be created. If PVP entry is not created, the vulnerability will show monthly as a repeat finding.
Reassigned: Use this when the device does not belong to you and you need to use the system assignment process to move the device to the appropriate owner.
Identify remediation actions based on vulnerability reports
The following guidance will assist medical device and SPS system owners in identifying the required actions that need to be taken to remediate vulnerabilities listed within the monthly medical device and SPS vulnerability reports. On the vulnerability tab of the reports, there are three columns described below that are most helpful in identifying remediation solutions.
1. Plugin Column:
The Plugin column (highlighted below) is key to the vulnerability, as each plugin represent a vulnerability identified by the Nessus scan and corresponds with the solution and plugin text. Additional details regarding the vulnerability can be found by placing the plugin number in the search field of this link: https://www.tenable.com/plugins/search
2. Solution Column
The Solution column (highlighted below) will provide the required action to be taken to remediate the vulnerability. The required action will also be listed as a solution in the Plugin Text Column, along with additional details.
3. Plugin Text Column
The Plugin Text column (highlighted below) will provide a synopsis of the vulnerability, along with remediation details. Details in this column include the “Plugin Output”, which identifies the scanned version currently on the device and the “Fixed Version”, which indicates the version that should be installed on the device to remediate the vulnerability. The solution to remediate the vulnerability is also listed.

## Vulnerability Data
BME Action Taken (Please use scroll bar) | Additional_Notes | IP_Address | Plugin | IP_Plugin | Plugin_Name | Solution | Severity | Plugin_Text | ip_description | ip_manufacturer | ip_ee | ip_location | Branch_Region | Branch_Network | Branch | HostName | vn_name | vn_description | vn_mdia | vn_acl_name | DNS_Name | OUPath | OS_via_Nessus | Protocol | Port | Family | Status | Vulnerability_Pub_Date | First_Discovered | Last_Observed | Months_With_Vulnerability | CISA_Top_10_Vulnerability | BOD
|  |  | 97833 |  | MS17-010: Security Update for Microsoft Windows SMB Server (4013389) (ETERNALBLUE) (ETERNALCHAMPION) (ETERNALROMANCE) (ETERNALSYNERGY) (WannaCry) (EternalRocks) (Petya) (uncredentialed check) | Microsoft has released a set of patches for Windows Vista, 2008, 7, 2008 R2, 2012, 8.1, RT 8.1, 2012 R2, 10, and 2016. Microsoft has also released emergency patches for Windows operating systems that are no longer supported, including Windows XP, 2003, and 8. For unsupported Windows operating systems, e.g. Windows XP, Microsoft recommends that users discontinue the use of SMBv1. SMBv1 lacks security features that were included in later SMB versions. SMBv1 can be disabled by following the vendor instructions provided in Microsoft KB2696547. Additionally, US-CERT recommends that users block SMB directly by blocking TCP port 445 on all network boundary devices. For SMB over the NetBIOS API, block TCP ports 137 / 139 and UDP ports 137 / 138 on all network boundary devices. | High | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is affected by the following vulnerabilities : - Multiple remote code execution vulnerabilities exist in Microsoft Server Message Block 1.0 (SMBv1) due to improper handling of certain requests. An unauthenticated, remote attacker can exploit these vulnerabilities, via a specially crafted packet, to execute arbitrary code. (CVE-2017-0143, CVE-2017-0144, CVE-2017-0145, CVE-2017-0146, CVE-2017-0148) - An information disclosure vulnerability exists in Microsoft Server Message Block 1.0 (SMBv1) due to improper handling of certain requests. An unauthenticated, remote attacker can exploit this, via a specially crafted packet, to disclose sensitive information. (CVE-2017-0147) ETERNALBLUE, ETERNALCHAMPION, ETERNALROMANCE, and ETERNALSYNERGY are four of multiple Equation Group vulnerabilities and exploits disclosed on 2017/04/14 by a group known as the Shadow Brokers. WannaCry / WannaCrypt is a ransomware program utilizing the ETERNALBLUE exploit, and EternalRocks is a worm that utilizes seven Equation Group vulnerabilities. Petya is a ransomware program that first utilizes CVE-2017-0199, a vulnerability in Microsoft Office, and then spreads via ETERNALBLUE. Solution: Microsoft has released a set of patches for Windows Vista, 2008, 7, 2008 R2, 2012, 8.1, RT 8.1, 2012 R2, 10, and 2016. Microsoft has also released emergency patches for Windows operating systems that are no longer supported, including Windows XP, 2003, and 8. For unsupported Windows operating systems, e.g. Windows XP, Microsoft recommends that users discontinue the use of SMBv1. SMBv1 lacks security features that were included in later SMB versions. SMBv1 can be disabled by following the vendor instructions provided in Microsoft KB2696547. Additionally, US-CERT recommends that users block SMB directly by blocking TCP port 445 on all network boundary devices. For SMB over the NetBIOS API, block TCP ports 137 / 139 and UDP ports 137 / 138 on all network boundary devices. See Also: http://www.nessus.org/u?68fc8eff http://www.nessus.org/u?321523eb http://www.nessus.org/u?065561d0 http://www.nessus.org/u?d9f569cf https://blogs.technet.microsoft.com/filecab/2016/09/16/stop-using-smb1/ http://www.nessus.org/u?b9d9ebf9 http://www.nessus.org/u?8dcab5e4 http://www.nessus.org/u?234f8ef8 http://www.nessus.org/u?4c7e0cf3 https://github.com/stamparm/EternalRocks/ http://www.nessus.org/u?59db5b5b CVSS Base Score: 8.1 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:C/I:C/A:C Plugin Output: Sent: 00000054ff534d4225000000001803c80000000000000000000000000008def70008000110000000 00ffffffff0000000000000000000000005400000054000200230000001100005c00500049005000 45005c0000000000 Received: ff534d4225050200c09803c80000000000000000000000000008def700080001000000 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2017-03-14 Patch Publication Date: 2017-03-14 Plugin Modification Date: 2022-05-25 Exploitability Ease: Exploits are available Plugin Type: remote Source File: | Sysmex SP-50 #1 (Host) | SYSMEX AMERICA INC |  |  | Region X | VISN X | VAMC | [7 empty] | Microsoft Windows Embedded Standard 7 | tcp | 445 | Windows | Repeat Vulnerability | 2017-03-20 | 2022-02-06 21:56 | 2023-05-09 07:18 | 15 | Yes
|  |  | 166555 |  | WinVerifyTrust Signature Validation CVE-2013-3900 Mitigation (EnableCertPaddingCheck) | Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck | High | Synopsis: The remote Windows host is potentially missing a mitigation for a remote code execution vulnerability. Description: The remote system may be in a vulnerable state to CVE-2013-3900 due to a missing or misconfigured registry keys: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck An unauthenticated, remote attacker could exploit this, by sending specially crafted requests, to execute arbitrary code on an affected host. Solution: Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck See Also: https://msrc.microsoft.com/update-guide/vulnerability/CVE-2013-3900 http://www.nessus.org/u?9780b9d2 CVSS Base Score: 7.4 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: Nessus detected the following potentially insecure registry key configuration: - Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. - Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2013-12-10 Patch Publication Date: NULL Plugin Modification Date: 2023-04-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-10-26 | 2022-11-08 03:32 | 2023-05-09 07:17 | 6 |  | Yes
|  |  | 166555 |  | WinVerifyTrust Signature Validation CVE-2013-3900 Mitigation (EnableCertPaddingCheck) | Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck | High | Synopsis: The remote Windows host is potentially missing a mitigation for a remote code execution vulnerability. Description: The remote system may be in a vulnerable state to CVE-2013-3900 due to a missing or misconfigured registry keys: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck An unauthenticated, remote attacker could exploit this, by sending specially crafted requests, to execute arbitrary code on an affected host. Solution: Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck See Also: https://msrc.microsoft.com/update-guide/vulnerability/CVE-2013-3900 http://www.nessus.org/u?9780b9d2 CVSS Base Score: 7.4 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: Nessus detected the following potentially insecure registry key configuration: - Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. - Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2013-12-10 Patch Publication Date: NULL Plugin Modification Date: 2023-04-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Compass Router - Prod | LAUREL BRIDGE SOFTWARE INC |  |  | Region X | VISN X | VAMC | VHAXXXAPPCROUTP |  |  |  |  | VHAXXXAPPCROUTP.vxx.med.va.gov |  | Microsoft Windows Server 2019 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-10-26 | 2022-11-07 00:55 | 2023-05-08 00:45 | 6 |  | Yes
|  |  | 166555 |  | WinVerifyTrust Signature Validation CVE-2013-3900 Mitigation (EnableCertPaddingCheck) | Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck | High | Synopsis: The remote Windows host is potentially missing a mitigation for a remote code execution vulnerability. Description: The remote system may be in a vulnerable state to CVE-2013-3900 due to a missing or misconfigured registry keys: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck An unauthenticated, remote attacker could exploit this, by sending specially crafted requests, to execute arbitrary code on an affected host. Solution: Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck See Also: https://msrc.microsoft.com/update-guide/vulnerability/CVE-2013-3900 http://www.nessus.org/u?9780b9d2 CVSS Base Score: 7.4 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: Nessus detected the following potentially insecure registry key configuration: - Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. - Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2013-12-10 Patch Publication Date: NULL Plugin Modification Date: 2023-04-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Compass Router - Hot Spare | LAUREL BRIDGE SOFTWARE INC |  |  | Region X | VISN X | VAMC | VHAXXXAPPCROUTS |  |  |  |  | VHAXXXAPPCROUTS.vxx.med.va.gov |  | Microsoft Windows Server 2019 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-10-26 | 2022-11-07 00:55 | 2023-05-08 00:46 | 6 |  | Yes
|  |  | 166555 |  | WinVerifyTrust Signature Validation CVE-2013-3900 Mitigation (EnableCertPaddingCheck) | Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck | High | Synopsis: The remote Windows host is potentially missing a mitigation for a remote code execution vulnerability. Description: The remote system may be in a vulnerable state to CVE-2013-3900 due to a missing or misconfigured registry keys: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck An unauthenticated, remote attacker could exploit this, by sending specially crafted requests, to execute arbitrary code on an affected host. Solution: Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck See Also: https://msrc.microsoft.com/update-guide/vulnerability/CVE-2013-3900 http://www.nessus.org/u?9780b9d2 CVSS Base Score: 7.4 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: Nessus detected the following potentially insecure registry key configuration: - Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. - Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2013-12-10 Patch Publication Date: NULL Plugin Modification Date: 2023-04-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Compass Router - Test | LAUREL BRIDGE SOFTWARE INC |  |  | Region X | VISN X | VAMC | VHAXXXAPPCROUTT |  |  |  |  | VHAXXXAPPCROUTT.vxx.med.va.gov |  | Microsoft Windows Server 2019 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-10-26 | 2022-11-07 00:55 | 2023-05-08 00:44 | 6 |  | Yes
|  |  | 166555 |  | WinVerifyTrust Signature Validation CVE-2013-3900 Mitigation (EnableCertPaddingCheck) | Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck | High | Synopsis: The remote Windows host is potentially missing a mitigation for a remote code execution vulnerability. Description: The remote system may be in a vulnerable state to CVE-2013-3900 due to a missing or misconfigured registry keys: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck An unauthenticated, remote attacker could exploit this, by sending specially crafted requests, to execute arbitrary code on an affected host. Solution: Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck See Also: https://msrc.microsoft.com/update-guide/vulnerability/CVE-2013-3900 http://www.nessus.org/u?9780b9d2 CVSS Base Score: 7.4 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: Nessus detected the following potentially insecure registry key configuration: - Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. - Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2013-12-10 Patch Publication Date: NULL Plugin Modification Date: 2023-04-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-10-26 | 2022-11-08 03:32 | 2023-05-09 07:10 | 6 |  | Yes
|  |  | 166555 |  | WinVerifyTrust Signature Validation CVE-2013-3900 Mitigation (EnableCertPaddingCheck) | Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck | High | Synopsis: The remote Windows host is potentially missing a mitigation for a remote code execution vulnerability. Description: The remote system may be in a vulnerable state to CVE-2013-3900 due to a missing or misconfigured registry keys: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck An unauthenticated, remote attacker could exploit this, by sending specially crafted requests, to execute arbitrary code on an affected host. Solution: Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck See Also: https://msrc.microsoft.com/update-guide/vulnerability/CVE-2013-3900 http://www.nessus.org/u?9780b9d2 CVSS Base Score: 7.4 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: Nessus detected the following potentially insecure registry key configuration: - Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. - Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2013-12-10 Patch Publication Date: NULL Plugin Modification Date: 2023-04-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Report Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE2 |  |  |  |  | VHAXXXAPPROCHE2.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-10-26 | 2022-11-08 03:32 | 2023-05-09 07:17 | 6 |  | Yes
|  |  | 166555 |  | WinVerifyTrust Signature Validation CVE-2013-3900 Mitigation (EnableCertPaddingCheck) | Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck | High | Synopsis: The remote Windows host is potentially missing a mitigation for a remote code execution vulnerability. Description: The remote system may be in a vulnerable state to CVE-2013-3900 due to a missing or misconfigured registry keys: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck An unauthenticated, remote attacker could exploit this, by sending specially crafted requests, to execute arbitrary code on an affected host. Solution: Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck See Also: https://msrc.microsoft.com/update-guide/vulnerability/CVE-2013-3900 http://www.nessus.org/u?9780b9d2 CVSS Base Score: 7.4 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: Nessus detected the following potentially insecure registry key configuration: - Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. - Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2013-12-10 Patch Publication Date: NULL Plugin Modification Date: 2023-04-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | COBAS Infinity Production Server (v3.03) | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXCCACOBASP |  |  |  |  | VHAXXXCCACOBASP.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-10-26 | 2022-11-08 03:32 | 2023-05-09 07:18 | 6 |  | Yes
|  |  | 166555 |  | WinVerifyTrust Signature Validation CVE-2013-3900 Mitigation (EnableCertPaddingCheck) | Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck | High | Synopsis: The remote Windows host is potentially missing a mitigation for a remote code execution vulnerability. Description: The remote system may be in a vulnerable state to CVE-2013-3900 due to a missing or misconfigured registry keys: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck An unauthenticated, remote attacker could exploit this, by sending specially crafted requests, to execute arbitrary code on an affected host. Solution: Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck See Also: https://msrc.microsoft.com/update-guide/vulnerability/CVE-2013-3900 http://www.nessus.org/u?9780b9d2 CVSS Base Score: 7.4 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: Nessus detected the following potentially insecure registry key configuration: - Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. - Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2013-12-10 Patch Publication Date: NULL Plugin Modification Date: 2023-04-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | VIRTUAL SERVER |  |  |  | Region X | VISN X | VAMC | VHAXXXCCAMIM |  |  |  |  | VHAXXXCCAMIM.vxx.med.va.gov |  | Microsoft Windows Server 2019 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-10-26 | 2022-11-06 13:08 | 2023-05-07 06:50 | 6 |  | Yes
|  |  | 166555 |  | WinVerifyTrust Signature Validation CVE-2013-3900 Mitigation (EnableCertPaddingCheck) | Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck | High | Synopsis: The remote Windows host is potentially missing a mitigation for a remote code execution vulnerability. Description: The remote system may be in a vulnerable state to CVE-2013-3900 due to a missing or misconfigured registry keys: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck An unauthenticated, remote attacker could exploit this, by sending specially crafted requests, to execute arbitrary code on an affected host. Solution: Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck See Also: https://msrc.microsoft.com/update-guide/vulnerability/CVE-2013-3900 http://www.nessus.org/u?9780b9d2 CVSS Base Score: 7.4 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: Nessus detected the following potentially insecure registry key configuration: - Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. - Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2013-12-10 Patch Publication Date: NULL Plugin Modification Date: 2023-04-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-10-26 | 2022-11-08 03:12 | 2023-05-09 06:58 | 6 |  | Yes
|  |  | 166555 |  | WinVerifyTrust Signature Validation CVE-2013-3900 Mitigation (EnableCertPaddingCheck) | Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck | High | Synopsis: The remote Windows host is potentially missing a mitigation for a remote code execution vulnerability. Description: The remote system may be in a vulnerable state to CVE-2013-3900 due to a missing or misconfigured registry keys: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck An unauthenticated, remote attacker could exploit this, by sending specially crafted requests, to execute arbitrary code on an affected host. Solution: Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck See Also: https://msrc.microsoft.com/update-guide/vulnerability/CVE-2013-3900 http://www.nessus.org/u?9780b9d2 CVSS Base Score: 7.4 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: Nessus detected the following potentially insecure registry key configuration: - Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. - Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2013-12-10 Patch Publication Date: NULL Plugin Modification Date: 2023-04-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | RALS I-Stat Data Exchange (DE) Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSD |  |  |  |  | VHAXXXCCARALSD.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-10-26 | 2022-11-08 03:06 | 2023-05-09 06:44 | 6 |  | Yes
|  |  | 166555 |  | WinVerifyTrust Signature Validation CVE-2013-3900 Mitigation (EnableCertPaddingCheck) | Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck | High | Synopsis: The remote Windows host is potentially missing a mitigation for a remote code execution vulnerability. Description: The remote system may be in a vulnerable state to CVE-2013-3900 due to a missing or misconfigured registry keys: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck An unauthenticated, remote attacker could exploit this, by sending specially crafted requests, to execute arbitrary code on an affected host. Solution: Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck See Also: https://msrc.microsoft.com/update-guide/vulnerability/CVE-2013-3900 http://www.nessus.org/u?9780b9d2 CVSS Base Score: 7.4 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: Nessus detected the following potentially insecure registry key configuration: - Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. - Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2013-12-10 Patch Publication Date: NULL Plugin Modification Date: 2023-04-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-10-26 | 2022-11-08 03:06 | 2023-05-09 06:45 | 6 |  | Yes
|  |  | 166555 |  | WinVerifyTrust Signature Validation CVE-2013-3900 Mitigation (EnableCertPaddingCheck) | Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck | High | Synopsis: The remote Windows host is potentially missing a mitigation for a remote code execution vulnerability. Description: The remote system may be in a vulnerable state to CVE-2013-3900 due to a missing or misconfigured registry keys: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck An unauthenticated, remote attacker could exploit this, by sending specially crafted requests, to execute arbitrary code on an affected host. Solution: Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck Additionally, on 64 Bit OS systems, Add and enable registry value EnableCertPaddingCheck: - HKEY_LOCAL_MACHINE\Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck See Also: https://msrc.microsoft.com/update-guide/vulnerability/CVE-2013-3900 http://www.nessus.org/u?9780b9d2 CVSS Base Score: 7.4 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: Nessus detected the following potentially insecure registry key configuration: - Software\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. - Software\Wow6432Node\Microsoft\Cryptography\Wintrust\Config\EnableCertPaddingCheck is not present in the registry. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2013-12-10 Patch Publication Date: NULL Plugin Modification Date: 2023-04-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | RALS I-Stat Data Exchange (DE) Server - Production | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSDE |  |  |  |  | VHAXXXCCARALSDE.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-10-26 | 2022-11-08 03:06 | 2023-05-09 06:43 | 6 |  | Yes
| Chosen |  | 108797 |  | Unsupported Windows OS (remote) | Upgrade to a supported service pack or operating system | Critical | Synopsis: The remote OS or service pack is no longer supported. Description: The remote version of Microsoft Windows is either missing a service pack or is no longer supported. As a result, it is likely to contain security vulnerabilities. Solution: Upgrade to a supported service pack or operating system See Also: https://support.microsoft.com/en-us/lifecycle CVSS Base Score: 10 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The following Windows version is installed and not supported: Microsoft Windows Embedded Standard 7601 Service Pack 1 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2022-07-05 Exploitability Ease: NULL Plugin Type: remote Source File: | Sysmex SP-50 #1 (Host) | SYSMEX AMERICA INC |  |  | Region X | VISN X | VAMC | [7 empty] | Microsoft Windows Embedded Standard 7 | tcp |  | Windows | Repeat Vulnerability | 2018-04-03 | 2022-07-10 09:00 | 2023-05-09 07:18 | 10
|  |  | 111068 |  | Apache Tomcat 8.5.0 < 8.5.32 Multiple Vulnerabilities | Upgrade to Apache Tomcat version 8.5.32 or later. | Critical | Synopsis: The remote Apache Tomcat server is affected multiple security vulnerabilities. Description: The version of Apache Tomcat installed on the remote host is 8.5.x prior to 8.5.32. It is, therefore, affected by multiple vulnerabilities. Solution: Upgrade to Apache Tomcat version 8.5.32 or later. See Also: http://www.nessus.org/u?5070a438 http://www.nessus.org/u?d5ab19d6 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.32 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-02-12 Patch Publication Date: 2018-02-13 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 80 | Web Servers | Repeat Vulnerability | 2018-07-13 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 111068 |  | Apache Tomcat 8.5.0 < 8.5.32 Multiple Vulnerabilities | Upgrade to Apache Tomcat version 8.5.32 or later. | Critical | Synopsis: The remote Apache Tomcat server is affected multiple security vulnerabilities. Description: The version of Apache Tomcat installed on the remote host is 8.5.x prior to 8.5.32. It is, therefore, affected by multiple vulnerabilities. Solution: Upgrade to Apache Tomcat version 8.5.32 or later. See Also: http://www.nessus.org/u?5070a438 http://www.nessus.org/u?d5ab19d6 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.32 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-02-12 Patch Publication Date: 2018-02-13 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 8080 | Web Servers | Repeat Vulnerability | 2018-07-13 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 12107 |  | McAfee Antivirus Detection and Status | Make sure that updates are working and the associated services are running. | Critical | Synopsis: An antivirus application is installed on the remote host, but it is not working properly. Description: McAfee VirusScan, an antivirus application, is installed on the remote host. However, there is a problem with the installation; either its services are not running or its virus definitions are out of date. Solution: Make sure that updates are working and the associated services are running. See Also: NULL CVSS Base Score: 10 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host has the McAfee antivirus installed. It has been fingerprinted as : McAfee VirusScan Enterprise : 8.8.0.2024 Engine version : 6000.8403 DAT version : 1111.0000 Updated date : 02022010 ePO Agent : not present. Path : C:\Program Files (x86)\McAfee\VirusScan Enterprise\ The remote host has an out-dated version of the McAfee virus database. Latest version is 10696 CPE: cpe:/a:mcafee:virusscan_enterprise CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2022-10-10 Exploitability Ease: NULL Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2004-03-16 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 72704 |  | Microsoft .NET Framework Unsupported | Upgrade to a version of the Microsoft .NET Framework that is currently supported. | Critical | Synopsis: An unsupported software framework is installed on the remote Windows host. Description: According to its self-reported version number, there is at least one version of Microsoft .NET Framework installed on the remote Windows host that is no longer supported. Lack of support implies that no new security patches for the product will be released by the vendor. As a result, it is likely to contain security vulnerabilities. Solution: Upgrade to a version of the Microsoft .NET Framework that is currently supported. See Also: http://www.nessus.org/u?930a7a91 http://www.nessus.org/u?3b10ac8d CVSS Base Score: 10 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The following Microsoft .NET Framework version is no longer supported : Installed version : Microsoft .NET Framework v4.6.1 EOL date : April 26, 2022 EOL URL : https://docs.microsoft.com/en-us/lifecycle/products/microsoft-net-framework Supported versions : 3.5 / 4.6.2 / 4.7 / 4.7.1 / 4.7.2/ 4.8 CPE: cpe:/a:microsoft:.net_framework CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2022-04-26 Exploitability Ease: NULL Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2014-02-26 | 2022-05-02 07:06 | 2023-05-09 07:17 | 12
|  |  | 91500 |  | McAfee VirusScan Enterprise Unsupported Version Detection | Upgrade to a version of McAfee VirusScan Enterprise (VSE) that is currently supported. | Critical | Synopsis: The remote host is running an unsupported version of McAfee VSE. Description: According to its version, the installation of McAfee VirusScan Enterprise (VSE) on the remote host is no longer supported. Lack of support implies that no new security patches for the product will be released by the vendor. As a result, it is likely to contain security vulnerabilities. Solution: Upgrade to a version of McAfee VirusScan Enterprise (VSE) that is currently supported. See Also: http://www.nessus.org/u?0c5d29f3 CVSS Base Score: 10 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Product : McAfee VirusScan Enterprise Installed version : 8.8.0.2024 End of life date : December 31 2021 EOL announcement : https://kc.mcafee.com/corporate/index?page=content&id=KB93335 CPE: cpe:/a:mcafee:virusscan_enterprise CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2022-06-08 Exploitability Ease: NULL Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2016-06-07 | 2022-07-10 08:46 | 2023-05-09 07:17 | 10
|  |  | 94138 |  | Oracle Java SE Multiple Vulnerabilities (October 2016 CPU) | Upgrade to Oracle JDK / JRE 8 Update 111 / 7 Update 121 / 6 Update 131 or later. If necessary, remove any affected versions. Note that an Extended Support contract with Oracle is needed to obtain JDK / JRE 6 Update 95 or later. | Critical | Synopsis: The remote Windows host contains a programming platform that is affected by multiple vulnerabilities. Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 8 Update 111, 7 Update 121, or 6 Update 131. It is, therefore, affected by multiple vulnerabilities : - An unspecified flaw exists in the Libraries subcomponent that allows an unauthenticated, remote attacker to impact integrity. (CVE-2016-5542) - An unspecified flaw exists in the JMX subcomponent that allows an unauthenticated, remote attacker to impact integrity. (CVE-2016-5554) - An unspecified flaw exists in the 2D subcomponent that allows an unauthenticated, remote attacker to execute arbitrary code. (CVE-2016-5556) - An unspecified flaw exists in the AWT subcomponent that allows an unauthenticated, remote attacker to execute arbitrary code. (CVE-2016-5568) - Multiple unspecified flaws exist in the Hotspot subcomponent that allow an unauthenticated, remote attacker to execute arbitrary code. (CVE-2016-5573, CVE-2016-5582) - An unspecified flaw exists in the Networking subcomponent that allows an unauthenticated, remote attacker to disclose sensitive information. (CVE-2016-5597) Solution: Upgrade to Oracle JDK / JRE 8 Update 111 / 7 Update 121 / 6 Update 131 or later. If necessary, remove any affected versions. Note that an Extended Support contract with Oracle is needed to obtain JDK / JRE 6 Update 95 or later. See Also: http://www.nessus.org/u?bac902d5 http://www.nessus.org/u?10d5f7a6 http://www.nessus.org/u?2fbcacca http://www.nessus.org/u?726f7054 CVSS Base Score: 9.6 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:C/I:C/A:C Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.6.0_131 / 1.7.0_121 / 1.8.0_111 CPE: cpe:/a:oracle:jre cpe:/a:oracle:jdk CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2016-10-18 Patch Publication Date: 2016-10-18 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2016-10-19 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 96628 |  | Oracle Java SE Multiple Vulnerabilities (January 2017 CPU) (SWEET32) | Upgrade to Oracle JDK / JRE 8 Update 121 / 7 Update 131 / 6 Update 141 or later. If necessary, remove any affected versions. Note that an Extended Support contract with Oracle is needed to obtain JDK / JRE 6 Update 95 or later. | Critical | Synopsis: The remote Windows host contains a programming platform that is affected by multiple vulnerabilities. Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 8 Update 121, 7 Update 131, or 6 Update 141. It is, therefore, affected by multiple vulnerabilities : - A vulnerability exists in the Libraries subcomponent, known as SWEET32, in the 3DES and Blowfish algorithms due to the use of weak 64-bit block ciphers by default. A man-in-the-middle attacker who has sufficient resources can exploit this vulnerability, via a 'birthday' attack, to detect a collision that leaks the XOR between the fixed secret and a known plaintext, allowing the disclosure of the secret text, such as secure HTTPS cookies, and possibly resulting in the hijacking of an authenticated session. (CVE-2016-2183) - An unspecified flaw exists in the Libraries subcomponent that allows an unauthenticated, remote attacker to impact integrity. (CVE-2016-5546) - An unspecified flaw exists in the Libraries subcomponent that allows an unauthenticated, remote attacker to cause a denial of service condition. (CVE-2016-5547) - Multiple unspecified flaws exist in the Libraries subcomponent that allow an unauthenticated, remote attacker to disclose sensitive information. (CVE-2016-5548, CVE-2016-5549) - An unspecified flaw exists in the Networking subcomponent that allows an unauthenticated, remote attacker to impact integrity. (CVE-2016-5552) - An unspecified flaw exists in the Mission Control subcomponent that allows an unauthenticated, remote attacker to impact integrity. (CVE-2016-8328) - Multiple unspecified flaws exist in the Networking subcomponent that allow an unauthenticated, remote attacker to disclose sensitive information. (CVE-2017-3231, CVE-2017-3261) - An unspecified flaw exists in the RMI subcomponent that allows an unauthenticated, remote attacker to execute arbitrary code. (CVE-2017-3241) - An unspecified flaw exists in the JAAS subcomponent that allows an unauthenticated, remote attacker to impact integrity. (CVE-2017-3252) - An unspecified flaw exists in the 2D subcomponent that allows an unauthenticated, remote attacker to cause a denial of service condition. (CVE-2017-3253) - An unspecified flaw exists in the Deployment subcomponent that allows an unauthenticated, remote attacker to disclose sensitive information. (CVE-2017-3259) - An unspecified flaw exists in the AWT subcomponent that allows an unauthenticated, remote attacker to execute arbitrary code. (CVE-2017-3260) - An unspecified flaw exists in the Java Mission Control subcomponent that allows an unauthenticated, remote attacker to disclose sensitive information. (CVE-2017-3262) - An unspecified flaw exists in the Libraries subcomponent that allows an unauthenticated, remote attacker to execute arbitrary code. (CVE-2017-3272) - An unspecified flaw exists in the Hotspot subcomponent that allows an unauthenticated, remote attacker to execute arbitrary code. (CVE-2017-3289) Note that CVE-2017-3241 can only be exploited by supplying data to APIs in the specified component without using untrusted Java Web Start applications or untrusted Java applets, such as through a web service. Note that CVE-2016-2183, CVE-2016-5546, CVE-2016-5547, CVE-2016-5552, CVE-2017-3252, and CVE-2017-3253 can be exploited through sandboxed Java Web Start applications and sandboxed Java applets. They can also be exploited by supplying data to APIs in the specified component without using sandboxed Java Web Start applications or sandboxed Java applets, such as through a web service. Solution: Upgrade to Oracle JDK / JRE 8 Update 121 / 7 Update 131 / 6 Update 141 or later. If necessary, remove any affected versions. Note that an Extended Suppor Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.6.0_141 / 1.7.0_131 / 1.8.0_121 | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2017-01-19 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 100995 |  | Apache 2.2.x < 2.2.33-dev / 2.4.x < 2.4.26 Multiple Vulnerabilities | Upgrade to Apache version 2.2.33-dev / 2.4.26 or later. | Critical | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: According to its banner, the version of Apache running on the remote host is 2.2.x prior to 2.2.33-dev or 2.4.x prior to 2.4.26. It is, therefore, affected by the following vulnerabilities : - An authentication bypass vulnerability exists due to third-party modules using the ap_get_basic_auth_pw() function outside of the authentication phase. An unauthenticated, remote attacker can exploit this to bypass authentication requirements. (CVE-2017-3167) - A NULL pointer dereference flaw exists due to third-party module calls to the mod_ssl ap_hook_process_connection() function during an HTTP request to an HTTPS port. An unauthenticated, remote attacker can exploit this to cause a denial of service condition. (CVE-2017-3169) - A NULL pointer dereference flaw exists in mod_http2 that is triggered when handling a specially crafted HTTP/2 request. An unauthenticated, remote attacker can exploit this to cause a denial of service condition. Note that this vulnerability does not affect 2.2.x. (CVE-2017-7659) - An out-of-bounds read error exists in the ap_find_token() function due to improper handling of header sequences. An unauthenticated, remote attacker can exploit this, via a specially crafted header sequence, to cause a denial of service condition. (CVE-2017-7668) - An out-of-bounds read error exists in mod_mime due to improper handling of Content-Type response headers. An unauthenticated, remote attacker can exploit this, via a specially crafted Content-Type response header, to cause a denial of service condition or the disclosure of sensitive information. (CVE-2017-7679) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.2.33-dev / 2.4.26 or later. See Also: https://archive.apache.org/dist/httpd/CHANGES_2.2.32 https://archive.apache.org/dist/httpd/CHANGES_2.4.26 https://httpd.apache.org/security/vulnerabilities_22.html https://httpd.apache.org/security/vulnerabilities_24.html CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.26 CPE: cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2017-06-19 Patch Publication Date: 2017-06-19 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2017-06-22 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 101788 |  | Apache 2.4.x < 2.4.27 Multiple Vulnerabilities | Upgrade to Apache version 2.4.27 or later. | Critical | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: According to its banner, the version of Apache running on the remote host is 2.4.x prior to 2.4.27. It is, therefore, affected by the following vulnerabilities : - A denial of service vulnerability exists in httpd due to a failure to initialize or reset the value placeholder in [Proxy-]Authorization headers of type 'Digest' before or between successive key=value assignments by mod_auth_digest. An unauthenticated, remote attacker can exploit this, by providing an initial key with no '=' assignment, to disclose sensitive information or cause a denial of service condition. (CVE-2017-9788) - A read-after-free error exists in httpd that is triggered when closing a large number of connections. An unauthenticated, remote attacker can exploit this to have an unspecified impact. (CVE-2017-9789) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.27 or later. See Also: https://archive.apache.org/dist/httpd/CHANGES_2.4.27 https://httpd.apache.org/security/vulnerabilities_24.html CVSS Base Score: 9.1 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:N/A:P Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.27 CPE: cpe:/a:apache:http_server cpe:/a:apache:httpd CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2017-07-11 Patch Publication Date: 2017-07-11 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2017-07-18 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 101843 |  | Oracle Java SE Multiple Vulnerabilities (July 2017 CPU) | Upgrade to Oracle JDK / JRE 8 Update 141 / 7 Update 151 / 6 Update 161 or later. If necessary, remove any affected versions. Note that an Extended Support contract with Oracle is needed to obtain JDK / JRE 6 Update 95 or later. | Critical | Synopsis: The remote Windows host contains a programming platform that is affected by multiple vulnerabilities. Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 8 Update 141, 7 Update 151, or 6 Update 161. It is, therefore, affected by multiple vulnerabilities : - An unspecified flaw exists in the 2D component that allows an unauthenticated, remote attacker to cause a denial of service condition. (CVE-2017-10053) - Multiple unspecified flaws exist in the Security component that allow an unauthenticated, remote attacker to execute arbitrary code. (CVE-2017-10067, CVE-2017-10116) - An unspecified flaw exists in the Hotspot component that allows an unauthenticated, remote attacker to execute arbitrary code. (CVE-2017-10074) - An unspecified flaw exists in the Scripting component that allows an authenticated, remote attacker to impact confidentiality and integrity. (CVE-2017-10078) - An unspecified flaw exists in the Hotspot component that allows an unauthenticated, remote attacker to impact integrity. (CVE-2017-10081) - Multiple unspecified flaws exist in the JavaFX component that allow an unauthenticated, remote attacker to execute arbitrary code. (CVE-2017-10086, CVE-2017-10114) - Multiple unspecified flaws exist in the Libraries component that allow an unauthenticated, remote attacker to execute arbitrary code. (CVE-2017-10087, CVE-2017-10090, CVE-2017-10111) - An unspecified flaw exists in the ImageIO component that allows an unauthenticated, remote attacker to execute arbitrary code. (CVE-2017-10089) - Multiple unspecified flaws exist in the JAXP component that allow an unauthenticated, remote attacker to execute arbitrary code. (CVE-2017-10096, CVE-2017-10101) - Multiple unspecified flaws exist in the RMI component that allow an unauthenticated, remote attacker to execute arbitrary code. (CVE-2017-10102, CVE-2017-10107) - Multiple unspecified flaws exist in the Server component of the Java Advanced Management Console that allow an authenticated, remote attacker to impact confidentiality, integrity, and availability. (CVE-2017-10104, CVE-2017-10145) - An unspecified flaw exists in the Deployment component that allows an unauthenticated, remote attacker to impact integrity. (CVE-2017-10105) - Multiple unspecified flaws exist in the Serialization component that allow an unauthenticated, remote attacker to exhaust available memory, resulting in a denial of service condition. (CVE-2017-10108, CVE-2017-10109) - An unspecified flaw exists in the AWT component that allows an unauthenticated, remote attacker to execute arbitrary code. (CVE-2017-10110) - Multiple unspecified flaws exist in the JCE component that allow an unauthenticated, remote attacker to disclose sensitive information. (CVE-2017-10115, CVE-2017-10118, CVE-2017-10135) - An unspecified flaw exists in the Server component of the Java Advanced Management Console that allows an unauthenticated, remote attacker to disclose sensitive information. (CVE-2017-10117) - An unspecified flaw exists in the Server component of the Java Advanced Management Console that allows an unauthenticated, remote attacker to impact confidentiality and integrity. (CVE-2017-10121) - An unspecified flaw exists in the Deployment component that allows a local attacker to impact confidentiality, integrity, and availability. (CVE-2017-10125) - Multiple unspecified flaws exist in the Security component that allow an unauthenticated, remote attacker to disclose sensitive information. (CVE-2017-10176, CVE-2017-10193, CVE-2017-10198) - An unspecified flaw exists in the JAX-WS component that allows an unauthenticated, remote attacker to impact confidentiality and availability. (CVE- Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.6.0_161 / 1.7.0_151 / 1.8.0_141 | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2017-07-20 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 103963 |  | Oracle Java SE Multiple Vulnerabilities (October 2017 CPU) | Upgrade to Oracle JDK / JRE 9 Update 1, 8 Update 151 / 7 Update 161 / 6 Update 171 or later. If necessary, remove any affected versions. Note that an Extended Support contract with Oracle is needed to obtain JDK / JRE 6 Update 95 or later. | Critical | Synopsis: The remote Windows host contains a programming platform that is affected by multiple vulnerabilities. Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 9 Update 1, 8 Update 151, 7 Update 161, or 6 Update 171. It is, therefore, affected by multiple vulnerabilities related to the following components : - 2D (Little CMS 2) - Deployment - Hotspot - JAX-WS - JAXP - Javadoc - Libraries - Networking - RMI - Security - Serialization - Smart Card IO - Util (zlib) Solution: Upgrade to Oracle JDK / JRE 9 Update 1, 8 Update 151 / 7 Update 161 / 6 Update 171 or later. If necessary, remove any affected versions. Note that an Extended Support contract with Oracle is needed to obtain JDK / JRE 6 Update 95 or later. See Also: http://www.nessus.org/u?ffb85cfa http://www.nessus.org/u?dfeae1af http://www.nessus.org/u?bbe7f5cf http://www.nessus.org/u?2fbcacca http://www.nessus.org/u?726f7054 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.6.0_171 / 1.7.0_161 / 1.8.0_151 / 1.9.0_1 CPE: cpe:/a:oracle:jre cpe:/a:oracle:jdk CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2017-10-17 Patch Publication Date: 2017-10-17 Plugin Modification Date: 2022-04-11 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2017-10-19 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 111163 |  | Oracle Java SE Multiple Vulnerabilities (July 2018 CPU) | Upgrade to Oracle JDK / JRE 10 Update 2, 8 Update 181 / 7 Update 191 / 6 Update 201 or later. If necessary, remove any affected versions. Note that an Extended Support contract with Oracle is needed to obtain JDK / JRE 6 Update 95 or later. | Critical | Synopsis: The remote Windows host contains a programming platform that is affected by multiple vulnerabilities. Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 10 Update 2, 8 Update 181, 7 Update 191, or 6 Update 201. It is, therefore, affected by multiple vulnerabilities related to the following components : - Concurrency. A difficult to exploit vulnerability allows an unauthenticated attacker with network access via multiple protocols to compromise Java SE (CVE-2018-2952) - Deployment. A difficult to exploit vulnerability allows an unauthenticated attacker with network access via multiple protocols to compromise Java SE (CVE-2018-2964) - JSSE. A difficult to exploit vulnerability allows an unauthenticated attacker with network access via multiple protocols to compromise Java SE (CVE-2018-2973) - Java DB. A difficult to exploit vulnerability allows an unauthenticated attacker with network access via multiple protocols to compromise Java SE. (CVE-2018-2938) - JavaFX. A difficult to exploit vulnerability allows an unauthenticated attacker with network access via multiple protocols to compromise Java SE. (CVE-2018-2941) - Libraries. An easily exploitable vulnerability allows an unauthenticated attacker with network access via multiple protocols to compromise Java SE. (CVE-2018-2940) - Security. A difficult to exploit vulnerability allows an unauthenticated attacker with network access via multiple protocols to compromise Java SE (CVE-2018-2972) - Windows DLL. A difficult to exploit vulnerability allows an unauthenticated attacker with network access via multiple protocols to compromise Java SE (CVE-2018-2942) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Oracle JDK / JRE 10 Update 2, 8 Update 181 / 7 Update 191 / 6 Update 201 or later. If necessary, remove any affected versions. Note that an Extended Support contract with Oracle is needed to obtain JDK / JRE 6 Update 95 or later. See Also: http://www.nessus.org/u?dbb3b1db http://www.nessus.org/u?8a11ccea http://www.nessus.org/u?6c975c0b http://www.nessus.org/u?2fbcacca http://www.nessus.org/u?726f7054 CVSS Base Score: 9 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:P/A:P Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.6.0_201 / 1.7.0_191 / 1.8.0_181 / 1.10.0_2 CPE: cpe:/a:oracle:jre cpe:/a:oracle:jdk CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-07-17 Patch Publication Date: 2018-07-17 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2018-07-20 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 118228 |  | Oracle Java SE Multiple Vulnerabilities (October 2018 CPU) | Upgrade to Oracle JDK / JRE 11 Update 1, 8 Update 191 / 7 Update 201 / 6 Update 211 or later. If necessary, remove any affected versions. Note that an Extended Support contract with Oracle is needed to obtain JDK / JRE 6 Update 95 or later. | Critical | Synopsis: The remote Windows host contains a programming platform that is affected by multiple vulnerabilities. Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 11 Update 1, 8 Update 191, 7 Update 201, or 6 Update 211. It is, therefore, affected by multiple vulnerabilities related to the following components : - An unspecified vulnerability in the Java SE, Java SE Embedded component of Oracle Java SE in the Deployment (libpng) subcomponent could allow an unauthenticated, remote attacker with network access via HTTP to compromise Java SE, Java SE Embedded. (CVE-2018-13785) - An unspecified vulnerability in the Java SE, Java SE Embedded component of Oracle Java SE in the Hotspot subcomponent could allow an unauthenticated, remote attacker with network access via multiple protocols to compromise Java SE, Java SE Embedded. (CVE-2018-3169) - An unspecified vulnerability in the Java SE component of Oracle Java SE in the JavaFX subcomponent could allow an unauthenticated, remote attacker with network access via multiple protocols to compromise Java SE. (CVE-2018-3209) - An unspecified vulnerability in the Java SE, Java SE Embedded, JRockit component of Oracle Java SE in the JNDI subcomponent could allow an unauthenticated, remote attacker with network access via multiple protocols to compromise Java SE, Java SE Embedded, JRockit. (CVE-2018-3149) - An unspecified vulnerability in the Java SE, Java SE Embedded, JRockit component of Oracle Java SE in the JSSE subcomponent could allow an unauthenticated, remote attacker with network access via SSL/TLS to compromise Java SE, Java SE Embedded, JRockit. (CVE-2018-3180) - An unspecified vulnerability in the Java SE, Java SE Embedded component of Oracle Java SE in the Networking subcomponent could allow an unauthenticated, remote attacker with network access via multiple protocols to compromise Java SE, Java SE Embedded. (CVE-2018-3139) - An unspecified vulnerability in the Java SE, Java SE Embedded, JRockit component of Oracle Java SE in the Scripting subcomponent could allow an unauthenticated, remote attacker with network access via multiple protocols to compromise Java SE, Java SE Embedded, JRockit. (CVE-2018-3183) - An unspecified vulnerability in the Java SE, Java SE Embedded component of Oracle Java SE in the Security subcomponent could allow an unauthenticated, remote attacker with network access via multiple protocols to compromise Java SE, Java SE Embedded. (CVE-2018-3136) - An unspecified vulnerability in the Java SE, Java SE Embedded component of Oracle Java SE in the Serviceability subcomponent could allow a low privileged attacker with logon to the infrastructure where Java SE, Java SE Embedded executes to compromise Java SE, Java SE Embedded. (CVE-2018-3211) - An unspecified vulnerability in the Java SE component of Oracle Java SE in the Sound subcomponent could allow an unauthenticated, remote attacker with network access via multiple protocols to compromise Java SE. (CVE-2018-3157) - An unspecified vulnerability in the Java SE component of Oracle Java SE in the Utility subcomponent could allow an unauthenticated, remote attacker with network access via multiple protocols to compromise Java SE. (CVE-2018-3150) Please consult the CVRF details for the applicable CVEs for additional information. Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Oracle JDK / JRE 11 Update 1, 8 Update 191 / 7 Update 201 / 6 Update 211 or later. If necessary, remove any affected versions. Note that an Extended Support contract with Oracle is needed to obtain JDK / JRE 6 Update 95 or later. See Also: http://www. Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.6.0_211 / 1.7.0_201 / 1.8.0_191 / 1.11.0_1 | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2018-10-19 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 122060 |  | Apache 2.4.x < 2.4.33 Multiple Vulnerabilities | Upgrade to Apache version 2.4.33 or later. | Critical | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: According to its banner, the version of Apache running on the remote host is 2.4.x prior to 2.4.33. It is, therefore, affected by multiple vulnerabilities: - An out of bounds write vulnerability exists in mod_authnz_ldap with AuthLDAPCharsetConfig enabled. An unauthenticated, remote attacker can exploit this, via the Accept-Language header value, to cause the application to stop responding. (CVE-2017-15710) - An arbitrary file upload vulnerability exists in the FilesMatch component where a malicious filename can be crafted to match the expression check for a newline character. An unauthenticated, remote attacker can exploit this, via newline character, to upload arbitrary files on the remote host subject to the privileges of the user. (CVE-2017-15715) - A session management vulnerability exists in the mod_session component due to SessionEnv being enabled and forwarding it's session data to the CGI Application. An unauthenticated, remote attacker can exploit this, via tampering the HTTP_SESSION and using a session header, to influence content. (CVE-2018-1283) - An out of bounds access vulnerability exists when the size limit is reached. An unauthenticated, remote attacker can exploit this, to cause the Apache HTTP Server to crash. (CVE-2018-1301) - A write after free vulnerability exists in HTTP/2 stream due to a NULL pointer being written to an area of freed memory. An unauthenticated, remote attacker can exploit this to execute arbitrary code. (CVE-2018-1302) - An out of bounds read vulnerability exists in mod_cache_socache. An unauthenticated, remote attacker can exploit this, via a specially crafted HTTP request header to cause the application to stop responding. (CVE-2018-1303) - A weak digest vulnerability exists in the HTTP digest authentication challenge. An unauthenticated, remote attacker can exploit this in a cluster of servers configured to use a common digest authentication, to replay HTTP requests across servers without being detected. (CVE-2018-1312) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.33 or later. See Also: https://archive.apache.org/dist/httpd/CHANGES_2.4.33 https://httpd.apache.org/security/vulnerabilities_24.html#2.4.33 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.33 CPE: cpe:/a:apache:http_server cpe:/a:apache:httpd CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-03-21 Patch Publication Date: 2018-03-21 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2019-02-08 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 124198 |  | Oracle Java SE 1.7.0_221 / 1.8.0_211 / 1.11.0_3 / 1.12.0_1 Multiple Vulnerabilities (Apr 2019 CPU) | Upgrade to Oracle JDK / JRE 12 Update 1 , 11 Update 3, 8 Update 211 / 7 Update 221 or later. If necessary, remove any affected versions. | Critical | Synopsis: The remote Windows host contains a programming platform that is affected by multiple vulnerabilities. Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 7 Update 221, 8 Update 211, 11 Update 3, or 12 Update 1. It is, therefore, affected by multiple vulnerabilities related to the following components : - 2D - Libraries - RMI - Windows DLL Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Oracle JDK / JRE 12 Update 1 , 11 Update 3, 8 Update 211 / 7 Update 221 or later. If necessary, remove any affected versions. See Also: http://www.nessus.org/u?9166970d http://www.nessus.org/u?aa0c1228 CVSS Base Score: 9 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:P/A:P Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.7.0_221 / 1.8.0_211 / 1.11.0_3 / 1.12.0_1 CPE: cpe:/a:oracle:jre cpe:/a:oracle:jdk CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2019-04-16 Patch Publication Date: 2019-04-16 Plugin Modification Date: 2022-04-11 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2019-04-19 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 126824 |  | McAfee Agent 5.0.x < 5.0.6 HF1267994 / 5.5.x < 5.5.1 HF1267994 / 5.6.x < 5.6.0 Use-After-Free Vulnerability (SB10258) | Upgrade to McAfee Agent version 5.0.6 HF1267994, 5.5.1 HF1267991, 5.6.0, or later. | Critical | Synopsis: A security management agent installed on the remote host is affected by a use-after-free vulnerability. Description: The version of McAfee Agent, formerly McAfee ePolicy Orchestrator (ePO) Agent, installed on the remote host is 5.0.x prior to 5.0.6 HF1267994, 5.5.x prior to 5.5.1 HF1267991, or 5.6.x prior to 5.6.0. It is, therefore, affected by a use-after-free vulnerability with the remote logging feature. An unauthenticated, remote attacker can exploit this issue, via specifically crafted HTTP header sent to the logging service, to cause the application to stop responding or remotely execute code. Solution: Upgrade to McAfee Agent version 5.0.6 HF1267994, 5.5.1 HF1267991, 5.6.0, or later. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10258 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files (x86)\McAfee\Common Framework\ Installed version : 5.0.6.220 Fixed version : 5.0.6 HF1267994 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2019-05-31 Patch Publication Date: 2019-05-31 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Misc. | Repeat Vulnerability | 2019-07-19 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 128033 |  | Apache 2.4.x < 2.4.41 Multiple Vulnerabilities | Upgrade to Apache version 2.4.41 or later. | Critical | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: The version of Apache httpd installed on the remote host is prior to 2.4.41. It is, therefore, affected by multiple vulnerabilities as referenced in the 2.4.41 advisory, including the following: - A limited cross-site scripting issue was reported affecting the mod_proxy error page. An attacker could cause the link on the error page to be malformed and instead point to a page of their choice. This would only be exploitable where a server was set up with proxying enabled but was misconfigured in such a way that the Proxy Error page was displayed. (CVE-2019-10092) - HTTP/2 (2.4.20 through 2.4.39) very early pushes, for example configured with H2PushResource, could lead to an overwrite of memory in the pushing request's pool, leading to crashes. The memory copied is that of the configured push link header values, not data supplied by the client. (CVE-2019-10081) - Some HTTP/2 implementations are vulnerable to unconstrained internal data buffering, potentially leading to a denial of service. The attacker opens the HTTP/2 window so the peer can send without constraint; however, they leave the TCP window closed so the peer cannot actually write (many of) the bytes on the wire. The attacker then sends a stream of requests for a large response object. Depending on how the servers queue the responses, this can consume excess memory, CPU, or both. (CVE-2019-9517) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.41 or later. See Also: NULL CVSS Base Score: 9.1 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:N/A:P Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.41 CPE: cpe:/a:apache:httpd cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2019-08-13 Patch Publication Date: 2019-08-14 Plugin Modification Date: 2022-12-05 Exploitability Ease: Exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2019-08-20 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 130011 |  | Oracle Java SE 1.7.0_241 / 1.8.0_231 / 1.11.0_5 / 1.13.0_1 Multiple Vulnerabilities (Oct 2019 CPU) (Windows) | Upgrade to Oracle JDK / JRE 13 Update 1, 11 Update 5, 8 Update 231 / 7 Update 241 or later. If necessary, remove any affected versions. | Critical | Synopsis: The remote Windows host contains a programming platform that is affected by multiple vulnerabilities. Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 7 Update 241, 8 Update 231, 11 Update 5, or 13 Update 1. It is, therefore, affected by multiple vulnerabilities related to the following components : - 2D - Libraries - Kerberos - Networking - JavaFX - Hotspot - Scripting - Javadoc - Deployment - Concurrency - JAXP - Serialization - Security Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Oracle JDK / JRE 13 Update 1, 11 Update 5, 8 Update 231 / 7 Update 241 or later. If necessary, remove any affected versions. See Also: http://www.nessus.org/u?2c94f8e4 http://www.nessus.org/u?144b1a0e CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.7.0_241 / 1.8.0_231 / 1.11.0_5 / 1.13.0_1 CPE: cpe:/a:oracle:jre cpe:/a:oracle:jdk CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2019-10-15 Patch Publication Date: 2019-10-15 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2019-10-17 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 150280 |  | Apache 2.4.x < 2.4.47 Multiple Vulnerabilities | Upgrade to Apache version 2.4.47 or later. | Critical | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: The version of Apache httpd installed on the remote host is prior to 2.4.47. It is, therefore, affected by multiple vulnerabilities as referenced in the 2.4.47 changelog: - Unexpected <Location> section matching with 'MergeSlashes OFF' (CVE-2021-30641) - mod_auth_digest: possible stack overflow by one nul byte while validating the Digest nonce. (CVE-2020-35452) - mod_session: Fix possible crash due to NULL pointer dereference, which could be used to cause a Denial of Service with a malicious backend server and SessionHeader. (CVE-2021-26691) - mod_session: Fix possible crash due to NULL pointer dereference, which could be used to cause a Denial of Service. (CVE-2021-26690) - mod_proxy_http: Fix possible crash due to NULL pointer dereference, which could be used to cause a Denial of Service. (CVE-2020-13950) - Windows: Prevent local users from stopping the httpd process (CVE-2020-13938) - mod_proxy_wstunnel, mod_proxy_http: Handle Upgradable protocols end-to-end negotiation. (CVE-2019-17567) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.47 or later. See Also: https://downloads.apache.org/httpd/CHANGES_2.4 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.47 CPE: cpe:/a:apache:httpd cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-06-01 Patch Publication Date: 2021-06-01 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2021-06-04 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 153584 |  | Apache < 2.4.49 Multiple Vulnerabilities | Upgrade to Apache version 2.4.49 or later. | Critical | Synopsis: The remote web server is affected by a vulnerability. Description: The version of Apache httpd installed on the remote host is prior to 2.4.49. It is, therefore, affected by multiple vulnerabilities as referenced in the 2.4.49 changelog. - ap_escape_quotes() may write beyond the end of a buffer when given malicious input. No included modules pass untrusted data to these functions, but third-party / external modules may. (CVE-2021-39275) - Malformed requests may cause the server to dereference a NULL pointer. (CVE-2021-34798) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.49 or later. See Also: https://downloads.apache.org/httpd/CHANGES_2.4 https://httpd.apache.org/security/vulnerabilities_24.html CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.49 CPE: cpe:/a:apache:httpd cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-09-16 Patch Publication Date: 2021-09-16 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2021-09-23 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 156255 |  | Apache 2.4.x >= 2.4.7 / < 2.4.52 Forward Proxy DoS / SSRF | Upgrade to Apache version 2.4.52 or later. | Critical | Synopsis: The remote web server is affected by a denial of service or server-side request forgery vulnerability. Description: The version of Apache httpd installed on the remote host is equal to or greater than 2.4.7 and prior to 2.4.52. It is, therefore, affected by a flaw related to acting as a forward proxy. A crafted URI sent to httpd configured as a forward proxy (ProxyRequests on) can cause a crash (NULL pointer dereference) or, for configurations mixing forward and reverse proxy declarations, can allow for requests to be directed to a declared Unix Domain Socket endpoint (Server Side Request Forgery). Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.52 or later. See Also: NULL CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.52 CPE: cpe:/a:apache:httpd cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-11-18 Patch Publication Date: 2021-12-20 Plugin Modification Date: 2023-04-03 Exploitability Ease: Exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2021-12-23 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 158900 |  | Apache 2.4.x < 2.4.53 Multiple Vulnerabilities | Upgrade to Apache version 2.4.53 or later. | Critical | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: The version of Apache httpd installed on the remote host is prior to 2.4.53. It is, therefore, affected by multiple vulnerabilities as referenced in the 2.4.53 advisory. - mod_lua Use of uninitialized value of in r:parsebody: A carefully crafted request body can cause a read to a random memory area which could cause the process to crash. This issue affects Apache HTTP Server 2.4.52 and earlier. Acknowledgements: Chamal De Silva (CVE-2022-22719) - HTTP request smuggling: Apache HTTP Server 2.4.52 and earlier fails to close inbound connection when errors are encountered discarding the request body, exposing the server to HTTP Request Smuggling Acknowledgements: James Kettle <james.kettle portswigger.net> (CVE-2022-22720) - Possible buffer overflow with very large or unlimited LimitXMLRequestBody in core: If LimitXMLRequestBody is set to allow request bodies larger than 350MB (defaults to 1M) on 32 bit systems an integer overflow happens which later causes out of bounds writes. This issue affects Apache HTTP Server 2.4.52 and earlier. Acknowledgements: Anonymous working with Trend Micro Zero Day Initiative (CVE-2022-22721) - Read/write beyond bounds in mod_sed: Out-of-bounds Write vulnerability in mod_sed of Apache HTTP Server allows an attacker to overwrite heap memory with possibly attacker provided data. This issue affects Apache HTTP Server 2.4 version 2.4.52 and prior versions. Acknowledgements: Ronald Crane (Zippenhop LLC) (CVE-2022-23943) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.53 or later. See Also: http://www.apache.org/dist/httpd/Announcement2.4.html https://httpd.apache.org/security/vulnerabilities_24.html CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.53 CPE: cpe:/a:apache:httpd cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-12-16 Patch Publication Date: 2022-03-14 Plugin Modification Date: 2022-06-15 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2022-03-14 | 2022-04-04 02:14 | 2023-05-09 07:17 | 13
|  |  | 161454 |  | Apache 2.4.x < 2.4.52 mod_lua Buffer Overflow | Upgrade to Apache version 2.4.52 or later. | Critical | Synopsis: The remote web server is affected by a buffer overflow vulnerability. Description: The version of Apache httpd installed on the remote host is prior to 2.4.52. It is, therefore, affected by a flaw related to mod_lua when handling multipart content. A carefully crafted request body can cause a buffer overflow in the mod_lua multipart parser (r:parsebody() called from Lua scripts). The Apache httpd team is not aware of an exploit for the vulnerability though it might be possible to craft one. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.52 or later. See Also: NULL CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.52 CPE: cpe:/a:apache:httpd cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-11-18 Patch Publication Date: 2021-12-20 Plugin Modification Date: 2023-04-03 Exploitability Ease: Exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2022-05-24 | 2022-07-10 08:46 | 2023-05-09 07:17 | 10
|  |  | 161948 |  | Apache 2.4.x < 2.4.54 Multiple Vulnerabilities | Upgrade to Apache version 2.4.54 or later. | Critical | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: The version of Apache httpd installed on the remote host is prior to 2.4.54. It is, therefore, affected by multiple vulnerabilities as referenced in the 2.4.54 advisory. - Possible request smuggling in mod_proxy_ajp: Inconsistent Interpretation of HTTP Requests ('HTTP Request Smuggling') vulnerability in mod_proxy_ajp of Apache HTTP Server allows an attacker to smuggle requests to the AJP server it forwards requests to. This issue affects Apache HTTP Server Apache HTTP Server 2.4 version 2.4.53 and prior versions. Acknowledgements: Ricter Z @ 360 Noah Lab (CVE-2022-26377) - Read beyond bounds in mod_isapi: Apache HTTP Server 2.4.53 and earlier on Windows may read beyond bounds when configured to process requests with the mod_isapi module. Acknowledgements: The Apache HTTP Server project would like to thank Ronald Crane (Zippenhop LLC) for reporting this issue (CVE-2022-28330) - Read beyond bounds via ap_rwrite(): The ap_rwrite() function in Apache HTTP Server 2.4.53 and earlier may read unintended memory if an attacker can cause the server to reflect very large input using ap_rwrite() or ap_rputs(), such as with mod_luas r:puts() function. Acknowledgements: The Apache HTTP Server project would like to thank Ronald Crane (Zippenhop LLC) for reporting this issue (CVE-2022-28614) - Read beyond bounds in ap_strcmp_match(): Apache HTTP Server 2.4.53 and earlier may crash or disclose information due to a read beyond bounds in ap_strcmp_match() when provided with an extremely large input buffer. While no code distributed with the server can be coerced into such a call, third-party modules or lua scripts that use ap_strcmp_match() may hypothetically be affected. Acknowledgements: The Apache HTTP Server project would like to thank Ronald Crane (Zippenhop LLC) for reporting this issue (CVE-2022-28615) - Denial of service in mod_lua r:parsebody: In Apache HTTP Server 2.4.53 and earlier, a malicious request to a lua script that calls r:parsebody(0) may cause a denial of service due to no default limit on possible input size. Acknowledgements: The Apache HTTP Server project would like to thank Ronald Crane (Zippenhop LLC) for reporting this issue (CVE-2022-29404) - Denial of Service mod_sed: If Apache HTTP Server 2.4.53 is configured to do transformations with mod_sed in contexts where the input to mod_sed may be very large, mod_sed may make excessively large memory allocations and trigger an abort. Acknowledgements: This issue was found by Brian Moussalli from the JFrog Security Research team (CVE-2022-30522) - Information Disclosure in mod_lua with websockets: Apache HTTP Server 2.4.53 and earlier may return lengths to applications calling r:wsread() that point past the end of the storage allocated for the buffer. Acknowledgements: The Apache HTTP Server project would like to thank Ronald Crane (Zippenhop LLC) for reporting this issue (CVE-2022-30556) - X-Forwarded-For dropped by hop-by-hop mechanism in mod_proxy: Apache HTTP Server 2.4.53 and earlier may not send the X-Forwarded-* headers to the origin server based on client side Connection header hop-by-hop mechanism. This may be used to bypass IP based authentication on the origin server/application. Acknowledgements: The Apache HTTP Server project would like to thank Gaetan Ferry (Synacktiv) for reporting this issue (CVE-2022-31813) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.54 or later. See Also: https://httpd.apache.org/security/vulnerabilities_24.html CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed ver | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2022-06-08 | 2022-07-10 08:46 | 2023-05-09 07:17 | 10
|  |  | 169789 |  | KB5022346: Windows Server 2012 R2 Security Update (January 2023) | Apply Security Update 5022346 or Cumulative Update 5022352 | Critical | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5022346. It is, therefore, affected by multiple vulnerabilities - Microsoft ODBC Driver Remote Code Execution Vulnerability (CVE-2023-21732) - Microsoft WDAC OLE DB provider for SQL Server Remote Code Execution Vulnerability (CVE-2023-21681) - Windows Advanced Local Procedure Call (ALPC) Elevation of Privilege Vulnerability (CVE-2023-21674) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5022346 or Cumulative Update 5022352 See Also: https://support.microsoft.com/help/5022346 https://support.microsoft.com/help/5022352 CVSS Base Score: 9.1 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5022352 - 5022346 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20778 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-01-10 Patch Publication Date: 2023-01-10 Plugin Modification Date: 2023-03-17 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2023-01-10 | 2023-02-06 05:03 | 2023-05-09 07:17 | 3
|  |  | 170113 |  | Apache 2.4.x < 2.4.55 Multiple Vulnerabilities | Upgrade to Apache version 2.4.55 or later. | Critical | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: The version of Apache httpd installed on the remote host is prior to 2.4.55. It is, therefore, affected by multiple vulnerabilities as referenced in the 2.4.55 advisory. - A carefully crafted If: request header can cause a memory read, or write of a single zero byte, in a pool (heap) memory location beyond the header value sent. This could cause the process to crash. This issue affects Apache HTTP Server 2.4.54 and earlier. (CVE-2006-20001) - Inconsistent Interpretation of HTTP Requests ('HTTP Request Smuggling') vulnerability in mod_proxy_ajp of Apache HTTP Server allows an attacker to smuggle requests to the AJP server it forwards requests to. This issue affects Apache HTTP Server Apache HTTP Server 2.4 version 2.4.54 and prior versions. (CVE-2022-36760) - Prior to Apache HTTP Server 2.4.55, a malicious backend can cause the response headers to be truncated early, resulting in some headers being incorporated into the response body. If the later headers have any security purpose, they will not be interpreted by the client. (CVE-2022-37436) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.55 or later. See Also: NULL CVSS Base Score: 9 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.55 CPE: cpe:/a:apache:httpd cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-07-12 Patch Publication Date: 2023-01-17 Plugin Modification Date: 2023-03-10 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2023-01-18 | 2023-02-06 05:03 | 2023-05-09 07:17 | 3
|  |  | 171453 |  | KB5022894: Windows Server 2012 R2 Security Update (February 2023) | Apply Security Update 5022894 or Cumulative Update 5022899 | Critical | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5022894. It is, therefore, affected by multiple vulnerabilities - Microsoft PostScript Printer Driver Remote Code Execution Vulnerability (CVE-2023-21684, CVE-2023-21801) - Microsoft WDAC OLE DB provider for SQL Server Remote Code Execution Vulnerability (CVE-2023-21685, CVE-2023-21686, CVE-2023-21799) - Microsoft Protected Extensible Authentication Protocol (PEAP) Remote Code Execution Vulnerability (CVE-2023-21689) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5022894 or Cumulative Update 5022899 See Also: https://support.microsoft.com/help/5022894 https://support.microsoft.com/help/5022899 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5022899 - 5022894 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20815 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-02-14 Patch Publication Date: 2023-02-14 Plugin Modification Date: 2023-03-16 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2023-02-14 | 2023-03-07 01:19 | 2023-05-09 07:17 | 2
|  |  | 172186 |  | Apache 2.4.x < 2.4.56 Multiple Vulnerabilities | Upgrade to Apache version 2.4.56 or later. | Critical | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: The version of Apache httpd installed on the remote host is prior to 2.4.56. It is, therefore, affected by multiple vulnerabilities as referenced in the 2.4.56 advisory. - HTTP request splitting with mod_rewrite and mod_proxy: Some mod_proxy configurations on Apache HTTP Server versions 2.4.0 through 2.4.55 allow a HTTP Request Smuggling attack. Configurations are affected when mod_proxy is enabled along with some form of RewriteRule or ProxyPassMatch in which a non-specific pattern matches some portion of the user-supplied request-target (URL) data and is then re-inserted into the proxied request-target using variable substitution. For example, something like: RewriteEngine on RewriteRule ^/here/(.*) http://example.com:8080/elsewhere?$1 http://example.com:8080/elsewhere ; [P] ProxyPassReverse /here/ http://example.com:8080/ http://example.com:8080/ Request splitting/smuggling could result in bypass of access controls in the proxy server, proxying unintended URLs to existing origin servers, and cache poisoning. Acknowledgements: finder: Lars Krapf of Adobe (CVE-2023-25690) - Apache HTTP Server: mod_proxy_uwsgi HTTP response splitting: HTTP Response Smuggling vulnerability in Apache HTTP Server via mod_proxy_uwsgi. This issue affects Apache HTTP Server: from 2.4.30 through 2.4.55. Special characters in the origin response header can truncate/split the response forwarded to the client. Acknowledgements: finder: Dimas Fariski Setyawan Putra (nyxsorcerer) (CVE-2023-27522) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.56 or later. See Also: NULL CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.56 CPE: cpe:/a:apache:httpd cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-01-29 Patch Publication Date: 2023-03-07 Plugin Modification Date: 2023-03-15 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2023-03-07 | 2023-04-04 04:16 | 2023-05-09 07:17 | 1
|  |  | 172535 |  | KB5023764: Windows Server 2012 R2 Security Update (March 2023) | Apply Security Update 5023764 or Cumulative Update 5023765 | Critical | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5023764. It is, therefore, affected by multiple vulnerabilities - Internet Control Message Protocol (ICMP) Remote Code Execution Vulnerability (CVE-2023-23415) - Remote Procedure Call Runtime Remote Code Execution Vulnerability (CVE-2023-21708, CVE-2023-23405, CVE-2023-24869, CVE-2023-24908) - Windows Point-to-Point Protocol over Ethernet (PPPoE) Elevation of Privilege Vulnerability (CVE-2023-23385) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5023764 or Cumulative Update 5023765 See Also: https://support.microsoft.com/help/5023764 https://support.microsoft.com/help/5023765 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5023765 - 5023764 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20876 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-03-14 Patch Publication Date: 2023-03-14 Plugin Modification Date: 2023-04-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2023-03-14 | 2023-04-04 04:16 | 2023-05-09 07:17 | 1
|  |  | 72704 |  | Microsoft .NET Framework Unsupported | Upgrade to a version of the Microsoft .NET Framework that is currently supported. | Critical | Synopsis: An unsupported software framework is installed on the remote Windows host. Description: According to its self-reported version number, there is at least one version of Microsoft .NET Framework installed on the remote Windows host that is no longer supported. Lack of support implies that no new security patches for the product will be released by the vendor. As a result, it is likely to contain security vulnerabilities. Solution: Upgrade to a version of the Microsoft .NET Framework that is currently supported. See Also: http://www.nessus.org/u?930a7a91 http://www.nessus.org/u?3b10ac8d CVSS Base Score: 10 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The following Microsoft .NET Framework version is no longer supported : Installed version : Microsoft .NET Framework v4.6.1 EOL date : April 26, 2022 EOL URL : https://docs.microsoft.com/en-us/lifecycle/products/microsoft-net-framework Supported versions : 3.5 / 4.6.2 / 4.7 / 4.7.1 / 4.7.2/ 4.8 CPE: cpe:/a:microsoft:.net_framework CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2022-04-26 Exploitability Ease: NULL Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2014-02-26 | 2022-05-02 07:06 | 2023-05-09 07:10 | 12
|  |  | 126824 |  | McAfee Agent 5.0.x < 5.0.6 HF1267994 / 5.5.x < 5.5.1 HF1267994 / 5.6.x < 5.6.0 Use-After-Free Vulnerability (SB10258) | Upgrade to McAfee Agent version 5.0.6 HF1267994, 5.5.1 HF1267991, 5.6.0, or later. | Critical | Synopsis: A security management agent installed on the remote host is affected by a use-after-free vulnerability. Description: The version of McAfee Agent, formerly McAfee ePolicy Orchestrator (ePO) Agent, installed on the remote host is 5.0.x prior to 5.0.6 HF1267994, 5.5.x prior to 5.5.1 HF1267991, or 5.6.x prior to 5.6.0. It is, therefore, affected by a use-after-free vulnerability with the remote logging feature. An unauthenticated, remote attacker can exploit this issue, via specifically crafted HTTP header sent to the logging service, to cause the application to stop responding or remotely execute code. Solution: Upgrade to McAfee Agent version 5.0.6 HF1267994, 5.5.1 HF1267991, 5.6.0, or later. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10258 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files\McAfee\Agent\ Installed version : 5.5.1.342 Fixed version : 5.5.1 HF1267991 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2019-05-31 Patch Publication Date: 2019-05-31 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Misc. | Repeat Vulnerability | 2019-07-19 | 2022-02-06 21:50 | 2023-05-09 07:10 | 15
|  |  | 169789 |  | KB5022346: Windows Server 2012 R2 Security Update (January 2023) | Apply Security Update 5022346 or Cumulative Update 5022352 | Critical | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5022346. It is, therefore, affected by multiple vulnerabilities - Microsoft ODBC Driver Remote Code Execution Vulnerability (CVE-2023-21732) - Microsoft WDAC OLE DB provider for SQL Server Remote Code Execution Vulnerability (CVE-2023-21681) - Windows Advanced Local Procedure Call (ALPC) Elevation of Privilege Vulnerability (CVE-2023-21674) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5022346 or Cumulative Update 5022352 See Also: https://support.microsoft.com/help/5022346 https://support.microsoft.com/help/5022352 CVSS Base Score: 9.1 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5022352 - 5022346 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20778 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-01-10 Patch Publication Date: 2023-01-10 Plugin Modification Date: 2023-03-17 Exploitability Ease: Exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2023-01-10 | 2023-02-06 05:03 | 2023-05-09 07:10 | 3
|  |  | 171453 |  | KB5022894: Windows Server 2012 R2 Security Update (February 2023) | Apply Security Update 5022894 or Cumulative Update 5022899 | Critical | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5022894. It is, therefore, affected by multiple vulnerabilities - Microsoft PostScript Printer Driver Remote Code Execution Vulnerability (CVE-2023-21684, CVE-2023-21801) - Microsoft WDAC OLE DB provider for SQL Server Remote Code Execution Vulnerability (CVE-2023-21685, CVE-2023-21686, CVE-2023-21799) - Microsoft Protected Extensible Authentication Protocol (PEAP) Remote Code Execution Vulnerability (CVE-2023-21689) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5022894 or Cumulative Update 5022899 See Also: https://support.microsoft.com/help/5022894 https://support.microsoft.com/help/5022899 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5022899 - 5022894 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20815 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-02-14 Patch Publication Date: 2023-02-14 Plugin Modification Date: 2023-03-16 Exploitability Ease: Exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2023-02-14 | 2023-03-07 01:19 | 2023-05-09 07:10 | 2
|  |  | 172535 |  | KB5023764: Windows Server 2012 R2 Security Update (March 2023) | Apply Security Update 5023764 or Cumulative Update 5023765 | Critical | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5023764. It is, therefore, affected by multiple vulnerabilities - Internet Control Message Protocol (ICMP) Remote Code Execution Vulnerability (CVE-2023-23415) - Remote Procedure Call Runtime Remote Code Execution Vulnerability (CVE-2023-21708, CVE-2023-23405, CVE-2023-24869, CVE-2023-24908) - Windows Point-to-Point Protocol over Ethernet (PPPoE) Elevation of Privilege Vulnerability (CVE-2023-23385) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5023764 or Cumulative Update 5023765 See Also: https://support.microsoft.com/help/5023764 https://support.microsoft.com/help/5023765 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5023765 - 5023764 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20876 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-03-14 Patch Publication Date: 2023-03-14 Plugin Modification Date: 2023-04-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2023-03-14 | 2023-04-04 04:16 | 2023-05-09 07:10 | 1
|  |  | 72704 |  | Microsoft .NET Framework Unsupported | Upgrade to a version of the Microsoft .NET Framework that is currently supported. | Critical | Synopsis: An unsupported software framework is installed on the remote Windows host. Description: According to its self-reported version number, there is at least one version of Microsoft .NET Framework installed on the remote Windows host that is no longer supported. Lack of support implies that no new security patches for the product will be released by the vendor. As a result, it is likely to contain security vulnerabilities. Solution: Upgrade to a version of the Microsoft .NET Framework that is currently supported. See Also: http://www.nessus.org/u?930a7a91 http://www.nessus.org/u?3b10ac8d CVSS Base Score: 10 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The following Microsoft .NET Framework version is no longer supported : Installed version : Microsoft .NET Framework v4.6.1 EOL date : April 26, 2022 EOL URL : https://docs.microsoft.com/en-us/lifecycle/products/microsoft-net-framework Supported versions : 3.5 / 4.6.2 / 4.7 / 4.7.1 / 4.7.2/ 4.8 CPE: cpe:/a:microsoft:.net_framework CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2022-04-26 Exploitability Ease: NULL Plugin Type: local Source File: | Report Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE2 |  |  |  |  | VHAXXXAPPROCHE2.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2014-02-26 | 2022-05-02 07:06 | 2023-05-09 07:17 | 12
|  |  | 169789 |  | KB5022346: Windows Server 2012 R2 Security Update (January 2023) | Apply Security Update 5022346 or Cumulative Update 5022352 | Critical | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5022346. It is, therefore, affected by multiple vulnerabilities - Microsoft ODBC Driver Remote Code Execution Vulnerability (CVE-2023-21732) - Microsoft WDAC OLE DB provider for SQL Server Remote Code Execution Vulnerability (CVE-2023-21681) - Windows Advanced Local Procedure Call (ALPC) Elevation of Privilege Vulnerability (CVE-2023-21674) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5022346 or Cumulative Update 5022352 See Also: https://support.microsoft.com/help/5022346 https://support.microsoft.com/help/5022352 CVSS Base Score: 9.1 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5022352 - 5022346 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20778 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-01-10 Patch Publication Date: 2023-01-10 Plugin Modification Date: 2023-03-17 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Report Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE2 |  |  |  |  | VHAXXXAPPROCHE2.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2023-01-10 | 2023-02-06 05:03 | 2023-05-09 07:17 | 3
|  |  | 171453 |  | KB5022894: Windows Server 2012 R2 Security Update (February 2023) | Apply Security Update 5022894 or Cumulative Update 5022899 | Critical | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5022894. It is, therefore, affected by multiple vulnerabilities - Microsoft PostScript Printer Driver Remote Code Execution Vulnerability (CVE-2023-21684, CVE-2023-21801) - Microsoft WDAC OLE DB provider for SQL Server Remote Code Execution Vulnerability (CVE-2023-21685, CVE-2023-21686, CVE-2023-21799) - Microsoft Protected Extensible Authentication Protocol (PEAP) Remote Code Execution Vulnerability (CVE-2023-21689) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5022894 or Cumulative Update 5022899 See Also: https://support.microsoft.com/help/5022894 https://support.microsoft.com/help/5022899 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5022899 - 5022894 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20815 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-02-14 Patch Publication Date: 2023-02-14 Plugin Modification Date: 2023-03-16 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Report Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE2 |  |  |  |  | VHAXXXAPPROCHE2.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2023-02-14 | 2023-03-07 01:19 | 2023-05-09 07:17 | 2
|  |  | 172535 |  | KB5023764: Windows Server 2012 R2 Security Update (March 2023) | Apply Security Update 5023764 or Cumulative Update 5023765 | Critical | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5023764. It is, therefore, affected by multiple vulnerabilities - Internet Control Message Protocol (ICMP) Remote Code Execution Vulnerability (CVE-2023-23415) - Remote Procedure Call Runtime Remote Code Execution Vulnerability (CVE-2023-21708, CVE-2023-23405, CVE-2023-24869, CVE-2023-24908) - Windows Point-to-Point Protocol over Ethernet (PPPoE) Elevation of Privilege Vulnerability (CVE-2023-23385) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5023764 or Cumulative Update 5023765 See Also: https://support.microsoft.com/help/5023764 https://support.microsoft.com/help/5023765 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5023765 - 5023764 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20876 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-03-14 Patch Publication Date: 2023-03-14 Plugin Modification Date: 2023-04-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Report Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE2 |  |  |  |  | VHAXXXAPPROCHE2.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2023-03-14 | 2023-04-04 04:16 | 2023-05-09 07:17 | 1
|  |  | 150280 |  | Apache 2.4.x < 2.4.47 Multiple Vulnerabilities | Upgrade to Apache version 2.4.47 or later. | Critical | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: The version of Apache httpd installed on the remote host is prior to 2.4.47. It is, therefore, affected by multiple vulnerabilities as referenced in the 2.4.47 changelog: - Unexpected <Location> section matching with 'MergeSlashes OFF' (CVE-2021-30641) - mod_auth_digest: possible stack overflow by one nul byte while validating the Digest nonce. (CVE-2020-35452) - mod_session: Fix possible crash due to NULL pointer dereference, which could be used to cause a Denial of Service with a malicious backend server and SessionHeader. (CVE-2021-26691) - mod_session: Fix possible crash due to NULL pointer dereference, which could be used to cause a Denial of Service. (CVE-2021-26690) - mod_proxy_http: Fix possible crash due to NULL pointer dereference, which could be used to cause a Denial of Service. (CVE-2020-13950) - Windows: Prevent local users from stopping the httpd process (CVE-2020-13938) - mod_proxy_wstunnel, mod_proxy_http: Handle Upgradable protocols end-to-end negotiation. (CVE-2019-17567) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.47 or later. See Also: https://downloads.apache.org/httpd/CHANGES_2.4 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Roche\CobasInfinity\Apache24\bin\httpd.exe Installed version : 2.4.46.0 Fixed version : 2.4.47 CPE: cpe:/a:apache:httpd cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-06-01 Patch Publication Date: 2021-06-01 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | COBAS Infinity Production Server (v3.03) | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXCCACOBASP |  |  |  |  | VHAXXXCCACOBASP.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2021-06-04 | 2022-09-04 23:24 | 2023-05-09 07:18 | 8
|  |  | 153584 |  | Apache < 2.4.49 Multiple Vulnerabilities | Upgrade to Apache version 2.4.49 or later. | Critical | Synopsis: The remote web server is affected by a vulnerability. Description: The version of Apache httpd installed on the remote host is prior to 2.4.49. It is, therefore, affected by multiple vulnerabilities as referenced in the 2.4.49 changelog. - ap_escape_quotes() may write beyond the end of a buffer when given malicious input. No included modules pass untrusted data to these functions, but third-party / external modules may. (CVE-2021-39275) - Malformed requests may cause the server to dereference a NULL pointer. (CVE-2021-34798) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.49 or later. See Also: https://downloads.apache.org/httpd/CHANGES_2.4 https://httpd.apache.org/security/vulnerabilities_24.html CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Roche\CobasInfinity\Apache24\bin\httpd.exe Installed version : 2.4.46.0 Fixed version : 2.4.49 CPE: cpe:/a:apache:httpd cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-09-16 Patch Publication Date: 2021-09-16 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | COBAS Infinity Production Server (v3.03) | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXCCACOBASP |  |  |  |  | VHAXXXCCACOBASP.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2021-09-23 | 2022-09-04 23:24 | 2023-05-09 07:18 | 8
|  |  | 156255 |  | Apache 2.4.x >= 2.4.7 / < 2.4.52 Forward Proxy DoS / SSRF | Upgrade to Apache version 2.4.52 or later. | Critical | Synopsis: The remote web server is affected by a denial of service or server-side request forgery vulnerability. Description: The version of Apache httpd installed on the remote host is equal to or greater than 2.4.7 and prior to 2.4.52. It is, therefore, affected by a flaw related to acting as a forward proxy. A crafted URI sent to httpd configured as a forward proxy (ProxyRequests on) can cause a crash (NULL pointer dereference) or, for configurations mixing forward and reverse proxy declarations, can allow for requests to be directed to a declared Unix Domain Socket endpoint (Server Side Request Forgery). Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.52 or later. See Also: NULL CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Roche\CobasInfinity\Apache24\bin\httpd.exe Installed version : 2.4.46.0 Fixed version : 2.4.52 CPE: cpe:/a:apache:httpd cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-11-18 Patch Publication Date: 2021-12-20 Plugin Modification Date: 2023-04-03 Exploitability Ease: Exploits are available Plugin Type: combined Source File: | COBAS Infinity Production Server (v3.03) | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXCCACOBASP |  |  |  |  | VHAXXXCCACOBASP.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2021-12-23 | 2022-09-04 23:24 | 2023-05-09 07:18 | 8
|  |  | 158900 |  | Apache 2.4.x < 2.4.53 Multiple Vulnerabilities | Upgrade to Apache version 2.4.53 or later. | Critical | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: The version of Apache httpd installed on the remote host is prior to 2.4.53. It is, therefore, affected by multiple vulnerabilities as referenced in the 2.4.53 advisory. - mod_lua Use of uninitialized value of in r:parsebody: A carefully crafted request body can cause a read to a random memory area which could cause the process to crash. This issue affects Apache HTTP Server 2.4.52 and earlier. Acknowledgements: Chamal De Silva (CVE-2022-22719) - HTTP request smuggling: Apache HTTP Server 2.4.52 and earlier fails to close inbound connection when errors are encountered discarding the request body, exposing the server to HTTP Request Smuggling Acknowledgements: James Kettle <james.kettle portswigger.net> (CVE-2022-22720) - Possible buffer overflow with very large or unlimited LimitXMLRequestBody in core: If LimitXMLRequestBody is set to allow request bodies larger than 350MB (defaults to 1M) on 32 bit systems an integer overflow happens which later causes out of bounds writes. This issue affects Apache HTTP Server 2.4.52 and earlier. Acknowledgements: Anonymous working with Trend Micro Zero Day Initiative (CVE-2022-22721) - Read/write beyond bounds in mod_sed: Out-of-bounds Write vulnerability in mod_sed of Apache HTTP Server allows an attacker to overwrite heap memory with possibly attacker provided data. This issue affects Apache HTTP Server 2.4 version 2.4.52 and prior versions. Acknowledgements: Ronald Crane (Zippenhop LLC) (CVE-2022-23943) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.53 or later. See Also: http://www.apache.org/dist/httpd/Announcement2.4.html https://httpd.apache.org/security/vulnerabilities_24.html CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Roche\CobasInfinity\Apache24\bin\httpd.exe Installed version : 2.4.46.0 Fixed version : 2.4.53 CPE: cpe:/a:apache:httpd cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-12-16 Patch Publication Date: 2022-03-14 Plugin Modification Date: 2022-06-15 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | COBAS Infinity Production Server (v3.03) | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXCCACOBASP |  |  |  |  | VHAXXXCCACOBASP.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2022-03-14 | 2022-09-04 23:24 | 2023-05-09 07:18 | 8
|  |  | 161454 |  | Apache 2.4.x < 2.4.52 mod_lua Buffer Overflow | Upgrade to Apache version 2.4.52 or later. | Critical | Synopsis: The remote web server is affected by a buffer overflow vulnerability. Description: The version of Apache httpd installed on the remote host is prior to 2.4.52. It is, therefore, affected by a flaw related to mod_lua when handling multipart content. A carefully crafted request body can cause a buffer overflow in the mod_lua multipart parser (r:parsebody() called from Lua scripts). The Apache httpd team is not aware of an exploit for the vulnerability though it might be possible to craft one. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.52 or later. See Also: NULL CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Roche\CobasInfinity\Apache24\bin\httpd.exe Installed version : 2.4.46.0 Fixed version : 2.4.52 CPE: cpe:/a:apache:httpd cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-11-18 Patch Publication Date: 2021-12-20 Plugin Modification Date: 2023-04-03 Exploitability Ease: Exploits are available Plugin Type: combined Source File: | COBAS Infinity Production Server (v3.03) | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXCCACOBASP |  |  |  |  | VHAXXXCCACOBASP.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2022-05-24 | 2022-09-04 23:24 | 2023-05-09 07:18 | 8
|  |  | 161948 |  | Apache 2.4.x < 2.4.54 Multiple Vulnerabilities | Upgrade to Apache version 2.4.54 or later. | Critical | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: The version of Apache httpd installed on the remote host is prior to 2.4.54. It is, therefore, affected by multiple vulnerabilities as referenced in the 2.4.54 advisory. - Possible request smuggling in mod_proxy_ajp: Inconsistent Interpretation of HTTP Requests ('HTTP Request Smuggling') vulnerability in mod_proxy_ajp of Apache HTTP Server allows an attacker to smuggle requests to the AJP server it forwards requests to. This issue affects Apache HTTP Server Apache HTTP Server 2.4 version 2.4.53 and prior versions. Acknowledgements: Ricter Z @ 360 Noah Lab (CVE-2022-26377) - Read beyond bounds in mod_isapi: Apache HTTP Server 2.4.53 and earlier on Windows may read beyond bounds when configured to process requests with the mod_isapi module. Acknowledgements: The Apache HTTP Server project would like to thank Ronald Crane (Zippenhop LLC) for reporting this issue (CVE-2022-28330) - Read beyond bounds via ap_rwrite(): The ap_rwrite() function in Apache HTTP Server 2.4.53 and earlier may read unintended memory if an attacker can cause the server to reflect very large input using ap_rwrite() or ap_rputs(), such as with mod_luas r:puts() function. Acknowledgements: The Apache HTTP Server project would like to thank Ronald Crane (Zippenhop LLC) for reporting this issue (CVE-2022-28614) - Read beyond bounds in ap_strcmp_match(): Apache HTTP Server 2.4.53 and earlier may crash or disclose information due to a read beyond bounds in ap_strcmp_match() when provided with an extremely large input buffer. While no code distributed with the server can be coerced into such a call, third-party modules or lua scripts that use ap_strcmp_match() may hypothetically be affected. Acknowledgements: The Apache HTTP Server project would like to thank Ronald Crane (Zippenhop LLC) for reporting this issue (CVE-2022-28615) - Denial of service in mod_lua r:parsebody: In Apache HTTP Server 2.4.53 and earlier, a malicious request to a lua script that calls r:parsebody(0) may cause a denial of service due to no default limit on possible input size. Acknowledgements: The Apache HTTP Server project would like to thank Ronald Crane (Zippenhop LLC) for reporting this issue (CVE-2022-29404) - Denial of Service mod_sed: If Apache HTTP Server 2.4.53 is configured to do transformations with mod_sed in contexts where the input to mod_sed may be very large, mod_sed may make excessively large memory allocations and trigger an abort. Acknowledgements: This issue was found by Brian Moussalli from the JFrog Security Research team (CVE-2022-30522) - Information Disclosure in mod_lua with websockets: Apache HTTP Server 2.4.53 and earlier may return lengths to applications calling r:wsread() that point past the end of the storage allocated for the buffer. Acknowledgements: The Apache HTTP Server project would like to thank Ronald Crane (Zippenhop LLC) for reporting this issue (CVE-2022-30556) - X-Forwarded-For dropped by hop-by-hop mechanism in mod_proxy: Apache HTTP Server 2.4.53 and earlier may not send the X-Forwarded-* headers to the origin server based on client side Connection header hop-by-hop mechanism. This may be used to bypass IP based authentication on the origin server/application. Acknowledgements: The Apache HTTP Server project would like to thank Gaetan Ferry (Synacktiv) for reporting this issue (CVE-2022-31813) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.54 or later. See Also: https://httpd.apache.org/security/vulnerabilities_24.html CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Roche\CobasInfinity\Apache24\bin\httpd.exe Installed version : 2.4.46.0 Fixed vers | COBAS Infinity Production Server (v3.03) | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXCCACOBASP |  |  |  |  | VHAXXXCCACOBASP.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2022-06-08 | 2022-09-04 23:24 | 2023-05-09 07:18 | 8
|  |  | 170113 |  | Apache 2.4.x < 2.4.55 Multiple Vulnerabilities | Upgrade to Apache version 2.4.55 or later. | Critical | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: The version of Apache httpd installed on the remote host is prior to 2.4.55. It is, therefore, affected by multiple vulnerabilities as referenced in the 2.4.55 advisory. - A carefully crafted If: request header can cause a memory read, or write of a single zero byte, in a pool (heap) memory location beyond the header value sent. This could cause the process to crash. This issue affects Apache HTTP Server 2.4.54 and earlier. (CVE-2006-20001) - Inconsistent Interpretation of HTTP Requests ('HTTP Request Smuggling') vulnerability in mod_proxy_ajp of Apache HTTP Server allows an attacker to smuggle requests to the AJP server it forwards requests to. This issue affects Apache HTTP Server Apache HTTP Server 2.4 version 2.4.54 and prior versions. (CVE-2022-36760) - Prior to Apache HTTP Server 2.4.55, a malicious backend can cause the response headers to be truncated early, resulting in some headers being incorporated into the response body. If the later headers have any security purpose, they will not be interpreted by the client. (CVE-2022-37436) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.55 or later. See Also: NULL CVSS Base Score: 9 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Roche\CobasInfinity\Apache24\bin\httpd.exe Installed version : 2.4.46.0 Fixed version : 2.4.55 CPE: cpe:/a:apache:httpd cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-07-12 Patch Publication Date: 2023-01-17 Plugin Modification Date: 2023-03-10 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | COBAS Infinity Production Server (v3.03) | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXCCACOBASP |  |  |  |  | VHAXXXCCACOBASP.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2023-01-18 | 2023-02-06 05:03 | 2023-05-09 07:18 | 3
|  |  | 172186 |  | Apache 2.4.x < 2.4.56 Multiple Vulnerabilities | Upgrade to Apache version 2.4.56 or later. | Critical | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: The version of Apache httpd installed on the remote host is prior to 2.4.56. It is, therefore, affected by multiple vulnerabilities as referenced in the 2.4.56 advisory. - HTTP request splitting with mod_rewrite and mod_proxy: Some mod_proxy configurations on Apache HTTP Server versions 2.4.0 through 2.4.55 allow a HTTP Request Smuggling attack. Configurations are affected when mod_proxy is enabled along with some form of RewriteRule or ProxyPassMatch in which a non-specific pattern matches some portion of the user-supplied request-target (URL) data and is then re-inserted into the proxied request-target using variable substitution. For example, something like: RewriteEngine on RewriteRule ^/here/(.*) http://example.com:8080/elsewhere?$1 http://example.com:8080/elsewhere ; [P] ProxyPassReverse /here/ http://example.com:8080/ http://example.com:8080/ Request splitting/smuggling could result in bypass of access controls in the proxy server, proxying unintended URLs to existing origin servers, and cache poisoning. Acknowledgements: finder: Lars Krapf of Adobe (CVE-2023-25690) - Apache HTTP Server: mod_proxy_uwsgi HTTP response splitting: HTTP Response Smuggling vulnerability in Apache HTTP Server via mod_proxy_uwsgi. This issue affects Apache HTTP Server: from 2.4.30 through 2.4.55. Special characters in the origin response header can truncate/split the response forwarded to the client. Acknowledgements: finder: Dimas Fariski Setyawan Putra (nyxsorcerer) (CVE-2023-27522) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.56 or later. See Also: NULL CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Roche\CobasInfinity\Apache24\bin\httpd.exe Installed version : 2.4.46.0 Fixed version : 2.4.56 CPE: cpe:/a:apache:httpd cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-01-29 Patch Publication Date: 2023-03-07 Plugin Modification Date: 2023-03-15 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | COBAS Infinity Production Server (v3.03) | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXCCACOBASP |  |  |  |  | VHAXXXCCACOBASP.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2023-03-07 | 2023-04-04 04:16 | 2023-05-09 07:18 | 1
|  |  | 172532 |  | KB5023697: Windows 10 Version 1607 and Windows Server 2016 Security Update (March 2023) | Apply Security Update 5023697 | Critical | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5023697. It is, therefore, affected by multiple vulnerabilities - An out-of-bounds write vulnerability exists in TPM2.0's Module Library allowing writing of a 2-byte data past the end of TPM2.0 command in the CryptParameterDecryption routine. An attacker who can successfully exploit this vulnerability can lead to denial of service (crashing the TPM chip/process or rendering it unusable) and/or arbitrary code execution in the TPM context. (CVE-2023-1017) - An out-of-bounds read vulnerability exists in TPM2.0's Module Library allowing a 2-byte read past the end of a TPM2.0 command in the CryptParameterDecryption routine. An attacker who can successfully exploit this vulnerability can read or access sensitive data stored in the TPM. (CVE-2023-1018) - Remote Procedure Call Runtime Remote Code Execution Vulnerability (CVE-2023-21708, CVE-2023-23405, CVE-2023-24869, CVE-2023-24908) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5023697 See Also: https://support.microsoft.com/help/5023697 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5023697 - C:\WINDOWS\system32\ntoskrnl.exe has not been patched. Remote version : 10.0.14393.5717 Should be : 10.0.14393.5786 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-02-28 Patch Publication Date: 2023-03-14 Plugin Modification Date: 2023-04-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | COBAS Infinity Production Server (v3.03) | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXCCACOBASP |  |  |  |  | VHAXXXCCACOBASP.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2023-03-14 | 2023-04-04 04:16 | 2023-05-09 07:18 | 1
|  |  | 34460 |  | Unsupported Web Server Detection | Remove the web server if it is no longer needed. Otherwise, upgrade to a supported version if possible or switch to another server. | Critical | Synopsis: The remote web server is obsolete / unsupported. Description: According to its version, the remote web server is obsolete and no longer maintained by its vendor or provider. Lack of support implies that no new security patches for the product will be released by the vendor. As a result, it may contain security vulnerabilities. Solution: Remove the web server if it is no longer needed. Otherwise, upgrade to a supported version if possible or switch to another server. See Also: NULL CVSS Base Score: 10 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Product : Microsoft IIS 7.5 Server response header : Microsoft-IIS/7.5 Support ended : 2020-01-14 Supported versions : Microsoft IIS 8.5 / 8.0 Additional information : http://www.nessus.org/u?d8353958 CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-02-10 Exploitability Ease: NULL Plugin Type: remote Source File: | Methadone Avatar Server | Dell |  |  | Region X | VISN X | VAMC | VHAXXXMETH1 |  |  |  |  | VHAXXXMETH1.vxx.med.va.gov |  | Microsoft Windows Server 2008 R2 | tcp | 80 | Web Servers | Repeat Vulnerability | 2008-10-21 | 2020-08-10 23:38 | 2023-05-09 06:52 | 33
|  |  | 73756 |  | Microsoft SQL Server Unsupported Version Detection (remote check) | Upgrade to a version of Microsoft SQL Server that is currently supported. | Critical | Synopsis: An unsupported version of a database server is running on the remote host. Description: According to its self-reported version number, the installation of Microsoft SQL Server on the remote host is no longer supported. Lack of support implies that no new security patches for the product will be released by the vendor. As a result, it is likely to contain security vulnerabilities. Solution: Upgrade to a version of Microsoft SQL Server that is currently supported. See Also: http://www.nessus.org/u?d4418a57 CVSS Base Score: 10 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The following unsupported installation of Microsoft SQL Server was detected : Installed version : 12.0.2000.0 Fixed version : 12.0.6024.0 (2014 SP3) CPE: cpe:/a:microsoft:sql_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2022-12-06 Exploitability Ease: NULL Plugin Type: remote Source File: | Spacelabs DB Server | SPACELABS HEALTHCARE INC |  |  | Region X | VISN X | VAMC | VHAXXXSQLICSDB |  |  |  |  | VHAXXXSQLICSDB.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 1433 | Databases | Repeat Vulnerability | 2014-04-29 | 2020-12-07 06:27 | 2023-05-07 09:49 | 29
|  |  | 100464 |  | Microsoft Windows SMBv1 Multiple Vulnerabilities | Apply the applicable security update for your Windows version : - Windows Server 2008 : KB4018466 - Windows 7 : KB4019264 - Windows Server 2008 R2 : KB4019264 - Windows Server 2012 : KB4019216 - Windows 8.1 / RT 8.1. : KB4019215 - Windows Server 2012 R2 : KB4019215 - Windows 10 : KB4019474 - Windows 10 Version 1511 : KB4019473 - Windows 10 Version 1607 : KB4019472 - Windows 10 Version 1703 : KB4016871 - Windows Server 2016 : KB4019472 | High | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host has Microsoft Server Message Block 1.0 (SMBv1) enabled. It is, therefore, affected by multiple vulnerabilities : - Multiple information disclosure vulnerabilities exist in Microsoft Server Message Block 1.0 (SMBv1) due to improper handling of SMBv1 packets. An unauthenticated, remote attacker can exploit these vulnerabilities, via a specially crafted SMBv1 packet, to disclose sensitive information. (CVE-2017-0267, CVE-2017-0268, CVE-2017-0270, CVE-2017-0271, CVE-2017-0274, CVE-2017-0275, CVE-2017-0276) - Multiple denial of service vulnerabilities exist in Microsoft Server Message Block 1.0 (SMBv1) due to improper handling of requests. An unauthenticated, remote attacker can exploit these vulnerabilities, via a specially crafted SMB request, to cause the system to stop responding. (CVE-2017-0269, CVE-2017-0273, CVE-2017-0280) - Multiple remote code execution vulnerabilities exist in Microsoft Server Message Block 1.0 (SMBv1) due to improper handling of SMBv1 packets. An unauthenticated, remote attacker can exploit these vulnerabilities, via a specially crafted SMBv1 packet, to execute arbitrary code. (CVE-2017-0272, CVE-2017-0277, CVE-2017-0278, CVE-2017-0279) Depending on the host's security policy configuration, this plugin cannot always correctly determine if the Windows host is vulnerable if the host is running a later Windows version (i.e., Windows 8.1, 10, 2012, 2012 R2, and 2016) specifically that named pipes and shares are allowed to be accessed remotely and anonymously. Tenable does not recommend this configuration, and the hosts should be checked locally for patches with one of the following plugins, depending on the Windows version : 100054, 100055, 100057, 100059, 100060, or 100061. Solution: Apply the applicable security update for your Windows version : - Windows Server 2008 : KB4018466 - Windows 7 : KB4019264 - Windows Server 2008 R2 : KB4019264 - Windows Server 2012 : KB4019216 - Windows 8.1 / RT 8.1. : KB4019215 - Windows Server 2012 R2 : KB4019215 - Windows 10 : KB4019474 - Windows 10 Version 1511 : KB4019473 - Windows 10 Version 1607 : KB4019472 - Windows 10 Version 1703 : KB4016871 - Windows Server 2016 : KB4019472 See Also: http://www.nessus.org/u?c21268d4 http://www.nessus.org/u?b9253982 http://www.nessus.org/u?23802c83 http://www.nessus.org/u?8313bb60 http://www.nessus.org/u?7677c678 http://www.nessus.org/u?36da236c http://www.nessus.org/u?0981b934 http://www.nessus.org/u?c88efefa http://www.nessus.org/u?695bf5cc http://www.nessus.org/u?459a1e8c http://www.nessus.org/u?ea45bbc5 http://www.nessus.org/u?4195776a http://www.nessus.org/u?fbf092cf http://www.nessus.org/u?8c0cc566 CVSS Base Score: 8.1 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:C/I:C/A:C Plugin Output: NULL CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2017-05-09 Patch Publication Date: 2017-05-09 Plugin Modification Date: 2019-11-13 Exploitability Ease: No known exploits are available Plugin Type: remote Source File: | Sysmex SP-50 #1 (Host) | SYSMEX AMERICA INC |  |  | Region X | VISN X | VAMC | [7 empty] | Microsoft Windows Embedded Standard 7 | tcp | 445 | Windows | Repeat Vulnerability | 2017-05-26 | 2022-02-06 21:56 | 2023-05-09 07:18 | 15
|  |  | 121124 |  | Apache Tomcat 8.0.x < 8.0.52 / 8.5.x < 8.5.31 / 9.0.x < 9.0.8 Denial of Service | Upgrade to Apache Tomcat version 8.0.52 / 8.5.31 / 9.0.8 or later. | High | Synopsis: The remote Apache Tomcat server is affected by a denial of service vulnerability. Description: According to its self-reported version number, the Apache Tomcat instance listening on the remote host is 8.0.x < 8.0.52, 8.5.x < 8.5.31 or 9.0.x < 9.0.8. It is, therefore, affected by the following vulnerability: - A denial of service (DoS) vulnerability exists in Tomcat due to improper overflow handling in the UTF-8 decoder. An unauthenticated, remote attacker can exploit this issue to cause an infinite loop in the decoder, leading to a denial of service condition. Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.0.52 / 8.5.31 / 9.0.8 or later. See Also: http://tomcat.apache.org/security-8.html#Fixed_in_Apache_Tomcat_8.0.52 http://tomcat.apache.org/security-8.html#Fixed_in_Apache_Tomcat_8.5.31 http://tomcat.apache.org/security-9.html#Fixed_in_Apache_Tomcat_9.0.8 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.31 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2017-12-07 Patch Publication Date: 2018-05-16 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 80 | Web Servers | Repeat Vulnerability | 2019-01-11 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 121124 |  | Apache Tomcat 8.0.x < 8.0.52 / 8.5.x < 8.5.31 / 9.0.x < 9.0.8 Denial of Service | Upgrade to Apache Tomcat version 8.0.52 / 8.5.31 / 9.0.8 or later. | High | Synopsis: The remote Apache Tomcat server is affected by a denial of service vulnerability. Description: According to its self-reported version number, the Apache Tomcat instance listening on the remote host is 8.0.x < 8.0.52, 8.5.x < 8.5.31 or 9.0.x < 9.0.8. It is, therefore, affected by the following vulnerability: - A denial of service (DoS) vulnerability exists in Tomcat due to improper overflow handling in the UTF-8 decoder. An unauthenticated, remote attacker can exploit this issue to cause an infinite loop in the decoder, leading to a denial of service condition. Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.0.52 / 8.5.31 / 9.0.8 or later. See Also: http://tomcat.apache.org/security-8.html#Fixed_in_Apache_Tomcat_8.0.52 http://tomcat.apache.org/security-8.html#Fixed_in_Apache_Tomcat_8.5.31 http://tomcat.apache.org/security-9.html#Fixed_in_Apache_Tomcat_9.0.8 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.31 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2017-12-07 Patch Publication Date: 2018-05-16 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 8080 | Web Servers | Repeat Vulnerability | 2019-01-11 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 126125 |  | Apache Tomcat 8.5.0 < 8.5.41 DoS | Upgrade to Apache Tomcat version 8.5.41 or later. | High | Synopsis: The remote Apache Tomcat server is affected by a denial of service vulnerability Description: The version of Tomcat installed on the remote host is prior to 8.5.41. It is, therefore, affected by a vulnerability as referenced in the fixed_in_apache_tomcat_8.5.41_security-8 advisory. - The fix for CVE-2019-0199 was incomplete and did not address HTTP/2 connection window exhaustion on write. By not sending WINDOW_UPDATE messages for the connection window (stream 0) clients were able to cause server-side threads to block eventually leading to thread exhaustion and a DoS. (CVE-2019-10072) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.41 or later. See Also: https://github.com/apache/tomcat/commit/0bcd69c https://github.com/apache/tomcat/commit/8d14c6f https://tomcat.apache.org/security-8.html#Fixed_in_Apache_Tomcat_8.5.41 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.41 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2019-04-10 Patch Publication Date: NULL Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 80 | Web Servers | Repeat Vulnerability | 2019-06-24 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 126125 |  | Apache Tomcat 8.5.0 < 8.5.41 DoS | Upgrade to Apache Tomcat version 8.5.41 or later. | High | Synopsis: The remote Apache Tomcat server is affected by a denial of service vulnerability Description: The version of Tomcat installed on the remote host is prior to 8.5.41. It is, therefore, affected by a vulnerability as referenced in the fixed_in_apache_tomcat_8.5.41_security-8 advisory. - The fix for CVE-2019-0199 was incomplete and did not address HTTP/2 connection window exhaustion on write. By not sending WINDOW_UPDATE messages for the connection window (stream 0) clients were able to cause server-side threads to block eventually leading to thread exhaustion and a DoS. (CVE-2019-10072) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.41 or later. See Also: https://github.com/apache/tomcat/commit/0bcd69c https://github.com/apache/tomcat/commit/8d14c6f https://tomcat.apache.org/security-8.html#Fixed_in_Apache_Tomcat_8.5.41 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.41 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2019-04-10 Patch Publication Date: NULL Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 8080 | Web Servers | Repeat Vulnerability | 2019-06-24 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 132413 |  | Apache Tomcat 8.5.0 < 8.5.49 Privilege Escalation | Upgrade to Apache Tomcat version 8.5.49 or later. | High | Synopsis: The remote Apache Tomcat server is affected by a vulnerability Description: The version of Tomcat installed on the remote host is prior to 8.5.49. It is, therefore, affected by a vulnerability as referenced in the fixed_in_apache_tomcat_8.5.49_security-8 advisory. - When Apache Tomcat is configured with the JMX Remote Lifecycle Listener, a local attacker without access to the Tomcat process or configuration files is able to manipulate the RMI registry to perform a man-in-the-middle attack to capture user names and passwords used to access the JMX interface. The attacker can then use these credentials to access the JMX interface and gain complete control over the Tomcat instance. (CVE-2019-12418) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.49 or later. See Also: https://github.com/apache/tomcat/commit/a91d7db http://www.nessus.org/u?ed6582f2 CVSS Base Score: 7 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:P/I:P/A:P Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.49 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2019-04-23 Patch Publication Date: NULL Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 80 | Web Servers | Repeat Vulnerability | 2019-12-27 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 132413 |  | Apache Tomcat 8.5.0 < 8.5.49 Privilege Escalation | Upgrade to Apache Tomcat version 8.5.49 or later. | High | Synopsis: The remote Apache Tomcat server is affected by a vulnerability Description: The version of Tomcat installed on the remote host is prior to 8.5.49. It is, therefore, affected by a vulnerability as referenced in the fixed_in_apache_tomcat_8.5.49_security-8 advisory. - When Apache Tomcat is configured with the JMX Remote Lifecycle Listener, a local attacker without access to the Tomcat process or configuration files is able to manipulate the RMI registry to perform a man-in-the-middle attack to capture user names and passwords used to access the JMX interface. The attacker can then use these credentials to access the JMX interface and gain complete control over the Tomcat instance. (CVE-2019-12418) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.49 or later. See Also: https://github.com/apache/tomcat/commit/a91d7db http://www.nessus.org/u?ed6582f2 CVSS Base Score: 7 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:P/I:P/A:P Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.49 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2019-04-23 Patch Publication Date: NULL Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 8080 | Web Servers | Repeat Vulnerability | 2019-12-27 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 132418 |  | Apache Tomcat 8.5.0 < 8.5.50 Privilege Escalation Vulnerability | Upgrade to Apache Tomcat version 8.5.50 or later. | High | Synopsis: The remote Apache Tomcat server is affected by a privilege escalation vulnerability Description: The version of Tomcat installed on the remote host is prior to 8.5.50. It is, therefore, affected by a privilege escalation vulnerability as referenced in the 'Fixed in Apache Tomcat 8.5.50' advisory. - When using FORM authentication there was a narrow window where an attacker could perform a session fixation attack. The window was considered too narrow for an exploit to be practical but, erring on the side of caution, this issue has been treated as a security vulnerability. (CVE-2019-17563) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.50 or later. See Also: https://github.com/apache/tomcat/commit/e19a202 http://www.nessus.org/u?e0b173ef CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:P/I:P/A:P Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.50 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2019-12-12 Patch Publication Date: 2019-12-12 Plugin Modification Date: 2022-12-05 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 80 | Web Servers | Repeat Vulnerability | 2019-12-27 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 132418 |  | Apache Tomcat 8.5.0 < 8.5.50 Privilege Escalation Vulnerability | Upgrade to Apache Tomcat version 8.5.50 or later. | High | Synopsis: The remote Apache Tomcat server is affected by a privilege escalation vulnerability Description: The version of Tomcat installed on the remote host is prior to 8.5.50. It is, therefore, affected by a privilege escalation vulnerability as referenced in the 'Fixed in Apache Tomcat 8.5.50' advisory. - When using FORM authentication there was a narrow window where an attacker could perform a session fixation attack. The window was considered too narrow for an exploit to be practical but, erring on the side of caution, this issue has been treated as a security vulnerability. (CVE-2019-17563) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.50 or later. See Also: https://github.com/apache/tomcat/commit/e19a202 http://www.nessus.org/u?e0b173ef CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:P/I:P/A:P Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.50 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2019-12-12 Patch Publication Date: 2019-12-12 Plugin Modification Date: 2022-12-05 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 8080 | Web Servers | Repeat Vulnerability | 2019-12-27 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 136807 |  | Apache Tomcat 8.5.x < 8.5.55 Remote Code Execution | Upgrade to Apache Tomcat version 8.5.55 or later. | High | Synopsis: The remote Apache Tomcat server is affected by a remote code execution vulnerability Description: The version of Tomcat installed on the remote host is prior to 8.5.x prior to 8.5.55. It is, therefore, affected by a remote code execution vulnerability as referenced in the fixed_in_apache_tomcat_8.5.55_security-8 advisory. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.55 or later. See Also: http://www.nessus.org/u?9502c510 CVSS Base Score: 7 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:P/I:P/A:P Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.55 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-05-11 Patch Publication Date: 2020-05-11 Plugin Modification Date: 2022-12-05 Exploitability Ease: Exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 80 | Web Servers | Repeat Vulnerability | 2020-05-22 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 136807 |  | Apache Tomcat 8.5.x < 8.5.55 Remote Code Execution | Upgrade to Apache Tomcat version 8.5.55 or later. | High | Synopsis: The remote Apache Tomcat server is affected by a remote code execution vulnerability Description: The version of Tomcat installed on the remote host is prior to 8.5.x prior to 8.5.55. It is, therefore, affected by a remote code execution vulnerability as referenced in the fixed_in_apache_tomcat_8.5.55_security-8 advisory. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.55 or later. See Also: http://www.nessus.org/u?9502c510 CVSS Base Score: 7 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:P/I:P/A:P Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.55 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-05-11 Patch Publication Date: 2020-05-11 Plugin Modification Date: 2022-12-05 Exploitability Ease: Exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 8080 | Web Servers | Repeat Vulnerability | 2020-05-22 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 138097 |  | Apache Tomcat 8.5.0 < 8.5.56 DoS | Upgrade to Apache Tomcat version 8.5.56 or later. | High | Synopsis: The remote Apache Tomcat server is affected by a denial of service vulnerability. Description: The version of Tomcat installed on the remote host is prior to 8.5.56. It is, therefore, affected by a denial of service vulnerability as referenced in the fixed_in_apache_tomcat_8.5.56_security-8 advisory. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.56 or later. See Also: http://www.nessus.org/u?9a501720 http://www.nessus.org/u?0ff2bf8c CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.56 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-06-26 Patch Publication Date: NULL Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 80 | Web Servers | Repeat Vulnerability | 2020-07-03 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 138097 |  | Apache Tomcat 8.5.0 < 8.5.56 DoS | Upgrade to Apache Tomcat version 8.5.56 or later. | High | Synopsis: The remote Apache Tomcat server is affected by a denial of service vulnerability. Description: The version of Tomcat installed on the remote host is prior to 8.5.56. It is, therefore, affected by a denial of service vulnerability as referenced in the fixed_in_apache_tomcat_8.5.56_security-8 advisory. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.56 or later. See Also: http://www.nessus.org/u?9a501720 http://www.nessus.org/u?0ff2bf8c CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.56 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-06-26 Patch Publication Date: NULL Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 8080 | Web Servers | Repeat Vulnerability | 2020-07-03 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 138574 |  | Apache Tomcat 8.5.0 < 8.5.57 Multiple Vulnerabilities | Upgrade to Apache Tomcat version 8.5.57 or later. | High | Synopsis: The remote Apache Tomcat server is affected by multiple vulnerabilities Description: The version of Tomcat installed on the remote host is 8.5.x prior to 8.5.57. It is, therefore, affected by multiple vulnerabilities as referenced in the Fixed in Apache Tomcat 8.5.57 security advisory. - The payload length in a WebSocket frame was not correctly validated. Invalid payload lengths could trigger an infinite loop. Multiple requests with invalid payload lengths could lead to a denial of service (DoS). (CVE-2020-13935) - An h2c direct connection did not release the HTTP/1.1 processor after the upgrade to HTTP/2. If a sufficient number of such requests were made, an OutOfMemoryException could occur leading to a denial of service (DoS). (CVE-2020-13934) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.57 or later. See Also: http://www.nessus.org/u?cd59de72 http://www.nessus.org/u?7358785a http://www.nessus.org/u?78f0e4ba CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.57 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-07-05 Patch Publication Date: 2020-07-05 Plugin Modification Date: 2022-12-05 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 80 | Web Servers | Repeat Vulnerability | 2020-07-17 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 138574 |  | Apache Tomcat 8.5.0 < 8.5.57 Multiple Vulnerabilities | Upgrade to Apache Tomcat version 8.5.57 or later. | High | Synopsis: The remote Apache Tomcat server is affected by multiple vulnerabilities Description: The version of Tomcat installed on the remote host is 8.5.x prior to 8.5.57. It is, therefore, affected by multiple vulnerabilities as referenced in the Fixed in Apache Tomcat 8.5.57 security advisory. - The payload length in a WebSocket frame was not correctly validated. Invalid payload lengths could trigger an infinite loop. Multiple requests with invalid payload lengths could lead to a denial of service (DoS). (CVE-2020-13935) - An h2c direct connection did not release the HTTP/1.1 processor after the upgrade to HTTP/2. If a sufficient number of such requests were made, an OutOfMemoryException could occur leading to a denial of service (DoS). (CVE-2020-13934) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.57 or later. See Also: http://www.nessus.org/u?cd59de72 http://www.nessus.org/u?7358785a http://www.nessus.org/u?78f0e4ba CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.57 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-07-05 Patch Publication Date: 2020-07-05 Plugin Modification Date: 2022-12-05 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 8080 | Web Servers | Repeat Vulnerability | 2020-07-17 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 144054 |  | Apache Tomcat 8.5.x < 8.5.60 Information Disclosure | Upgrade to Apache Tomcat version 8.5.60 or later. | High | Synopsis: The remote Apache Tomcat server is affected by an information disclosure vulnerability. Description: The version of Tomcat installed on the remote host is prior to 8.5.60. It is, therefore, affected by multiple vulnerabilities as referenced in the fixed_in_apache_tomcat_8.5.60_security-8 advisory. - When serving resources from a network location using the NTFS file system, Apache Tomcat versions 10.0.0-M1 to 10.0.0-M9, 9.0.0.M1 to 9.0.39, 8.5.0 to 8.5.59 and 7.0.0 to 7.0.106 were susceptible to JSP source code disclosure in some configurations. The root cause was the unexpected behaviour of the JRE API File.getCanonicalPath() which in turn was caused by the inconsistent behaviour of the Windows API (FindFirstFileW) in some circumstances. (CVE-2021-24122) - While investigating bug 64830 it was discovered that Apache Tomcat 10.0.0-M1 to 10.0.0-M9, 9.0.0-M1 to 9.0.39 and 8.5.0 to 8.5.59 could re-use an HTTP request header value from the previous stream received on an HTTP/2 connection for the request associated with the subsequent stream. While this would most likely lead to an error and the closure of the HTTP/2 connection, it is possible that information could leak between requests. (CVE-2020-17527) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.60 or later. See Also: http://www.nessus.org/u?05c4b1e2 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:N/A:N Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.60 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-11-17 Patch Publication Date: 2020-11-17 Plugin Modification Date: 2022-12-05 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 80 | Web Servers | Repeat Vulnerability | 2020-12-10 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 144054 |  | Apache Tomcat 8.5.x < 8.5.60 Information Disclosure | Upgrade to Apache Tomcat version 8.5.60 or later. | High | Synopsis: The remote Apache Tomcat server is affected by an information disclosure vulnerability. Description: The version of Tomcat installed on the remote host is prior to 8.5.60. It is, therefore, affected by multiple vulnerabilities as referenced in the fixed_in_apache_tomcat_8.5.60_security-8 advisory. - When serving resources from a network location using the NTFS file system, Apache Tomcat versions 10.0.0-M1 to 10.0.0-M9, 9.0.0.M1 to 9.0.39, 8.5.0 to 8.5.59 and 7.0.0 to 7.0.106 were susceptible to JSP source code disclosure in some configurations. The root cause was the unexpected behaviour of the JRE API File.getCanonicalPath() which in turn was caused by the inconsistent behaviour of the Windows API (FindFirstFileW) in some circumstances. (CVE-2021-24122) - While investigating bug 64830 it was discovered that Apache Tomcat 10.0.0-M1 to 10.0.0-M9, 9.0.0-M1 to 9.0.39 and 8.5.0 to 8.5.59 could re-use an HTTP request header value from the previous stream received on an HTTP/2 connection for the request associated with the subsequent stream. While this would most likely lead to an error and the closure of the HTTP/2 connection, it is possible that information could leak between requests. (CVE-2020-17527) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.60 or later. See Also: http://www.nessus.org/u?05c4b1e2 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:N/A:N Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.60 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-11-17 Patch Publication Date: 2020-11-17 Plugin Modification Date: 2022-12-05 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 8080 | Web Servers | Repeat Vulnerability | 2020-12-10 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 147019 |  | Apache Tomcat 8.5.0 < 8.5.63 Multiple Vulnerabilities | Upgrade to Apache Tomcat version 8.5.63 or later. | High | Synopsis: The remote Apache Tomcat server is affected by multiple vulnerabilities Description: The version of Tomcat installed on the remote host is prior to 8.5.63. It is, therefore, affected by multiple vulnerabilities as referenced in the vendor advisory. - When using Apache Tomcat versions 10.0.0-M1 to 10.0.0-M4, 9.0.0.M1 to 9.0.34, 8.5.0 to 8.5.54 and 7.0.0 to 7.0.103 if a) an attacker is able to control the contents and name of a file on the server; and b) the server is configured to use the PersistenceManager with a FileStore; and c) the PersistenceManager is configured with sessionAttributeValueClassNameFilter=null (the default unless a SecurityManager is used) or a sufficiently lax filter to allow the attacker provided object to be deserialized; and d) the attacker knows the relative file path from the storage location used by FileStore to the file the attacker has control over; then, using a specifically crafted request, the attacker will be able to trigger remote code execution via deserialization of the file under their control. Note that all of conditions a) to d) must be true for the attack to succeed. (CVE-2020-9484) - When responding to new h2c connection requests, Apache Tomcat versions 10.0.0-M1 to 10.0.0, 9.0.0.M1 to 9.0.41 and 8.5.0 to 8.5.61 could duplicate request headers and a limited amount of request body from one request to another meaning user A and user B could both see the results of user A's request. (CVE-2021-25122) - When using Apache Tomcat 10.0.0-M1 to 10.0.0, 9.0.0.M1 to 9.0.41, 8.5.0 to 8.5.61 or 7.0.0. to 7.0.107 with a configuration edge case that was highly unlikely to be used, the Tomcat instance was still vulnerable to CVE-2020-9494. Note that both the previously published prerequisites for CVE-2020-9484 and the previously published mitigations for CVE-2020-9484 also apply to this issue. (CVE-2021-25329) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.63 or later. See Also: http://www.nessus.org/u?b6278e74 http://www.nessus.org/u?0be223a3 http://www.nessus.org/u?15b6baad CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:N/A:N Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.63 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-02-02 Patch Publication Date: 2021-01-28 Plugin Modification Date: 2022-12-05 Exploitability Ease: Exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 80 | Web Servers | Repeat Vulnerability | 2021-03-04 | 2021-03-09 02:03 | 2023-05-07 06:58 | 26
|  |  | 147019 |  | Apache Tomcat 8.5.0 < 8.5.63 Multiple Vulnerabilities | Upgrade to Apache Tomcat version 8.5.63 or later. | High | Synopsis: The remote Apache Tomcat server is affected by multiple vulnerabilities Description: The version of Tomcat installed on the remote host is prior to 8.5.63. It is, therefore, affected by multiple vulnerabilities as referenced in the vendor advisory. - When using Apache Tomcat versions 10.0.0-M1 to 10.0.0-M4, 9.0.0.M1 to 9.0.34, 8.5.0 to 8.5.54 and 7.0.0 to 7.0.103 if a) an attacker is able to control the contents and name of a file on the server; and b) the server is configured to use the PersistenceManager with a FileStore; and c) the PersistenceManager is configured with sessionAttributeValueClassNameFilter=null (the default unless a SecurityManager is used) or a sufficiently lax filter to allow the attacker provided object to be deserialized; and d) the attacker knows the relative file path from the storage location used by FileStore to the file the attacker has control over; then, using a specifically crafted request, the attacker will be able to trigger remote code execution via deserialization of the file under their control. Note that all of conditions a) to d) must be true for the attack to succeed. (CVE-2020-9484) - When responding to new h2c connection requests, Apache Tomcat versions 10.0.0-M1 to 10.0.0, 9.0.0.M1 to 9.0.41 and 8.5.0 to 8.5.61 could duplicate request headers and a limited amount of request body from one request to another meaning user A and user B could both see the results of user A's request. (CVE-2021-25122) - When using Apache Tomcat 10.0.0-M1 to 10.0.0, 9.0.0.M1 to 9.0.41, 8.5.0 to 8.5.61 or 7.0.0. to 7.0.107 with a configuration edge case that was highly unlikely to be used, the Tomcat instance was still vulnerable to CVE-2020-9494. Note that both the previously published prerequisites for CVE-2020-9484 and the previously published mitigations for CVE-2020-9484 also apply to this issue. (CVE-2021-25329) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.63 or later. See Also: http://www.nessus.org/u?b6278e74 http://www.nessus.org/u?0be223a3 http://www.nessus.org/u?15b6baad CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:N/A:N Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.63 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-02-02 Patch Publication Date: 2021-01-28 Plugin Modification Date: 2022-12-05 Exploitability Ease: Exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 8080 | Web Servers | Repeat Vulnerability | 2021-03-04 | 2021-03-09 02:03 | 2023-05-07 06:58 | 26
|  |  | 161181 |  | Apache Tomcat 8.5.0 < 8.5.76 vulnerability | Upgrade to Apache Tomcat version 8.5.76 or later. | High | Synopsis: The remote Apache Tomcat server is affected by a vulnerability Description: The version of Tomcat installed on the remote host is prior to 8.5.76. It is, therefore, affected by a vulnerability as referenced in the fixed_in_apache_tomcat_8.5.76_security-8 advisory. - If a web application sends a WebSocket message concurrently with the WebSocket connection closing when running on Apache Tomcat 8.5.0 to 8.5.75 or Apache Tomcat 9.0.0.M1 to 9.0.20, it is possible that the application will continue to use the socket after it has been closed. The error handling triggered in this case could cause the a pooled object to be placed in the pool twice. This could result in subsequent connections using the same object concurrently which could result in data being returned to the wrong use and/or other errors. (CVE-2022-25762) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.76 or later. See Also: http://www.nessus.org/u?d8992e36 http://www.nessus.org/u?97cadf8c CVSS Base Score: 8.6 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.76 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2019-06-07 Patch Publication Date: 2022-02-28 Plugin Modification Date: 2022-05-25 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 80 | Web Servers | Repeat Vulnerability | 2022-05-13 | 2022-06-04 14:06 | 2023-05-07 06:58 | 11
|  |  | 161181 |  | Apache Tomcat 8.5.0 < 8.5.76 vulnerability | Upgrade to Apache Tomcat version 8.5.76 or later. | High | Synopsis: The remote Apache Tomcat server is affected by a vulnerability Description: The version of Tomcat installed on the remote host is prior to 8.5.76. It is, therefore, affected by a vulnerability as referenced in the fixed_in_apache_tomcat_8.5.76_security-8 advisory. - If a web application sends a WebSocket message concurrently with the WebSocket connection closing when running on Apache Tomcat 8.5.0 to 8.5.75 or Apache Tomcat 9.0.0.M1 to 9.0.20, it is possible that the application will continue to use the socket after it has been closed. The error handling triggered in this case could cause the a pooled object to be placed in the pool twice. This could result in subsequent connections using the same object concurrently which could result in data being returned to the wrong use and/or other errors. (CVE-2022-25762) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.76 or later. See Also: http://www.nessus.org/u?d8992e36 http://www.nessus.org/u?97cadf8c CVSS Base Score: 8.6 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.76 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2019-06-07 Patch Publication Date: 2022-02-28 Plugin Modification Date: 2022-05-25 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 8080 | Web Servers | Repeat Vulnerability | 2022-05-13 | 2022-06-04 14:06 | 2023-05-07 06:58 | 11
|  |  | 166807 |  | Apache Tomcat 8.5.x < 8.5.83 Request Smuggling Vulnerability | Upgrade to Apache Tomcat version 8.5.83 or later. | High | Synopsis: The remote Apache Tomcat server is affected by a request smuggling vulnerability Description: The version of Tomcat installed on the remote host is 8.5.x prior to 8.5.83. It is, therefore, affected by a request smuggling vulnerability as referenced in the fixed_in_apache_tomcat_8.5.83_security-8 advisory. If Tomcat was configured to ignore invalid HTTP headers via setting rejectIllegalHeader to false (the default), Tomcat did not reject a request containing an invalid Content-Length header making a request smuggling attack possible if Tomcat was located behind a reverse proxy that also failed to reject the request with the invalid header. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.83 or later. See Also: http://www.nessus.org/u?6b7e7ac8 http://www.nessus.org/u?7ff4a0f9 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:C/A:N Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.83 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-10-31 Patch Publication Date: 2022-10-31 Plugin Modification Date: 2023-01-05 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 80 | Web Servers | Repeat Vulnerability | 2022-11-02 | 2022-11-06 13:29 | 2023-05-07 06:58 | 6
|  |  | 166807 |  | Apache Tomcat 8.5.x < 8.5.83 Request Smuggling Vulnerability | Upgrade to Apache Tomcat version 8.5.83 or later. | High | Synopsis: The remote Apache Tomcat server is affected by a request smuggling vulnerability Description: The version of Tomcat installed on the remote host is 8.5.x prior to 8.5.83. It is, therefore, affected by a request smuggling vulnerability as referenced in the fixed_in_apache_tomcat_8.5.83_security-8 advisory. If Tomcat was configured to ignore invalid HTTP headers via setting rejectIllegalHeader to false (the default), Tomcat did not reject a request containing an invalid Content-Length header making a request smuggling attack possible if Tomcat was located behind a reverse proxy that also failed to reject the request with the invalid header. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.83 or later. See Also: http://www.nessus.org/u?6b7e7ac8 http://www.nessus.org/u?7ff4a0f9 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:C/A:N Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.83 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-10-31 Patch Publication Date: 2022-10-31 Plugin Modification Date: 2023-01-05 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 8080 | Web Servers | Repeat Vulnerability | 2022-11-02 | 2022-11-06 13:29 | 2023-05-07 06:58 | 6
|  |  | 171656 |  | Apache Tomcat 8.5.0 < 8.5.85 | Upgrade to Apache Tomcat version 8.5.85 or later. | High | Synopsis: The remote Apache Tomcat server is affected by a vulnerability Description: The version of Tomcat installed on the remote host is prior to 8.5.85. It is, therefore, affected by a vulnerability as referenced in the fixed_in_apache_tomcat_8.5.85_security-8 advisory. - Apache Commons FileUpload before 1.5 does not limit the number of request parts to be processed resulting in the possibility of an attacker triggering a DoS with a malicious upload or series of uploads. (CVE-2023-24998) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.85 or later. See Also: http://www.nessus.org/u?476daae8 http://www.nessus.org/u?3131512d CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:C Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.85 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-02-20 Patch Publication Date: NULL Plugin Modification Date: 2023-03-27 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 80 | Web Servers | Repeat Vulnerability | 2023-02-20 | 2023-03-05 07:58 | 2023-05-07 06:58 | 2
|  |  | 171656 |  | Apache Tomcat 8.5.0 < 8.5.85 | Upgrade to Apache Tomcat version 8.5.85 or later. | High | Synopsis: The remote Apache Tomcat server is affected by a vulnerability Description: The version of Tomcat installed on the remote host is prior to 8.5.85. It is, therefore, affected by a vulnerability as referenced in the fixed_in_apache_tomcat_8.5.85_security-8 advisory. - Apache Commons FileUpload before 1.5 does not limit the number of request parts to be processed resulting in the possibility of an attacker triggering a DoS with a malicious upload or series of uploads. (CVE-2023-24998) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.85 or later. See Also: http://www.nessus.org/u?476daae8 http://www.nessus.org/u?3131512d CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:C Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.85 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-02-20 Patch Publication Date: NULL Plugin Modification Date: 2023-03-27 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 8080 | Web Servers | Repeat Vulnerability | 2023-02-20 | 2023-03-05 07:58 | 2023-05-07 06:58 | 2
|  |  | 63155 |  | Microsoft Windows Unquoted Service Path Enumeration | Ensure that any services that contain a space in the path enclose the path in quotes. | High | Synopsis: The remote Windows host has at least one service installed that uses an unquoted service path. Description: The remote Windows host has at least one service installed that uses an unquoted service path, which contains at least one whitespace. A local attacker can gain elevated privileges by inserting an executable file in the path of the affected service. Note that this is a generic test that will flag any application affected by the described vulnerability. Solution: Ensure that any services that contain a space in the path enclose the path in quotes. See Also: http://www.nessus.org/u?84a4cc1c http://cwe.mitre.org/data/definitions/428.html https://www.commonexploits.com/unquoted-service-paths/ http://www.nessus.org/u?4aa6acbc CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:C/A:C Plugin Output: Nessus found the following services with an untrusted path : AxedaWatchDog : C:\Program Files (x86)\RVA\Gateway\xWatchDog.exe AxedaGateway : C:\Program Files (x86)\RVA\Gateway\xGate.exe Roche.DP.RSS.DataDistributionClient.StandAloneWinService : C:\Program Files (x86)\RVA\DDC\Roche.DP.RSS.DataDistributionClient.StandAloneWinService\Roche.DP.RSS.DataDistributionClient.StandAloneWinService.exe CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2012-09-15 Patch Publication Date: NULL Plugin Modification Date: 2020-06-12 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2012-12-05 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 65057 |  | Insecure Windows Service Permissions | Ensure the groups listed above do not have permissions to modify or write service executables. Additionally, ensure these groups do not have Full Control permission to any directories that contain service executables. | High | Synopsis: At least one improperly configured Windows service may have a privilege escalation vulnerability. Description: At least one Windows service executable with insecure permissions was detected on the remote host. Services configured to use an executable with weak permissions are vulnerable to privilege escalation attacks. An unprivileged user could modify or overwrite the executable with arbitrary code, which would be executed the next time the service is started. Depending on the user that the service runs as, this could result in privilege escalation. This plugin checks if any of the following groups have permissions to modify executable files that are started by Windows services : - Everyone - Users - Domain Users - Authenticated Users Solution: Ensure the groups listed above do not have permissions to modify or write service executables. Additionally, ensure these groups do not have Full Control permission to any directories that contain service executables. See Also: http://www.nessus.org/u?e4e766b2 CVSS Base Score: 8.4 CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : c:\intersystems\hsfoundation\bin\cservice.exe Used by services : Cache_c-_intersystems_hsfoundation File write allowed for groups : Authenticated Users (S-1-5-11) Full control of directory allowed for groups : Authenticated Users (S-1-5-11) Path : c:\intersystems\hsfoundation\httpd\bin\httpd.exe Used by services : HSFoundationhttpd File write allowed for groups : Authenticated Users (S-1-5-11) Full control of directory allowed for groups : Authenticated Users (S-1-5-11) Path : c:\program files\roche\roche it solutions\rits-reportsserver\bin\rits-reportsserver_service.exe Used by services : ReportsServer_RocheService File write allowed for groups : Authenticated Users (S-1-5-11) Full control of directory allowed for groups : Authenticated Users (S-1-5-11) CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2022-08-19 Exploitability Ease: NULL Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2013-03-06 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 81264 |  | MS15-011: Vulnerability in Group Policy Could Allow Remote Code Execution (3000483) | Microsoft has released a set of patches for Windows Vista, 2008, 7, 2008 R2, 8, 2012, 8.1, and 2012 R2. | High | Synopsis: The remote Windows host is affected by a remote code execution vulnerability. Description: The remote Windows host is affected by a remote code execution vulnerability due to how the Group Policy service manages policy data when a domain-joined system connects to a domain controller. An attacker, using a controlled network, can exploit this to gain complete control of the host. Note that Microsoft has no plans to release an update for Windows 2003 even though it is affected by this vulnerability. Solution: Microsoft has released a set of patches for Windows Vista, 2008, 7, 2008 R2, 8, 2012, 8.1, and 2012 R2. See Also: http://www.nessus.org/u?6533d970 http://www.nessus.org/u?b8230f41 CVSS Base Score: NULL CVSS Vector: CVSS2#AV:A/AC:L/Au:N/C:C/I:C/A:C Plugin Output: KB 3000483 or a related, subsequent update was successfully installed, but the GPO setting "Hardened UNC Paths" has not been enabled. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2015-02-10 Patch Publication Date: 2015-02-10 Plugin Modification Date: 2022-06-20 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2015-02-10 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 92320 |  | Apache 2.4.18 / 2.4.20 X.509 Certificate Authentication Bypass | Upgrade to Apache version 2.4.23 or later. Alternatively, as a temporary workaround, HTTP/2 can be disabled by changing the configuration by removing 'h2' and 'h2c' from the Protocols line(s) in the configuration file. | High | Synopsis: The remote web server is affected by an authentication bypass vulnerability. Description: According to its banner, the version of Apache running on the remote host is either 2.4.18 or 2.4.20. Additionally, HTTP/2 is enabled over TLS or SSL. It is, therefore, affected by the an authentication bypass vulnerability in the experimental module for the HTTP/2 protocol due to a failure to correctly validate X.509 certificates, allowing access to resources that otherwise would not be allowed. An unauthenticated, remote attacker can exploit this to disclose potentially sensitive information. Solution: Upgrade to Apache version 2.4.23 or later. Alternatively, as a temporary workaround, HTTP/2 can be disabled by changing the configuration by removing 'h2' and 'h2c' from the Protocols line(s) in the configuration file. See Also: https://archive.apache.org/dist/httpd/CHANGES_2.4.23 https://httpd.apache.org/security/vulnerabilities_24.html https://seclists.org/fulldisclosure/2016/Jul/11 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:P/A:N Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.23 CPE: cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2016-07-05 Patch Publication Date: 2016-07-05 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2016-07-15 | 2022-04-04 02:14 | 2023-05-09 07:17 | 13
|  |  | 96451 |  | Apache 2.4.x < 2.4.25 Multiple Vulnerabilities (httpoxy) | Upgrade to Apache version 2.4.25 or later. Note that the 'httpoxy' vulnerability can be mitigated by applying the workarounds or patches as referenced in the vendor advisory asf-httpoxy-response.txt. Furthermore, to mitigate the other vulnerabilities, ensure that the affected modules (mod_session_crypto, mod_auth_digest, and mod_http2) are not in use. | High | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: According to its banner, the version of Apache running on the remote host is 2.4.x prior to 2.4.25. It is, therefore, affected by the following vulnerabilities : - A flaw exists in the mod_session_crypto module due to encryption for data and cookies using the configured ciphers with possibly either CBC or ECB modes of operation (AES256-CBC by default). An unauthenticated, remote attacker can exploit this, via a padding oracle attack, to decrypt information without knowledge of the encryption key, resulting in the disclosure of potentially sensitive information. (CVE-2016-0736) - A denial of service vulnerability exists in the mod_auth_digest module during client entry allocation. An unauthenticated, remote attacker can exploit this, via specially crafted input, to exhaust shared memory resources, resulting in a server crash. (CVE-2016-2161) - The Apache HTTP Server is affected by a man-in-the-middle vulnerability known as 'httpoxy' due to a failure to properly resolve namespace conflicts in accordance with RFC 3875 section 4.1.18. The HTTP_PROXY environment variable is set based on untrusted user data in the 'Proxy' header of HTTP requests. The HTTP_PROXY environment variable is used by some web client libraries to specify a remote proxy server. An unauthenticated, remote attacker can exploit this, via a crafted 'Proxy' header in an HTTP request, to redirect an application's internal HTTP traffic to an arbitrary proxy server where it may be observed or manipulated. (CVE-2016-5387) - A denial of service vulnerability exists in the mod_http2 module due to improper handling of the LimitRequestFields directive. An unauthenticated, remote attacker can exploit this, via specially crafted CONTINUATION frames in an HTTP/2 request, to inject unlimited request headers into the server, resulting in the exhaustion of memory resources. (CVE-2016-8740) - A flaw exists due to improper handling of whitespace patterns in user-agent headers. An unauthenticated, remote attacker can exploit this, via a specially crafted user-agent header, to cause the program to incorrectly process sequences of requests, resulting in interpreting responses incorrectly, polluting the cache, or disclosing the content from one request to a second downstream user-agent. (CVE-2016-8743) - A CRLF injection allowing HTTP response splitting attacks for sites which use mod_userdir (CVE-2016-4975) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.25 or later. Note that the 'httpoxy' vulnerability can be mitigated by applying the workarounds or patches as referenced in the vendor advisory asf-httpoxy-response.txt. Furthermore, to mitigate the other vulnerabilities, ensure that the affected modules (mod_session_crypto, mod_auth_digest, and mod_http2) are not in use. See Also: https://httpd.apache.org/dev/dist/Announcement2.4.html http://httpd.apache.org/security/vulnerabilities_24.html https://github.com/apache/httpd/blob/2.4.x/CHANGES https://www.apache.org/security/asf-httpoxy-response.txt https://httpoxy.org CVSS Base Score: 8.1 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.25 CPE: cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2016-07-18 Patch Publication Date: 2016-12-20 Plugin Modification Date: 2022-04-11 Exploitability Ease: Exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2017-01-12 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 99588 |  | Oracle Java SE Multiple Vulnerabilities (April 2017 CPU) | Upgrade to Oracle JDK / JRE 6 Update 151 / 7 Update 141 / 8 Update 131 or later. If necessary, remove any affected versions. Note that an Extended Support contract with Oracle is needed to obtain JDK / JRE 6 Update 95 or later. | High | Synopsis: The remote Windows host contains a programming platform that is affected by multiple vulnerabilities. Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 6 Update 151, 7 Update 141, or 8 Update 131. It is, therefore, affected by multiple vulnerabilities : - An unspecified flaw exists in the Networking subcomponent that allows an unauthenticated, remote attacker to impact confidentiality and integrity. (CVE-2017-3509) - An unspecified flaw exists in the JCE subcomponent that allows a local attacker to gain elevated privileges. This vulnerability does not affect Java SE version 6. (CVE-2017-3511) - An unspecified flaw exists in the AWT subcomponent that allows an unauthenticated, remote attacker to execute arbitrary code. This vulnerability does not affect Java SE version 6. (CVE-2017-3512) - An unspecified flaw exists in the AWT subcomponent that allows an unauthenticated, remote attacker to execute arbitrary code. (CVE-2017-3514) - An unspecified flaw exists in the JAXP subcomponent that allows an unauthenticated, remote attacker to cause a denial of service condition. (CVE-2017-3526) - Multiple unspecified flaws exist in the Networking subcomponent that allow an unauthenticated, remote attacker to gain update, insert, or delete access to unauthorized data. (CVE-2017-3533, CVE-2017-3544) - An unspecified flaw exists in the Security subcomponent that allows an unauthenticated, remote attacker to gain update, insert, or delete access to unauthorized data. (CVE-2017-3539) Solution: Upgrade to Oracle JDK / JRE 6 Update 151 / 7 Update 141 / 8 Update 131 or later. If necessary, remove any affected versions. Note that an Extended Support contract with Oracle is needed to obtain JDK / JRE 6 Update 95 or later. See Also: http://www.nessus.org/u?02dc6498 http://www.nessus.org/u?ce35fa3a http://www.nessus.org/u?2fbcacca http://www.nessus.org/u?726f7054 http://www.nessus.org/u?eb4db3c7 CVSS Base Score: 8.3 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:P/I:P/A:P Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.6.0_151 / 1.7.0_141 / 1.8.0_131 CPE: cpe:/a:oracle:jre cpe:/a:oracle:jdk CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2017-04-18 Patch Publication Date: 2017-04-18 Plugin Modification Date: 2022-04-11 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2017-04-21 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 103838 |  | Apache 2.4.x < 2.4.28 HTTP Vulnerability (OptionsBleed) | Upgrade to Apache version 2.4.28 or later. | High | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: According to its banner, the version of Apache running on the remote host is 2.4.x prior to 2.4.28. It is, therefore, affected by an HTTP vulnerability related to the <Limit {method}> directive in an .htaccess file. Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.28 or later. See Also: https://archive.apache.org/dist/httpd/CHANGES_2.4.28 https://httpd.apache.org/security/vulnerabilities_24.html CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:N/A:N Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.28 CPE: cpe:/a:apache:http_server cpe:/a:apache:httpd CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2017-09-18 Patch Publication Date: 2017-10-05 Plugin Modification Date: 2022-04-11 Exploitability Ease: Exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2017-10-13 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 104896 |  | Security Updates for Internet Explorer (September 2017) | Microsoft has released security updates for the affected versions of Internet Explorer. | High | Synopsis: The Internet Explorer installation on the remote host is affected by multiple vulnerabilities. Description: The Internet Explorer installation on the remote host is missing security updates. It is, therefore, affected by multiple vulnerabilities : - An information disclosure vulnerability exists when affected Microsoft scripting engines do not properly handle objects in memory. The vulnerability could allow an attacker to detect specific files on the user's computer. (CVE-2017-8529) - A remote code execution vulnerability exists when Microsoft browsers improperly access objects in memory. The vulnerability could corrupt memory in such a way that an attacker could execute arbitrary code in the context of the current user. (CVE-2017-8750) - A spoofing vulnerability exists when Internet Explorer improperly handles specific HTML content. An attacker who successfully exploited this vulnerability could trick a user into believing that the user was visiting a legitimate website. The specially crafted website could either spoof content or serve as a pivot to chain an attack with other vulnerabilities in web services. (CVE-2017-8733) - A remote code execution vulnerability exists when Internet Explorer improperly accesses objects in memory. The vulnerability could corrupt memory in such a way that an attacker could execute arbitrary code in the context of the current user. (CVE-2017-8747, CVE-2017-8749) - A remote code execution vulnerability exists in the way that Microsoft browser JavaScript engines render content when handling objects in memory. The vulnerability could corrupt memory in such a way that an attacker could execute arbitrary code in the context of the current user. In a web-based attack scenario, an attacker could host a specially crafted website that is designed to exploit the vulnerability through Microsoft browsers and then convince a user to view the website. An attacker could also embed an ActiveX control marked &quot;safe for initialization&quot; in an application or Microsoft Office document that hosts the related rendering engine. The attacker could also take advantage of compromised websites, and websites that accept or host user-provided content or advertisements. These websites could contain specially crafted content that could exploit the vulnerability. An attacker who successfully exploited the vulnerability could gain the same user rights as the current user. (CVE-2017-8741, CVE-2017-8748) - An information disclosure vulnerability exists in Microsoft browsers due to improper parent domain verification in certain functionality. An attacker who successfully exploited the vulnerability could obtain specific information that is used in the parent domain. (CVE-2017-8736) Solution: Microsoft has released security updates for the affected versions of Internet Explorer. See Also: http://www.nessus.org/u?26b484bb http://www.nessus.org/u?085e4d22 http://www.nessus.org/u?35364720 http://www.nessus.org/u?1dbb18cc CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: The following registry key is required to enable the fix for CVE-2017-8529 and is missing. HKLM\SOFTWARE\Microsoft\Internet Explorer\Main\FeatureControl\FEATURE_ENABLE_PRINT_INFO_DISCLOSURE_FIX\iexplore.exe The following registry key is required to enable the fix for CVE-2017-8529 and is missing. HKLM\SOFTWARE\WOW6432Node\Microsoft\Internet Explorer\Main\FeatureControl\FEATURE_ENABLE_PRINT_INFO_DISCLOSURE_FIX\iexplore.exe CPE: cpe:/o:microsoft:windows cpe:/a:microsoft:ie CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2017-11-23 Patch Publication Date: 2017-11-23 Plugin Modification Date: 2020-07-17 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2017-11-30 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 106190 |  | Oracle Java SE Multiple Vulnerabilities (January 2018 CPU) | Upgrade to Oracle JDK / JRE 9 Update 4, 8 Update 161 / 7 Update 171 / 6 Update 181 or later. If necessary, remove any affected versions. Note that an Extended Support contract with Oracle is needed to obtain JDK / JRE 6 Update 95 or later. | High | Synopsis: The remote Windows host contains a programming platform that is affected by multiple vulnerabilities. Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 9 Update 4, 8 Update 161, 7 Update 171, or 6 Update 181. It is, therefore, affected by multiple vulnerabilities related to the following components : - AWT - Deployment - Hotspot - I18n - Installer - JCE - JGSS - JMX - JNDI - JavaFX - LDAP - Libraries - Serialization Solution: Upgrade to Oracle JDK / JRE 9 Update 4, 8 Update 161 / 7 Update 171 / 6 Update 181 or later. If necessary, remove any affected versions. Note that an Extended Support contract with Oracle is needed to obtain JDK / JRE 6 Update 95 or later. See Also: http://www.nessus.org/u?29ce2b01 http://www.nessus.org/u?793c3773 http://www.nessus.org/u?cc061f9a http://www.nessus.org/u?2fbcacca http://www.nessus.org/u?726f7054 CVSS Base Score: 8.3 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:P/A:P Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.6.0_181 / 1.7.0_171 / 1.8.0_161 / 1.9.0_4 CPE: cpe:/a:oracle:jre cpe:/a:oracle:jdk CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-01-16 Patch Publication Date: 2018-01-16 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2018-01-19 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 109202 |  | Oracle Java SE Multiple Vulnerabilities (April 2018 CPU) | Upgrade to Oracle JDK / JRE 10 Update 1, 8 Update 171 / 7 Update 181 / 6 Update 191 or later. If necessary, remove any affected versions. Note that an Extended Support contract with Oracle is needed to obtain JDK / JRE 6 Update 95 or later. | High | Synopsis: The remote Windows host contains a programming platform that is affected by multiple vulnerabilities. Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 10 Update 1, 8 Update 171, 7 Update 181, or 6 Update 191. It is, therefore, affected by multiple vulnerabilities related to the following components : - AWT - Concurrency - Hotspot - Install - JAXP - JMX - Libraries - RMI - Security - Serialization Solution: Upgrade to Oracle JDK / JRE 10 Update 1, 8 Update 171 / 7 Update 181 / 6 Update 191 or later. If necessary, remove any affected versions. Note that an Extended Support contract with Oracle is needed to obtain JDK / JRE 6 Update 95 or later. See Also: http://www.nessus.org/u?76507bf8 http://www.nessus.org/u?6f630e2b http://www.nessus.org/u?9bf6e180 http://www.nessus.org/u?2fbcacca http://www.nessus.org/u?726f7054 CVSS Base Score: 7.4 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:P/A:N Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.6.0_191 / 1.7.0_181 / 1.8.0_171 / 1.10.0_1 CPE: cpe:/a:oracle:jre cpe:/a:oracle:jdk CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-04-17 Patch Publication Date: 2018-04-17 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2018-04-20 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 111788 |  | Apache 2.4.x < 2.4.34 Multiple Vulnerabilities | Upgrade to Apache version 2.4.34 or later. | High | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: According to its banner, the version of Apache running on the remote host is 2.4.x prior to 2.4.34. It is, therefore, affected by the following vulnerabilities: - By specially crafting HTTP/2 requests, workers would be allocated 60 seconds longer than necessary, leading to worker exhaustion and a denial of service. (CVE-2018-1333) - By specially crafting HTTP requests, the mod_md challenge handler would dereference a NULL pointer and cause the child process to segfault. This could be used to DoS the server. (CVE-2018-8011) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.34 or later. See Also: https://archive.apache.org/dist/httpd/CHANGES_2.4.34 https://httpd.apache.org/security/vulnerabilities_24.html#2.4.34 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.34 CPE: cpe:/a:apache:http_server cpe:/a:apache:httpd CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-07-15 Patch Publication Date: 2018-07-15 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2018-08-16 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 121355 |  | Apache 2.4.x < 2.4.38 Multiple Vulnerabilities | Upgrade to Apache version 2.4.38 or later. | High | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: According to its banner, the version of Apache running on the remote host is 2.4.x prior to 2.4.38. It is, therefore, affected by multiple vulnerabilities: - A denial of service (DoS) vulnerability exists in HTTP/2 steam handling. An unauthenticated, remote attacker can exploit this issue, via sending request bodies in a slow loris way to plain resources, to occupy a server thread. (CVE-2018-17189) - A vulnerability exists in mod_sesion_cookie, as it does not properly check the expiry time of cookies. (CVE-2018-17199) - A denial of service (DoS) vulnerability exists in mod_ssl when used with OpenSSL 1.1.1 due to an interaction in changes to handling of renegotiation attempts. An unauthenticated, remote attacker can exploit this issue to cause mod_ssl to stop responding. (CVE-2019-0190) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.38 or later. See Also: https://archive.apache.org/dist/httpd/CHANGES_2.4.38 https://httpd.apache.org/security/vulnerabilities_24.html#2.4.38 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:P/A:N Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.38 CPE: cpe:/a:apache:http_server cpe:/a:apache:httpd CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2019-01-22 Patch Publication Date: 2019-01-22 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2019-01-24 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 125780 |  | McAfee Agent < 5.6.1 Information Disclosure Vulnerability (SB10271) | Upgrade to McAfee Agent version 5.6.1, or later. | High | Synopsis: A security management agent installed on the remote host is affected by an information disclosure vulnerability. Description: The version of McAfee Agent, formerly McAfee ePolicy Orchestrator (ePO) Agent, installed on the remote host is 5.0.x, 5.5.x, or 5.6.x < 5.6.1. It is, therefore, affected by an information disclosure vulnerability. An unauthenticated, remote attacker can exploit this, via the remote logger, to disclose potentially sensitive information. Solution: Upgrade to McAfee Agent version 5.6.1, or later. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10271 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:N/A:N Plugin Output: Path : C:\Program Files (x86)\McAfee\Common Framework\ Installed version : 5.0.6.220 Fixed version : 5.6.1 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2019-02-26 Patch Publication Date: 2019-02-26 Plugin Modification Date: 2022-04-27 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Misc. | Repeat Vulnerability | 2019-06-07 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 132992 |  | Oracle Java SE 1.7.0_251 / 1.8.0_241 / 1.11.0_6 / 1.13.0_2 Multiple Vulnerabilities (Jan 2020 CPU) | Upgrade to Oracle JDK / JRE 13 Update 2 , 11 Update 6, 8 Update 241 / 7 Update 251 or later. If necessary, remove any affected versions. | High | Synopsis: The remote Windows host contains a programming platform that is affected by multiple vulnerabilities. Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 7 Update 251, 8 Update 241, 11 Update 6, or 13 Update 2. It is, therefore, affected by multiple vulnerabilities: - Oracle Java SE and Java SE Embedded are prone to a severe division by zero, over 'Multiple' protocol. This issue affects the 'SQLite' component.(CVE-2019-16168) - Oracle Java SE and Java SE Embedded are prone to format string vulnerability, leading to a read uninitialized stack data over 'Multiple' protocol. This issue affects the 'libxst' component. (CVE-2019-13117, CVE-2019-13118) - Oracle Java SE and Java SE Embedded are prone to a remote security vulnerability. An unauthenticated remote attacker can exploit this over 'Kerberos' protocol. This issue affects the 'Security' component. (CVE-2020-2601, CVE-2020-2590) - Oracle Java SE/Java SE Embedded are prone to a remote security vulnerability. An unauthenticated remote attacker can exploit this overmultiple protocols. This issue affects the 'Serialization' component. (CVE-2020-2604, CVE-2020-2583) - Oracle Java SE/Java SE Embedded are prone to a remote security vulnerability. Tn unauthenticated remote attacker can exploit this over multiple protocols. This issue affects the 'Networking' component. (CVE-2020-2593, CVE-2020-2659) - Oracle Java SE are prone to a remote security vulnerability. An unauthenticated remote attacker can exploit this over multiple protocols. This issue affects the 'Libraries' component. (CVE-2020-2654) - Oracle Java SE are prone to a multiple security vulnerability. An unauthenticated remote attacker can exploit this over multiple protocols. This issue affects the 'JavaFX' component. (CVE-2020-2585) - Oracle Java SE are prone to a multiple security vulnerability. An unauthenticate remote attacker can exploit this over 'HTTPS' protocols. This issue affects the 'JSSE' component. (CVE-2020-2655) Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Oracle JDK / JRE 13 Update 2 , 11 Update 6, 8 Update 241 / 7 Update 251 or later. If necessary, remove any affected versions. See Also: http://www.nessus.org/u?d22a1e87 CVSS Base Score: 8.1 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:P/A:P Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.7.0_251 / 1.8.0_241 / 1.11.0_6 / 1.13.0_2 CPE: cpe:/a:oracle:jre cpe:/a:oracle:jdk CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-01-14 Patch Publication Date: 2020-01-14 Plugin Modification Date: 2022-10-21 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2020-01-16 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 135592 |  | Oracle Java SE 1.7.0_261 / 1.8.0_251 / 1.11.0_7 / 1.14.0_1 Multiple Vulnerabilities (Apr 2020 CPU) | Upgrade to Oracle JDK / JRE 14 Update 1 , 11 Update 7, 8 Update 251 , 7 Update 261 or later. If necessary, remove any affected versions. | High | Synopsis: The remote host is affected by multiple vulnerabilities Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 7 Update 261, 8 Update 251, 11 Update 7, or 14 Update 1. It is, therefore, affected by multiple vulnerabilities related to the following components : - Oracle Java SE and Java SE Embedded are prone to a buffer overflow attack, over 'Multiple' protocol. This issue affects the 'JavaFX (libxslt)' component. Successful attacks of this vulnerability allow unauthenticated attacker with network access to takeover of Java SE. (CVE-2019-18197) - Oracle Java SE and Java SE Embedded are prone to partial denial of service (partial DOS) vulnerability. An unauthenticated remote attacker can exploit this over 'Multiple' protocol. This issue affects the 'Scripting' component. (CVE-2020-2754, CVE-2020-2755) - Oracle Java SE and Java SE Embedded are prone to partial denial of service (partial DOS) vulnerability. An unauthenticated remote attacker can exploit this over 'Multiple' protocol. This issue affects the 'Serialization' component. (CVE-2020-2756, CVE-2020-2757) - Oracle Java SE prone to unauthorized read access vulnerability. An unauthenticated remote attacker can exploit this over 'Multiple' protocol can result in unauthorized read access to a subset of Java SE accessible data. This issue affects the 'Advanced Management Console' component. (CVE-2020-2764) - Oracle Java SE and Java SE Embedded are prone to unauthorized write/read access vulnerability. An unauthenticated remote attacker over 'HTTPS' can read, update, insert or delete access to some of Java SE accessible data. This issue affects the 'JSSE' component. (CVE-2020-2767) - Oracle Java SE and Java SE Embedded are prone to partial denial of service (partial DOS) vulnerability. An unauthenticated remote attacker can exploit this over 'Multiple' protocol. This issue affects the 'Scripting' component. (CVE-2020-2773) It is also affected by other vulnerabilities; please see vendor advisories for more information. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Oracle JDK / JRE 14 Update 1 , 11 Update 7, 8 Update 251 , 7 Update 261 or later. If necessary, remove any affected versions. See Also: https://www.oracle.com/a/tech/docs/cpuapr2020cvrf.xml https://www.oracle.com/security-alerts/cpuapr2020.html CVSS Base Score: 8.3 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:P/A:N Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.7.0_261 / 1.8.0_251 / 1.11.0_7 / 1.14.0_1 CPE: cpe:/a:oracle:jre cpe:/a:oracle:jdk CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-04-14 Patch Publication Date: 2020-04-14 Plugin Modification Date: 2022-05-13 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2020-04-16 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 136668 |  | McAfee VirusScan Enterprise < 8.8 Patch 14 Hotfix 116778 Privilege Escalation Vulnerability (SB10316) | Upgrade to McAfee VirusScan Enterprise version 8.8 Patch 14 Hotfix 116778. | High | Synopsis: The antivirus application installed on the remote Windows host is affected by a privilege escalation vulnerability. Description: The version of McAfee VirusScan Enterprise (VSE) installed on the remote Windows host is prior to 8.8 Patch 14 Hotfix 116778. It is, therefore, affected by a privilege escalation vulnerability. Solution: Upgrade to McAfee VirusScan Enterprise version 8.8 Patch 14 Hotfix 116778. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10316 CVSS Base Score: 8.4 CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:N/I:P/A:P Plugin Output: Installed version : 8.8.0.2024 Fixed version : 8.8.0.2190 (8.8 Patch 14 with Hotfix 116778) CPE: cpe:/a:mcafee:virusscan_enterprise CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-04-23 Patch Publication Date: 2020-04-23 Plugin Modification Date: 2020-07-02 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2020-05-15 | 2022-09-04 23:24 | 2023-05-09 07:17 | 8
|  |  | 137648 |  | McAfee VirusScan Enterprise < 8.8 Patch 15 Multiple Vulnerabilities (SB10302) | Upgrade to McAfee VirusScan Enterprise version 8.8 Patch 15. | High | Synopsis: The antivirus application installed on the remote Windows host is affected by multiple vulnerabilities. Description: The version of McAfee VirusScan Enterprise (VSE) installed on the remote Windows host is prior to 8.8 Patch 15. It is, therefore, affected by multiple vulnerabilites: - Privilege Escalation vulnerability in Microsoft Windows client (McTray.exe) in McAfee VirusScan Enterprise (VSE) 8.8 prior to Patch 14 may allow local users to interact with the On-Access Scan Messages - Threat Alert Window with elevated privileges via running McAfee Tray with elevated privileges (CVE-2019-3585 ). - Privilege Escalation vulnerability in Microsoft Windows client (McTray.exe) in McAfee VirusScan Enterprise (VSE) 8.8 prior to Patch 14 may allow unauthorized users to interact with the On-Access Scan Messages - Threat Alert Window when the Windows Login Screen is locked (CVE-2019-3588). - Privilege Escalation vulnerability during daily DAT updates when using McAfee Virus Scan Enterprise (VSE) prior to 8.8 Patch 15 allows local users to cause the deletion and creation of files they would not normally have permission to through altering the target of symbolic links. This is timing dependent (CVE-2020-7280). Solution: Upgrade to McAfee VirusScan Enterprise version 8.8 Patch 15. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10302 CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Installed version : 8.8.0.2024 Fixed version : 8.8.0.2232 CPE: cpe:/a:mcafee:virusscan_enterprise CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-06-09 Patch Publication Date: 2020-06-09 Plugin Modification Date: 2020-12-18 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2020-06-19 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 138522 |  | Oracle Java SE 1.7.0_271 / 1.8.0_261 / 1.11.0_8 / 1.14.0_2 Multiple Vulnerabilities (Jul 2020 CPU) | Apply the appropriate patch according to the July 2020 Oracle Critical Patch Update advisory. | High | Synopsis: The remote host is affected by multiple vulnerabilities Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 7 Update 271, 8 Update 261, 11 Update 8, or 14 Update 2. It is, therefore, affected by multiple vulnerabilities related to the following components as referenced in the July 2020 CPU advisory: - Vulnerability in the Java SE product of Oracle Java SE (component: JavaFX). The supported version that is affected is Java SE: 8u251. Difficult to exploit vulnerability allows unauthenticated attacker with network access via multiple protocols to compromise Java SE. Successful attacks require human interaction from a person other than the attacker and while the vulnerability is in Java SE, attacks may significantly impact additional products. Successful attacks of this vulnerability can result in takeover of Java SE. Note: This vulnerability applies to Java deployments, typically in clients running sandboxed Java Web Start applications or sandboxed Java applets, that load and run untrusted code (e.g., code that comes from the internet) and rely on the Java sandbox for security. This vulnerability does not apply to Java deployments, typically in servers, that load and run only trusted code (e.g., code installed by an administrator). (CVE-2020-14664) - Vulnerability in the Java SE, Java SE Embedded product of Oracle Java SE (component: Libraries). Supported versions that are affected are Java SE: 7u261, 8u251, 11.0.7 and 14.0.1; Java SE Embedded: 8u251. Difficult to exploit vulnerability allows unauthenticated attacker with network access via multiple protocols to compromise Java SE, Java SE Embedded. Successful attacks require human interaction from a person other than the attacker and while the vulnerability is in Java SE, Java SE Embedded, attacks may significantly impact additional products. Successful attacks of this vulnerability can result in takeover of Java SE, Java SE Embedded. Note: This vulnerability applies to Java deployments, typically in clients running sandboxed Java Web Start applications or sandboxed Java applets, that load and run untrusted code (e.g., code that comes from the internet) and rely on the Java sandbox for security. This vulnerability does not apply to Java deployments, typically in servers, that load and run only trusted code (e.g., code installed by an administrator). (CVE-2020-14583) - Vulnerability in the Java SE, Java SE Embedded product of Oracle Java SE (component: 2D). Supported versions that are affected are Java SE: 7u261, 8u251, 11.0.7 and 14.0.1; Java SE Embedded: 8u251. Easily exploitable vulnerability allows unauthenticated attacker with network access via multiple protocols to compromise Java SE, Java SE Embedded. Successful attacks require human interaction from a person other than the attacker and while the vulnerability is in Java SE, Java SE Embedded, attacks may significantly impact additional products. Successful attacks of this vulnerability can result in unauthorized creation, deletion or modification access to critical data or all Java SE, Java SE Embedded accessible data. Note: This vulnerability applies to Java deployments, typically in clients running sandboxed Java Web Start applications or sandboxed Java applets, that load and run untrusted code (e.g., code that comes from the internet) and rely on the Java sandbox for security. This vulnerability does not apply to Java deployments, typically in servers, that load and run only trusted code (e.g., code installed by an administrator). (CVE-2020-14593) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Apply the appropriate patch according to the July 2020 Oracle Critical Patch Update advisory. See Also: https://www.oracle.com/a/tech Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.7.0_271 / 1.8.0_261 / 1.11.0_8 / 1.14.0_2 | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2020-07-16 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 150860 |  | McAfee Agent 5.x prior to 5.7.3 Multiple Vulnerabilities (SB10362) | Upgrade to McAfee Agent version 5.7.3 or later. | High | Synopsis: A security management agent installed on the remote host is affected by a missing authorization vulnerability. Description: The version of McAfee Agent, formerly McAfee ePolicy Orchestrator (ePO) Agent, installed on the remote host is 5.x prior to 5.7.3. It is, therefore, affected by the following vulnerabilities: - An improper privilege management vulnerability exists due to a local user having the ability to edit the event log. An authenticated, attacker can exploit this, via by editing the agents event log prior to the events being sent to the ePO server. (CVE-2021-31839) - An unsigned DLL preloading attack vulnerability exists. An authenticated, local attacker can exploit this and load malicious libraries of functionality on the remote host due to an insufficient validation of resources loaded by the application at runtime. (CVE-2021-31840) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to McAfee Agent version 5.7.3 or later. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10362 CVSS Base Score: 7.3 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files (x86)\McAfee\Common Framework\ Installed version : 5.0.6.220 Fixed version : 5.7.3 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-06-08 Patch Publication Date: 2021-06-08 Plugin Modification Date: 2021-09-27 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2021-06-18 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 152020 |  | Oracle Java SE 1.7.0_311 / 1.8.0_301 / 1.11.0_12 / 1.16.0_2 Multiple Vulnerabilities (July 2021 CPU) | Apply the appropriate patch according to the July 2021 Oracle Critical Patch Update advisory. | High | Synopsis: The remote host is affected by multiple vulnerabilities. Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 7 Update 301, 8 Update 291, 11 Update 11, or 16 Update 1. It is, therefore, affected by multiple vulnerabilities as referenced in the July 2021 CPU advisory: - Vulnerability in the Java SE, Oracle GraalVM Enterprise Edition product of Oracle Java SE (component: Networking). Supported versions that are affected are Java SE: 7u301, 8u291, 11.0.11, 16.0.1; Oracle GraalVM Enterprise Edition: 20.3.2 and 21.1.0. Difficult to exploit vulnerability allows unauthenticated attacker with network access via multiple protocols to compromise Java SE, Oracle GraalVM Enterprise Edition. Successful attacks require human interaction from a person other than the attacker. Successful attacks of this vulnerability can result in unauthorized read access to a subset of Java SE, Oracle GraalVM Enterprise Edition accessible data. This vulnerability applies to Java deployments, typically in clients running sandboxed Java Web Start applications or sandboxed Java applets, that load and run untrusted code (e.g., code that comes from the internet) and rely on the Java sandbox for security. This vulnerability does not apply to Java deployments, typically in servers, that load and run only trusted code (e.g., code installed by an administrator). (CVE-2021-2341) - Vulnerability in the Java SE, Oracle GraalVM Enterprise Edition product of Oracle Java SE (component: Library). Supported versions that are affected are Java SE: 7u301, 8u291, 11.0.11, 16.0.1; Oracle GraalVM Enterprise Edition: 20.3.2 and 21.1.0. Easily exploitable vulnerability allows unauthenticated attacker with network access via multiple protocols to compromise Java SE, Oracle GraalVM Enterprise Edition. Successful attacks require human interaction from a person other than the attacker. Successful attacks of this vulnerability can result in unauthorized update, insert or delete access to some of Java SE, Oracle GraalVM Enterprise Edition accessible data. This vulnerability applies to Java deployments, typically in clients running sandboxed Java Web Start applications or sandboxed Java applets, that load and run untrusted code (e.g., code that comes from the internet) and rely on the Java sandbox for security. This vulnerability does not apply to Java deployments, typically in servers, that load and run only trusted code (e.g., code installed by an administrator). (CVE-2021-2369) - Vulnerability in the Java SE, Oracle GraalVM Enterprise Edition product of Oracle Java SE (component: Hotspot). Supported versions that are affected are Java SE: 8u291, 11.0.11, 16.0.1; Oracle GraalVM Enterprise Edition: 20.3.2 and 21.1.0. Difficult to exploit vulnerability allows unauthenticated attacker with network access via multiple protocols to compromise Java SE, Oracle GraalVM Enterprise Edition. Successful attacks require human interaction from a person other than the attacker. Successful attacks of this vulnerability can result in takeover of Java SE, Oracle GraalVM Enterprise Edition. This vulnerability applies to Java deployments, typically in clients running sandboxed Java Web Start applications or sandboxed Java applets, that load and run untrusted code (e.g., code that comes from the internet) and rely on the Java sandbox for security. This vulnerability does not apply to Java deployments, typically in servers, that load and run only trusted code (e.g., code installed by an administrator). (CVE-2021-2388) - Vulnerability in the Java SE product of Oracle Java SE (component: JNDI). The supported version that is affected is Java SE: 7u301. Difficult to exploit vulnerability allows unauthenticated attacker with network access via multiple protocols to compromise Java SE. Successful attacks of this Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.7.0_311 / 1.8.0_301 / 1.11.0_12 / 1.16.0_2 | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2021-07-23 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 152042 |  | McAfee Agent 5.5.x prior to 5.5.4 / 5.6.x prior to 5.6.4 DLL Search Order Hijacking (SB10320) | Upgrade to McAfee Agent version 5.5.4, 5.6.4 or later. | High | Synopsis: A security management agent installed on the remote host is affected by a DLL search order hijacking vulnerability. Description: The version of McAfee Agent, formerly McAfee ePolicy Orchestrator (ePO) Agent, installed on the remote host is 5.5.x prior to 5.5.4 / 5.6.x prior to 5.6.4. It is, therefore, affected by a DLL search order hijacking vulnerability. it allows attackers with local access to execute arbitrary code via execution from a compromised folder. Solution: Upgrade to McAfee Agent version 5.5.4, 5.6.4 or later. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10320 CVSS Base Score: 7.3 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files (x86)\McAfee\Common Framework\ Installed version : 5.0.6.220 Fixed version : 5.5.4 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-06-10 Patch Publication Date: 2020-06-10 Plugin Modification Date: 2021-09-27 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2021-07-23 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 153617 |  | McAfee Agent < 5.7.4 Multiple Vulnerabilities (SB10369) | Upgrade to McAfee Agent version 5.7.4 or later. | High | Synopsis: A security management agent installed on the remote host is affected by multiple vulnerabilities. Description: The version of McAfee Agent, formerly McAfee ePolicy Orchestrator (ePO) Agent, installed on the remote host is prior to 5.7.4. It is, therefore, affected by the following vulnerabilities: - Improper access control vulnerability in the repair process for McAfee Agent for Windows prior to 5.7.4 could allow a local attacker to perform a DLL preloading attack using unsigned DLLs. This would result in elevation of privileges and the ability to execute arbitrary code as the system user, through not correctly protecting a temporary directory used in the repair process and not checking the DLL signature. (CVE-2021-31847) - A DLL sideloading vulnerability in McAfee Agent for Windows prior to 5.7.4 could allow a local user to perform a DLL sideloading attack with an unsigned DLL with a specific name and in a specific location. This would result in the user gaining elevated permissions and the ability to execute arbitrary code as the system user, through not checking the DLL signature. (CVE-2021-31841) - Improper privilege management vulnerability in maconfig for McAfee Agent for Windows prior to 5.7.4 allows a local user to gain access to sensitive information. The utility was able to be run from any location on the file system and by a low privileged user. (CVE-2021-31836) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to McAfee Agent version 5.7.4 or later. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10369 CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files (x86)\McAfee\Common Framework\ Installed version : 5.0.6.220 Fixed version : 5.7.4 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-09-21 Patch Publication Date: 2021-09-21 Plugin Modification Date: 2022-02-03 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2021-09-24 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 154344 |  | Oracle Java SE 1.7.0_321 / 1.8.0_311 / 1.11.0_13 / 1.17.0_1 Multiple Vulnerabilities (October 2021 CPU) | Apply the appropriate patch according to the October 2021 Oracle Critical Patch Update advisory. | High | Synopsis: The remote host is affected by multiple vulnerabilities. Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 7 Update 321, 8 Update 311, 11 Update 13, or 17 Update 1. It is, therefore, affected by multiple vulnerabilities as referenced in the October 2021 CPU advisory: - Vulnerability in the Java SE product of Oracle Java SE (component: JavaFX (libxml)). The supported version that is affected is Java SE: 8u301. Easily exploitable vulnerability allows unauthenticated attacker with network access via multiple protocols to compromise Java SE. Successful attacks of this vulnerability can result in unauthorized ability to cause a hang or frequently repeatable crash (complete DOS) of Java SE as well as unauthorized update, insert or delete access to some of Java SE accessible data and unauthorized read access to a subset of Java SE accessible data. This vulnerability applies to Java deployments, typically in clients running sandboxed Java Web Start applications or sandboxed Java applets, that load and run untrusted code (e.g., code that comes from the internet) and rely on the Java sandbox for security. This vulnerability does not apply to Java deployments, typically in servers, that load and run only trusted code (e.g., code installed by an administrator). (CVE-2021-3517) - Vulnerability in the Java SE product of Oracle Java SE (component: Deployment). The supported version that is affected is Java SE: 8u301. Difficult to exploit vulnerability allows unauthenticated attacker with network access via multiple protocols to compromise Java SE. Successful attacks require human interaction from a person other than the attacker. Successful attacks of this vulnerability can result in takeover of Java SE. This vulnerability applies to Java deployments, typically in clients running sandboxed Java Web Start applications or sandboxed Java applets, that load and run untrusted code (e.g., code that comes from the internet) and rely on the Java sandbox for security. This vulnerability does not apply to Java deployments, typically in servers, that load and run only trusted code (e.g., code installed by an administrator). (CVE-2021-35560) - Vulnerability in the Java SE, Oracle GraalVM Enterprise Edition product of Oracle Java SE (component: Libraries). Supported versions that are affected are Java SE: 8u301, 11.0.12, 17; Oracle GraalVM Enterprise Edition: 20.3.3 and 21.2.0. Easily exploitable vulnerability allows low privileged attacker with network access via Kerberos to compromise Java SE, Oracle GraalVM Enterprise Edition. Successful attacks require human interaction from a person other than the attacker and while the vulnerability is in Java SE, Oracle GraalVM Enterprise Edition, attacks may significantly impact additional products. Successful attacks of this vulnerability can result in unauthorized access to critical data or complete access to all Java SE, Oracle GraalVM Enterprise Edition accessible data. (CVE-2021-35567) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Apply the appropriate patch according to the October 2021 Oracle Critical Patch Update advisory. See Also: https://www.oracle.com/a/tech/docs/cpuoct2021cvrf.xml https://www.oracle.com/security-alerts/cpuoct2021.html#AppendixJAVA CVSS Base Score: 8.6 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.7.0_321 / 1.8.0_311 / 1.11.0_13 / 1.17.0_1 CPE: cpe:/a:oracle:jre cpe:/a:oracle:jdk CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-05-19 Patch Publication Date: 2021-10-19 Plugin Mod | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2021-10-22 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 157126 |  | McAfee Agent < 5.7.5 Multiple Vulnerabilities (SB10378) | Upgrade to McAfee Agent version 5.7.5 or later. | High | Synopsis: A security management agent installed on the remote host is affected by multiple vulnerabilities. Description: The version of McAfee Agent, formerly McAfee ePolicy Orchestrator (ePO) Agent, installed on the remote host is prior to 5.7.5. It is, therefore, affected by the following vulnerabilities: - A command Injection Vulnerability in McAfee Agent (MA) for Windows prior to 5.7.5 allows local users to inject arbitrary shell code into the file cleanup.exe. The malicious clean.exe file is placed into the relevant folder and executed by running the McAfee Agent deployment feature located in the System Tree. An attacker may exploit the vulnerability to obtain a reverse shell which can lead to privilege escalation to obtain root privileges. (CVE-2021-31854) - A privilege escalation vulnerability in the McAfee Agent prior to 5.7.5. McAfee Agent uses openssl.cnf during the build process to specify the OPENSSLDIR variable as a subdirectory within the installation directory. A low privilege user could have created subdirectories and executed arbitrary code with SYSTEM privileges by creating the appropriate pathway to the specifically created malicious openssl.cnf file. (CVE-2022-0166) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to McAfee Agent version 5.7.5 or later. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10378 CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files (x86)\McAfee\Common Framework\ Installed version : 5.0.6.220 Fixed version : 5.7.5 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-01-18 Patch Publication Date: 2022-01-18 Plugin Modification Date: 2022-05-06 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2022-01-26 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 160054 |  | McAfee Agent < 5.7.6 Multiple Vulnerabilities (SB10382) | Upgrade to McAfee Agent version 5.7.6 or later. | High | Synopsis: A security management agent installed on the remote host is affected by multiple vulnerabilities. Description: The version of McAfee Agent, formerly McAfee ePolicy Orchestrator (ePO) Agent, installed on the remote host is prior to 5.7.6. It is, therefore, affected by the following vulnerabilities: - A local privilege escalation vulnerability in MA for Windows prior to 5.7.6 allows a local low privileged user to gain system privileges through running the repair functionality. Temporary file actions were performed on the local user's %TEMP% directory with System privileges through manipulation of symbolic links. (CVE-2022-1256) - Insecure storage of sensitive information vulnerability in MA for Linux, macOS, and Windows prior to 5.7.6 allows a local user to gain access to sensitive information through storage in ma.db. The sensitive information has been moved to encrypted database files. (CVE-2022-1257) - A blind SQL injection vulnerability in the ePolicy Orchestrator (ePO) extension of MA prior to 5.7.6 can be exploited by an authenticated administrator on ePO to perform arbitrary SQL queries in the back-end database, potentially leading to command execution on the server. (CVE-2022-1258) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to McAfee Agent version 5.7.6 or later. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10382 CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files (x86)\McAfee\Common Framework\ Installed version : 5.0.6.220 Fixed version : 5.7.6 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-04-14 Patch Publication Date: 2022-04-14 Plugin Modification Date: 2022-08-09 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2022-04-21 | 2022-05-02 07:06 | 2023-05-09 07:17 | 12
|  |  | 163587 |  | McAfee Agent < 5.7.7 DLL Hijacking (SB10385) | Upgrade to McAfee Agent version 5.7.7 or later. | High | Synopsis: A security management agent installed on the remote host is affected by a DLL hijacking vulnerability. Description: A DLL hijacking vulnerability in the MA Smart Installer for Windows prior to 5.7.7, which allows local users to execute arbitrary code and obtain higher privileges via careful placement of a malicious DLL into the folder from where the Smart installer is being executed. Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to McAfee Agent version 5.7.7 or later. See Also: https://kcm.trellix.com/corporate/index?page=content&id=SB10385 CVSS Base Score: 7.3 CVSS Vector: CVSS2#AV:L/AC:L/Au:S/C:C/I:C/A:C Plugin Output: Path : C:\Program Files (x86)\McAfee\Common Framework\ Installed version : 5.0.6.220 Fixed version : 5.7.7 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-07-26 Patch Publication Date: 2022-07-26 Plugin Modification Date: 2023-01-05 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2022-07-29 | 2022-08-07 23:57 | 2023-05-09 07:17 | 9
|  |  | 166030 |  | KB5018476: Windows Server 2012 R2 Security Update (October 2022) | Apply Security Update 5018476 or Cumulative Update 5018474 | High | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5018476. It is, therefore, affected by multiple vulnerabilities - Server Service Remote Protocol Elevation of Privilege Vulnerability (CVE-2022-38045) - Microsoft ODBC Driver Remote Code Execution Vulnerability (CVE-2022-38040) - Microsoft WDAC OLE DB provider for SQL Server Remote Code Execution Vulnerability (CVE-2022-37982, CVE-2022-38031) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5018476 or Cumulative Update 5018474 See Also: https://support.microsoft.com/en-us/help/5018474 https://support.microsoft.com/en-us/help/5018476 https://support.microsoft.com/help/5018474 https://support.microsoft.com/help/5018476 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5018476 - 5018474 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20625 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-10-11 Patch Publication Date: 2022-10-11 Plugin Modification Date: 2023-02-09 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-10-11 | 2022-11-08 03:32 | 2023-05-09 07:17 | 6
|  |  | 167109 |  | KB5020010: Windows Server 2012 R2 Security Update (November 2022) | Apply Security Update 5020010 or Cumulative Update 5020023 | High | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5020010. It is, therefore, affected by multiple vulnerabilities - AMD: CVE-2022-23824 IBPB and Return Address Predictor Interactions (CVE-2022-23824) - Windows Kerberos RC4-HMAC Elevation of Privilege Vulnerability (CVE-2022-37966) - Windows Kerberos Elevation of Privilege Vulnerability (CVE-2022-37967) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5020010 or Cumulative Update 5020023 See Also: https://support.microsoft.com/help/5020010 https://support.microsoft.com/help/5020023 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5020023 - 5020010 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20666 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-11-08 Patch Publication Date: 2022-11-08 Plugin Modification Date: 2023-01-03 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-11-08 | 2022-12-05 20:44 | 2023-05-09 07:17 | 5
|  |  | 168687 |  | KB5021296: Windows Server 2012 R2 Security Update (December 2022) | Apply Security Update 5021296 or Cumulative Update 5021294 | High | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5021296. It is, therefore, affected by multiple vulnerabilities - PowerShell Remote Code Execution Vulnerability (CVE-2022-41076) - A remote code execution vulnerability. An attacker can exploit this to bypass authentication and execute unauthorized arbitrary commands. (CVE-2022-44670, CVE-2022-44676) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5021296 or Cumulative Update 5021294 See Also: https://support.microsoft.com/en-us/help/5021294 https://support.microsoft.com/en-us/help/5021296 https://support.microsoft.com/help/5021294 https://support.microsoft.com/help/5021296 CVSS Base Score: 8.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5021296 - 5021294 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20716 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-12-13 Patch Publication Date: 2022-12-13 Plugin Modification Date: 2023-01-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-12-13 | 2023-01-09 04:57 | 2023-05-09 07:17 | 4
|  |  | 172137 |  | Wireshark 4.0.x < 4.0.4 A Vulnerability | Upgrade to Wireshark version 4.0.4 or later. | High | Synopsis: An application installed on the remote Windows host is affected by a vulnerability. Description: The version of Wireshark installed on the remote Windows host is prior to 4.0.4. It is, therefore, affected by a vulnerability as referenced in the wireshark-4.0.4 advisory. - The ISO 15765 and ISO 10681 dissectors could crash. It may be possible to make Wireshark crash by injecting a malformed packet onto the wire or by convincing someone to read a malformed packet trace file. (CVE-2023-1161) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Wireshark version 4.0.4 or later. See Also: https://www.wireshark.org/docs/relnotes/wireshark-4.0.4.html https://gitlab.com/wireshark/wireshark/-/issues/18839 https://www.wireshark.org/security/wnpa-sec-2023-08 CVSS Base Score: 7.1 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Path : C:\Program Files\Wireshark Installed version : 4.0.3 Fixed version : 4.0.4 CPE: cpe:/a:wireshark:wireshark CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-03-02 Patch Publication Date: 2023-03-02 Plugin Modification Date: 2023-03-22 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Compass Router - Prod | LAUREL BRIDGE SOFTWARE INC |  |  | Region X | VISN X | VAMC | VHAXXXAPPCROUTP |  |  |  |  | VHAXXXAPPCROUTP.vxx.med.va.gov |  | Microsoft Windows Server 2019 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2023-03-06 | 2023-04-02 20:36 | 2023-05-08 00:45 | 1
|  |  | 172137 |  | Wireshark 4.0.x < 4.0.4 A Vulnerability | Upgrade to Wireshark version 4.0.4 or later. | High | Synopsis: An application installed on the remote Windows host is affected by a vulnerability. Description: The version of Wireshark installed on the remote Windows host is prior to 4.0.4. It is, therefore, affected by a vulnerability as referenced in the wireshark-4.0.4 advisory. - The ISO 15765 and ISO 10681 dissectors could crash. It may be possible to make Wireshark crash by injecting a malformed packet onto the wire or by convincing someone to read a malformed packet trace file. (CVE-2023-1161) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Wireshark version 4.0.4 or later. See Also: https://www.wireshark.org/docs/relnotes/wireshark-4.0.4.html https://gitlab.com/wireshark/wireshark/-/issues/18839 https://www.wireshark.org/security/wnpa-sec-2023-08 CVSS Base Score: 7.1 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Path : C:\Program Files\Wireshark Installed version : 4.0.3 Fixed version : 4.0.4 CPE: cpe:/a:wireshark:wireshark CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-03-02 Patch Publication Date: 2023-03-02 Plugin Modification Date: 2023-03-22 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Compass Router - Hot Spare | LAUREL BRIDGE SOFTWARE INC |  |  | Region X | VISN X | VAMC | VHAXXXAPPCROUTS |  |  |  |  | VHAXXXAPPCROUTS.vxx.med.va.gov |  | Microsoft Windows Server 2019 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2023-03-06 | 2023-04-02 20:36 | 2023-05-08 00:46 | 1
|  |  | 172137 |  | Wireshark 4.0.x < 4.0.4 A Vulnerability | Upgrade to Wireshark version 4.0.4 or later. | High | Synopsis: An application installed on the remote Windows host is affected by a vulnerability. Description: The version of Wireshark installed on the remote Windows host is prior to 4.0.4. It is, therefore, affected by a vulnerability as referenced in the wireshark-4.0.4 advisory. - The ISO 15765 and ISO 10681 dissectors could crash. It may be possible to make Wireshark crash by injecting a malformed packet onto the wire or by convincing someone to read a malformed packet trace file. (CVE-2023-1161) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Wireshark version 4.0.4 or later. See Also: https://www.wireshark.org/docs/relnotes/wireshark-4.0.4.html https://gitlab.com/wireshark/wireshark/-/issues/18839 https://www.wireshark.org/security/wnpa-sec-2023-08 CVSS Base Score: 7.1 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Path : C:\Program Files\Wireshark Installed version : 4.0.3 Fixed version : 4.0.4 CPE: cpe:/a:wireshark:wireshark CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-03-02 Patch Publication Date: 2023-03-02 Plugin Modification Date: 2023-03-22 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Compass Router - Test | LAUREL BRIDGE SOFTWARE INC |  |  | Region X | VISN X | VAMC | VHAXXXAPPCROUTT |  |  |  |  | VHAXXXAPPCROUTT.vxx.med.va.gov |  | Microsoft Windows Server 2019 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2023-03-06 | 2023-04-02 20:36 | 2023-05-08 00:44 | 1
|  |  | 63155 |  | Microsoft Windows Unquoted Service Path Enumeration | Ensure that any services that contain a space in the path enclose the path in quotes. | High | Synopsis: The remote Windows host has at least one service installed that uses an unquoted service path. Description: The remote Windows host has at least one service installed that uses an unquoted service path, which contains at least one whitespace. A local attacker can gain elevated privileges by inserting an executable file in the path of the affected service. Note that this is a generic test that will flag any application affected by the described vulnerability. Solution: Ensure that any services that contain a space in the path enclose the path in quotes. See Also: http://www.nessus.org/u?84a4cc1c http://cwe.mitre.org/data/definitions/428.html https://www.commonexploits.com/unquoted-service-paths/ http://www.nessus.org/u?4aa6acbc CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:C/A:C Plugin Output: Nessus found the following services with an untrusted path : AxedaWatchDog : C:\Program Files (x86)\RVA\Gateway\xWatchDog.exe AxedaGateway : C:\Program Files (x86)\RVA\Gateway\xGate.exe CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2012-09-15 Patch Publication Date: NULL Plugin Modification Date: 2020-06-12 Exploitability Ease: Exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2012-12-05 | 2022-02-06 21:50 | 2023-05-09 07:10 | 15
|  |  | 65057 |  | Insecure Windows Service Permissions | Ensure the groups listed above do not have permissions to modify or write service executables. Additionally, ensure these groups do not have Full Control permission to any directories that contain service executables. | High | Synopsis: At least one improperly configured Windows service may have a privilege escalation vulnerability. Description: At least one Windows service executable with insecure permissions was detected on the remote host. Services configured to use an executable with weak permissions are vulnerable to privilege escalation attacks. An unprivileged user could modify or overwrite the executable with arbitrary code, which would be executed the next time the service is started. Depending on the user that the service runs as, this could result in privilege escalation. This plugin checks if any of the following groups have permissions to modify executable files that are started by Windows services : - Everyone - Users - Domain Users - Authenticated Users Solution: Ensure the groups listed above do not have permissions to modify or write service executables. Additionally, ensure these groups do not have Full Control permission to any directories that contain service executables. See Also: http://www.nessus.org/u?e4e766b2 CVSS Base Score: 8.4 CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : d:\intersystems\cache\bin\cservice.exe Used by services : Cache_d-_intersystems_cache File write allowed for groups : Authenticated Users (S-1-5-11), Everyone (S-1-1-0) Full control of directory allowed for groups : Authenticated Users (S-1-5-11), Everyone (S-1-1-0) Path : d:\intersystems\cache\httpd\bin\httpd.exe Used by services : CACHEhttpd File write allowed for groups : Authenticated Users (S-1-5-11), Everyone (S-1-1-0) Full control of directory allowed for groups : Authenticated Users (S-1-5-11), Everyone (S-1-1-0) CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2022-08-19 Exploitability Ease: NULL Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2013-03-06 | 2022-02-06 21:50 | 2023-05-09 07:10 | 15
|  |  | 81264 |  | MS15-011: Vulnerability in Group Policy Could Allow Remote Code Execution (3000483) | Microsoft has released a set of patches for Windows Vista, 2008, 7, 2008 R2, 8, 2012, 8.1, and 2012 R2. | High | Synopsis: The remote Windows host is affected by a remote code execution vulnerability. Description: The remote Windows host is affected by a remote code execution vulnerability due to how the Group Policy service manages policy data when a domain-joined system connects to a domain controller. An attacker, using a controlled network, can exploit this to gain complete control of the host. Note that Microsoft has no plans to release an update for Windows 2003 even though it is affected by this vulnerability. Solution: Microsoft has released a set of patches for Windows Vista, 2008, 7, 2008 R2, 8, 2012, 8.1, and 2012 R2. See Also: http://www.nessus.org/u?6533d970 http://www.nessus.org/u?b8230f41 CVSS Base Score: NULL CVSS Vector: CVSS2#AV:A/AC:L/Au:N/C:C/I:C/A:C Plugin Output: KB 3000483 or a related, subsequent update was successfully installed, but the GPO setting "Hardened UNC Paths" has not been enabled. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2015-02-10 Patch Publication Date: 2015-02-10 Plugin Modification Date: 2022-06-20 Exploitability Ease: Exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2015-02-10 | 2019-02-01 21:38 | 2023-05-09 07:10 | 51
|  |  | 104896 |  | Security Updates for Internet Explorer (September 2017) | Microsoft has released security updates for the affected versions of Internet Explorer. | High | Synopsis: The Internet Explorer installation on the remote host is affected by multiple vulnerabilities. Description: The Internet Explorer installation on the remote host is missing security updates. It is, therefore, affected by multiple vulnerabilities : - An information disclosure vulnerability exists when affected Microsoft scripting engines do not properly handle objects in memory. The vulnerability could allow an attacker to detect specific files on the user's computer. (CVE-2017-8529) - A remote code execution vulnerability exists when Microsoft browsers improperly access objects in memory. The vulnerability could corrupt memory in such a way that an attacker could execute arbitrary code in the context of the current user. (CVE-2017-8750) - A spoofing vulnerability exists when Internet Explorer improperly handles specific HTML content. An attacker who successfully exploited this vulnerability could trick a user into believing that the user was visiting a legitimate website. The specially crafted website could either spoof content or serve as a pivot to chain an attack with other vulnerabilities in web services. (CVE-2017-8733) - A remote code execution vulnerability exists when Internet Explorer improperly accesses objects in memory. The vulnerability could corrupt memory in such a way that an attacker could execute arbitrary code in the context of the current user. (CVE-2017-8747, CVE-2017-8749) - A remote code execution vulnerability exists in the way that Microsoft browser JavaScript engines render content when handling objects in memory. The vulnerability could corrupt memory in such a way that an attacker could execute arbitrary code in the context of the current user. In a web-based attack scenario, an attacker could host a specially crafted website that is designed to exploit the vulnerability through Microsoft browsers and then convince a user to view the website. An attacker could also embed an ActiveX control marked &quot;safe for initialization&quot; in an application or Microsoft Office document that hosts the related rendering engine. The attacker could also take advantage of compromised websites, and websites that accept or host user-provided content or advertisements. These websites could contain specially crafted content that could exploit the vulnerability. An attacker who successfully exploited the vulnerability could gain the same user rights as the current user. (CVE-2017-8741, CVE-2017-8748) - An information disclosure vulnerability exists in Microsoft browsers due to improper parent domain verification in certain functionality. An attacker who successfully exploited the vulnerability could obtain specific information that is used in the parent domain. (CVE-2017-8736) Solution: Microsoft has released security updates for the affected versions of Internet Explorer. See Also: http://www.nessus.org/u?26b484bb http://www.nessus.org/u?085e4d22 http://www.nessus.org/u?35364720 http://www.nessus.org/u?1dbb18cc CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: The following registry key is required to enable the fix for CVE-2017-8529 and is missing. HKLM\SOFTWARE\Microsoft\Internet Explorer\Main\FeatureControl\FEATURE_ENABLE_PRINT_INFO_DISCLOSURE_FIX\iexplore.exe The following registry key is required to enable the fix for CVE-2017-8529 and is missing. HKLM\SOFTWARE\WOW6432Node\Microsoft\Internet Explorer\Main\FeatureControl\FEATURE_ENABLE_PRINT_INFO_DISCLOSURE_FIX\iexplore.exe CPE: cpe:/o:microsoft:windows cpe:/a:microsoft:ie CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2017-11-23 Patch Publication Date: 2017-11-23 Plugin Modification Date: 2020-07-17 Exploitability Ease: Exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2017-11-30 | 2022-02-06 21:50 | 2023-05-09 07:10 | 15
|  |  | 125780 |  | McAfee Agent < 5.6.1 Information Disclosure Vulnerability (SB10271) | Upgrade to McAfee Agent version 5.6.1, or later. | High | Synopsis: A security management agent installed on the remote host is affected by an information disclosure vulnerability. Description: The version of McAfee Agent, formerly McAfee ePolicy Orchestrator (ePO) Agent, installed on the remote host is 5.0.x, 5.5.x, or 5.6.x < 5.6.1. It is, therefore, affected by an information disclosure vulnerability. An unauthenticated, remote attacker can exploit this, via the remote logger, to disclose potentially sensitive information. Solution: Upgrade to McAfee Agent version 5.6.1, or later. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10271 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:N/A:N Plugin Output: Path : C:\Program Files\McAfee\Agent\ Installed version : 5.5.1.342 Fixed version : 5.6.1 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2019-02-26 Patch Publication Date: 2019-02-26 Plugin Modification Date: 2022-04-27 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Misc. | Repeat Vulnerability | 2019-06-07 | 2022-02-06 21:50 | 2023-05-09 07:10 | 15
|  |  | 150860 |  | McAfee Agent 5.x prior to 5.7.3 Multiple Vulnerabilities (SB10362) | Upgrade to McAfee Agent version 5.7.3 or later. | High | Synopsis: A security management agent installed on the remote host is affected by a missing authorization vulnerability. Description: The version of McAfee Agent, formerly McAfee ePolicy Orchestrator (ePO) Agent, installed on the remote host is 5.x prior to 5.7.3. It is, therefore, affected by the following vulnerabilities: - An improper privilege management vulnerability exists due to a local user having the ability to edit the event log. An authenticated, attacker can exploit this, via by editing the agents event log prior to the events being sent to the ePO server. (CVE-2021-31839) - An unsigned DLL preloading attack vulnerability exists. An authenticated, local attacker can exploit this and load malicious libraries of functionality on the remote host due to an insufficient validation of resources loaded by the application at runtime. (CVE-2021-31840) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to McAfee Agent version 5.7.3 or later. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10362 CVSS Base Score: 7.3 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files\McAfee\Agent\ Installed version : 5.5.1.342 Fixed version : 5.7.3 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-06-08 Patch Publication Date: 2021-06-08 Plugin Modification Date: 2021-09-27 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2021-06-18 | 2022-02-06 21:50 | 2023-05-09 07:10 | 15
|  |  | 152041 |  | McAfee Host Intrusion Prevention Services < 8.0.0.5257 DLL Search Order Hijacking (SB10320) | Upgrade to McAfee Host Intrusion Prevention Services 8.0 Patch 15 (8.0.0.5257) or later. | High | Synopsis: A security application installed on the remote host is affected by a DLL search order hijacking vulnerability. Description: The version of McAfee Host Intrusion Prevention Services (HIPS) is prior to 8.0.0.5257. It is, therefore, affected by a DLL search order hijacking vulnerability. It allows attackers with local access to execute arbitrary code via execution from a compromised folder. Solution: Upgrade to McAfee Host Intrusion Prevention Services 8.0 Patch 15 (8.0.0.5257) or later. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10320 https://kc.mcafee.com/corporate/index?page=content&id=KB56057 CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files (x86)\McAfee\Host Intrusion Prevention\ Installed version : 8.0.0.4480 Fixed version : 8.0.0.5257 CPE: cpe:/a:mcafee:host_intrusion_prevention CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-06-10 Patch Publication Date: 2020-06-10 Plugin Modification Date: 2021-07-27 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2021-07-23 | 2022-02-06 21:50 | 2023-05-09 07:10 | 15
|  |  | 152042 |  | McAfee Agent 5.5.x prior to 5.5.4 / 5.6.x prior to 5.6.4 DLL Search Order Hijacking (SB10320) | Upgrade to McAfee Agent version 5.5.4, 5.6.4 or later. | High | Synopsis: A security management agent installed on the remote host is affected by a DLL search order hijacking vulnerability. Description: The version of McAfee Agent, formerly McAfee ePolicy Orchestrator (ePO) Agent, installed on the remote host is 5.5.x prior to 5.5.4 / 5.6.x prior to 5.6.4. It is, therefore, affected by a DLL search order hijacking vulnerability. it allows attackers with local access to execute arbitrary code via execution from a compromised folder. Solution: Upgrade to McAfee Agent version 5.5.4, 5.6.4 or later. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10320 CVSS Base Score: 7.3 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files\McAfee\Agent\ Installed version : 5.5.1.342 Fixed version : 5.5.4 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-06-10 Patch Publication Date: 2020-06-10 Plugin Modification Date: 2021-09-27 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2021-07-23 | 2022-02-06 21:50 | 2023-05-09 07:10 | 15
|  |  | 153617 |  | McAfee Agent < 5.7.4 Multiple Vulnerabilities (SB10369) | Upgrade to McAfee Agent version 5.7.4 or later. | High | Synopsis: A security management agent installed on the remote host is affected by multiple vulnerabilities. Description: The version of McAfee Agent, formerly McAfee ePolicy Orchestrator (ePO) Agent, installed on the remote host is prior to 5.7.4. It is, therefore, affected by the following vulnerabilities: - Improper access control vulnerability in the repair process for McAfee Agent for Windows prior to 5.7.4 could allow a local attacker to perform a DLL preloading attack using unsigned DLLs. This would result in elevation of privileges and the ability to execute arbitrary code as the system user, through not correctly protecting a temporary directory used in the repair process and not checking the DLL signature. (CVE-2021-31847) - A DLL sideloading vulnerability in McAfee Agent for Windows prior to 5.7.4 could allow a local user to perform a DLL sideloading attack with an unsigned DLL with a specific name and in a specific location. This would result in the user gaining elevated permissions and the ability to execute arbitrary code as the system user, through not checking the DLL signature. (CVE-2021-31841) - Improper privilege management vulnerability in maconfig for McAfee Agent for Windows prior to 5.7.4 allows a local user to gain access to sensitive information. The utility was able to be run from any location on the file system and by a low privileged user. (CVE-2021-31836) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to McAfee Agent version 5.7.4 or later. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10369 CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files\McAfee\Agent\ Installed version : 5.5.1.342 Fixed version : 5.7.4 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-09-21 Patch Publication Date: 2021-09-21 Plugin Modification Date: 2022-02-03 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2021-09-24 | 2022-02-06 21:50 | 2023-05-09 07:10 | 15
|  |  | 157126 |  | McAfee Agent < 5.7.5 Multiple Vulnerabilities (SB10378) | Upgrade to McAfee Agent version 5.7.5 or later. | High | Synopsis: A security management agent installed on the remote host is affected by multiple vulnerabilities. Description: The version of McAfee Agent, formerly McAfee ePolicy Orchestrator (ePO) Agent, installed on the remote host is prior to 5.7.5. It is, therefore, affected by the following vulnerabilities: - A command Injection Vulnerability in McAfee Agent (MA) for Windows prior to 5.7.5 allows local users to inject arbitrary shell code into the file cleanup.exe. The malicious clean.exe file is placed into the relevant folder and executed by running the McAfee Agent deployment feature located in the System Tree. An attacker may exploit the vulnerability to obtain a reverse shell which can lead to privilege escalation to obtain root privileges. (CVE-2021-31854) - A privilege escalation vulnerability in the McAfee Agent prior to 5.7.5. McAfee Agent uses openssl.cnf during the build process to specify the OPENSSLDIR variable as a subdirectory within the installation directory. A low privilege user could have created subdirectories and executed arbitrary code with SYSTEM privileges by creating the appropriate pathway to the specifically created malicious openssl.cnf file. (CVE-2022-0166) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to McAfee Agent version 5.7.5 or later. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10378 CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files\McAfee\Agent\ Installed version : 5.5.1.342 Fixed version : 5.7.5 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-01-18 Patch Publication Date: 2022-01-18 Plugin Modification Date: 2022-05-06 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2022-01-26 | 2022-02-06 21:50 | 2023-05-09 07:10 | 15
|  |  | 160054 |  | McAfee Agent < 5.7.6 Multiple Vulnerabilities (SB10382) | Upgrade to McAfee Agent version 5.7.6 or later. | High | Synopsis: A security management agent installed on the remote host is affected by multiple vulnerabilities. Description: The version of McAfee Agent, formerly McAfee ePolicy Orchestrator (ePO) Agent, installed on the remote host is prior to 5.7.6. It is, therefore, affected by the following vulnerabilities: - A local privilege escalation vulnerability in MA for Windows prior to 5.7.6 allows a local low privileged user to gain system privileges through running the repair functionality. Temporary file actions were performed on the local user's %TEMP% directory with System privileges through manipulation of symbolic links. (CVE-2022-1256) - Insecure storage of sensitive information vulnerability in MA for Linux, macOS, and Windows prior to 5.7.6 allows a local user to gain access to sensitive information through storage in ma.db. The sensitive information has been moved to encrypted database files. (CVE-2022-1257) - A blind SQL injection vulnerability in the ePolicy Orchestrator (ePO) extension of MA prior to 5.7.6 can be exploited by an authenticated administrator on ePO to perform arbitrary SQL queries in the back-end database, potentially leading to command execution on the server. (CVE-2022-1258) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to McAfee Agent version 5.7.6 or later. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10382 CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files\McAfee\Agent\ Installed version : 5.5.1.342 Fixed version : 5.7.6 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-04-14 Patch Publication Date: 2022-04-14 Plugin Modification Date: 2022-08-09 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2022-04-21 | 2022-05-02 07:06 | 2023-05-09 07:10 | 12
|  |  | 163587 |  | McAfee Agent < 5.7.7 DLL Hijacking (SB10385) | Upgrade to McAfee Agent version 5.7.7 or later. | High | Synopsis: A security management agent installed on the remote host is affected by a DLL hijacking vulnerability. Description: A DLL hijacking vulnerability in the MA Smart Installer for Windows prior to 5.7.7, which allows local users to execute arbitrary code and obtain higher privileges via careful placement of a malicious DLL into the folder from where the Smart installer is being executed. Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to McAfee Agent version 5.7.7 or later. See Also: https://kcm.trellix.com/corporate/index?page=content&id=SB10385 CVSS Base Score: 7.3 CVSS Vector: CVSS2#AV:L/AC:L/Au:S/C:C/I:C/A:C Plugin Output: Path : C:\Program Files\McAfee\Agent\ Installed version : 5.5.1.342 Fixed version : 5.7.7 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-07-26 Patch Publication Date: 2022-07-26 Plugin Modification Date: 2023-01-05 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2022-07-29 | 2022-08-07 23:57 | 2023-05-09 07:10 | 9
|  |  | 166030 |  | KB5018476: Windows Server 2012 R2 Security Update (October 2022) | Apply Security Update 5018476 or Cumulative Update 5018474 | High | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5018476. It is, therefore, affected by multiple vulnerabilities - Server Service Remote Protocol Elevation of Privilege Vulnerability (CVE-2022-38045) - Microsoft ODBC Driver Remote Code Execution Vulnerability (CVE-2022-38040) - Microsoft WDAC OLE DB provider for SQL Server Remote Code Execution Vulnerability (CVE-2022-37982, CVE-2022-38031) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5018476 or Cumulative Update 5018474 See Also: https://support.microsoft.com/en-us/help/5018474 https://support.microsoft.com/en-us/help/5018476 https://support.microsoft.com/help/5018474 https://support.microsoft.com/help/5018476 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5018476 - 5018474 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20625 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-10-11 Patch Publication Date: 2022-10-11 Plugin Modification Date: 2023-02-09 Exploitability Ease: Exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-10-11 | 2022-11-08 03:32 | 2023-05-09 07:10 | 6
|  |  | 167109 |  | KB5020010: Windows Server 2012 R2 Security Update (November 2022) | Apply Security Update 5020010 or Cumulative Update 5020023 | High | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5020010. It is, therefore, affected by multiple vulnerabilities - AMD: CVE-2022-23824 IBPB and Return Address Predictor Interactions (CVE-2022-23824) - Windows Kerberos RC4-HMAC Elevation of Privilege Vulnerability (CVE-2022-37966) - Windows Kerberos Elevation of Privilege Vulnerability (CVE-2022-37967) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5020010 or Cumulative Update 5020023 See Also: https://support.microsoft.com/help/5020010 https://support.microsoft.com/help/5020023 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5020023 - 5020010 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20666 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-11-08 Patch Publication Date: 2022-11-08 Plugin Modification Date: 2023-01-03 Exploitability Ease: Exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-11-08 | 2022-12-05 20:44 | 2023-05-09 07:10 | 5
|  |  | 168687 |  | KB5021296: Windows Server 2012 R2 Security Update (December 2022) | Apply Security Update 5021296 or Cumulative Update 5021294 | High | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5021296. It is, therefore, affected by multiple vulnerabilities - PowerShell Remote Code Execution Vulnerability (CVE-2022-41076) - A remote code execution vulnerability. An attacker can exploit this to bypass authentication and execute unauthorized arbitrary commands. (CVE-2022-44670, CVE-2022-44676) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5021296 or Cumulative Update 5021294 See Also: https://support.microsoft.com/en-us/help/5021294 https://support.microsoft.com/en-us/help/5021296 https://support.microsoft.com/help/5021294 https://support.microsoft.com/help/5021296 CVSS Base Score: 8.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5021296 - 5021294 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20716 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-12-13 Patch Publication Date: 2022-12-13 Plugin Modification Date: 2023-01-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-12-13 | 2023-01-09 04:57 | 2023-05-09 07:10 | 4
|  |  | 168745 |  | Security Updates for Microsoft .NET Framework (December 2022) | Microsoft has released security updates for Microsoft .NET Framework. | High | Synopsis: The Microsoft .NET Framework installation on the remote host is missing a security update. Description: The Microsoft .NET Framework installation on the remote host is missing a security update. It is, therefore, affected by a remote code execution vulnerability in the handling of XPS files. Solution: Microsoft has released security updates for Microsoft .NET Framework. See Also: http://www.nessus.org/u?0d29de7c http://www.nessus.org/u?e40dadbd https://support.microsoft.com/en-us/help/5020859 https://support.microsoft.com/en-us/help/5020860 https://support.microsoft.com/en-us/help/5020861 https://support.microsoft.com/en-us/help/5020862 https://support.microsoft.com/en-us/help/5020866 https://support.microsoft.com/en-us/help/5020867 https://support.microsoft.com/en-us/help/5020868 https://support.microsoft.com/en-us/help/5020869 https://support.microsoft.com/en-us/help/5020872 https://support.microsoft.com/en-us/help/5020873 https://support.microsoft.com/en-us/help/5020874 https://support.microsoft.com/en-us/help/5020875 https://support.microsoft.com/en-us/help/5020876 https://support.microsoft.com/en-us/help/5020877 https://support.microsoft.com/en-us/help/5020878 https://support.microsoft.com/en-us/help/5020879 https://support.microsoft.com/en-us/help/5020880 https://support.microsoft.com/en-us/help/5020881 https://support.microsoft.com/en-us/help/5020882 https://support.microsoft.com/en-us/help/5020883 https://support.microsoft.com/en-us/help/5020894 https://support.microsoft.com/en-us/help/5020895 https://support.microsoft.com/en-us/help/5020896 https://support.microsoft.com/en-us/help/5020897 https://support.microsoft.com/en-us/help/5020898 https://support.microsoft.com/en-us/help/5020899 https://support.microsoft.com/en-us/help/5020900 https://support.microsoft.com/en-us/help/5020901 https://support.microsoft.com/en-us/help/5020902 https://support.microsoft.com/en-us/help/5020903 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Microsoft .NET Framework 3.5 The remote host is missing one of the following rollup KBs : Security Only - 5020897 Cumulative - 5020862 None of the versions of 'System.Printing.dll' under C:\windows\WinSxS have been patched. Fixed version : 3.0.6920.8953 CPE: cpe:/a:microsoft:.net_framework CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-12-13 Patch Publication Date: 2022-12-13 Plugin Modification Date: 2023-02-16 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-12-15 | 2023-01-09 04:57 | 2023-05-09 07:10 | 4
|  |  | 171598 |  | Security Updates for Microsoft .NET Framework (February 2023) | Microsoft has released security updates for Microsoft .NET Framework. | High | Synopsis: The Microsoft .NET Framework installation on the remote host is missing a security update. Description: The Microsoft .NET Framework installation on the remote host is missing a security update. It is, therefore, affected by multiple vulnerabilities, as follows: - A denial of service (DoS) vulnerability. (CVE-2023-21722) - A remote code execution vulnerability. (CVE-2023-21808) Solution: Microsoft has released security updates for Microsoft .NET Framework. See Also: http://www.nessus.org/u?5bd7d30c http://www.nessus.org/u?42dae88f http://www.nessus.org/u?db0b1765 https://support.microsoft.com/en-us/help/5022497 https://support.microsoft.com/en-us/help/5022498 https://support.microsoft.com/en-us/help/5022499 https://support.microsoft.com/en-us/help/5022501 https://support.microsoft.com/en-us/help/5022502 https://support.microsoft.com/en-us/help/5022503 https://support.microsoft.com/en-us/help/5022504 https://support.microsoft.com/en-us/help/5022505 https://support.microsoft.com/en-us/help/5022506 https://support.microsoft.com/en-us/help/5022507 https://support.microsoft.com/en-us/help/5022508 https://support.microsoft.com/en-us/help/5022509 https://support.microsoft.com/en-us/help/5022511 https://support.microsoft.com/en-us/help/5022512 https://support.microsoft.com/en-us/help/5022513 https://support.microsoft.com/en-us/help/5022514 https://support.microsoft.com/en-us/help/5022515 https://support.microsoft.com/en-us/help/5022516 https://support.microsoft.com/en-us/help/5022520 https://support.microsoft.com/en-us/help/5022521 https://support.microsoft.com/en-us/help/5022522 https://support.microsoft.com/en-us/help/5022523 https://support.microsoft.com/en-us/help/5022524 https://support.microsoft.com/en-us/help/5022525 https://support.microsoft.com/en-us/help/5022526 https://support.microsoft.com/en-us/help/5022529 https://support.microsoft.com/en-us/help/5022530 https://support.microsoft.com/en-us/help/5022531 https://support.microsoft.com/en-us/help/5022574 https://support.microsoft.com/en-us/help/5022575 CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Microsoft .NET Framework 3.5 The remote host is missing one of the following rollup KBs : Security Only - 5022531 Cumulative - 5022525 C:\windows\Microsoft.NET\Framework\v2.0.50727\mscorlib.dll has not been patched. Remote version : 2.0.50727.8964 Should be : 2.0.50727.8966 CPE: cpe:/a:microsoft:.net_framework CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-02-14 Patch Publication Date: 2023-02-14 Plugin Modification Date: 2023-02-20 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2023-02-17 | 2023-03-07 01:19 | 2023-05-09 07:10 | 2
|  |  | 63155 |  | Microsoft Windows Unquoted Service Path Enumeration | Ensure that any services that contain a space in the path enclose the path in quotes. | High | Synopsis: The remote Windows host has at least one service installed that uses an unquoted service path. Description: The remote Windows host has at least one service installed that uses an unquoted service path, which contains at least one whitespace. A local attacker can gain elevated privileges by inserting an executable file in the path of the affected service. Note that this is a generic test that will flag any application affected by the described vulnerability. Solution: Ensure that any services that contain a space in the path enclose the path in quotes. See Also: http://www.nessus.org/u?84a4cc1c http://cwe.mitre.org/data/definitions/428.html https://www.commonexploits.com/unquoted-service-paths/ http://www.nessus.org/u?4aa6acbc CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:C/A:C Plugin Output: Nessus found the following services with an untrusted path : AxedaWatchdog : C:\Program Files (x86)\RVA_2.2.0.1825\Gateway\xWatchDog.exe AxedaGateway : C:\Program Files (x86)\RVA_2.2.0.1825\Gateway\xGate.exe CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2012-09-15 Patch Publication Date: NULL Plugin Modification Date: 2020-06-12 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Report Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE2 |  |  |  |  | VHAXXXAPPROCHE2.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2012-12-05 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 65057 |  | Insecure Windows Service Permissions | Ensure the groups listed above do not have permissions to modify or write service executables. Additionally, ensure these groups do not have Full Control permission to any directories that contain service executables. | High | Synopsis: At least one improperly configured Windows service may have a privilege escalation vulnerability. Description: At least one Windows service executable with insecure permissions was detected on the remote host. Services configured to use an executable with weak permissions are vulnerable to privilege escalation attacks. An unprivileged user could modify or overwrite the executable with arbitrary code, which would be executed the next time the service is started. Depending on the user that the service runs as, this could result in privilege escalation. This plugin checks if any of the following groups have permissions to modify executable files that are started by Windows services : - Everyone - Users - Domain Users - Authenticated Users Solution: Ensure the groups listed above do not have permissions to modify or write service executables. Additionally, ensure these groups do not have Full Control permission to any directories that contain service executables. See Also: http://www.nessus.org/u?e4e766b2 CVSS Base Score: 8.4 CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : d:\intersystems\cache\bin\cservice.exe Used by services : Cache_d-_intersystems_cache File write allowed for groups : Authenticated Users (S-1-5-11) Full control of directory allowed for groups : Authenticated Users (S-1-5-11) Path : d:\intersystems\cache\httpd\bin\httpd.exe Used by services : CACHEhttpd File write allowed for groups : Authenticated Users (S-1-5-11) Full control of directory allowed for groups : Authenticated Users (S-1-5-11) CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2022-08-19 Exploitability Ease: NULL Plugin Type: local Source File: | Report Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE2 |  |  |  |  | VHAXXXAPPROCHE2.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2013-03-06 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 81264 |  | MS15-011: Vulnerability in Group Policy Could Allow Remote Code Execution (3000483) | Microsoft has released a set of patches for Windows Vista, 2008, 7, 2008 R2, 8, 2012, 8.1, and 2012 R2. | High | Synopsis: The remote Windows host is affected by a remote code execution vulnerability. Description: The remote Windows host is affected by a remote code execution vulnerability due to how the Group Policy service manages policy data when a domain-joined system connects to a domain controller. An attacker, using a controlled network, can exploit this to gain complete control of the host. Note that Microsoft has no plans to release an update for Windows 2003 even though it is affected by this vulnerability. Solution: Microsoft has released a set of patches for Windows Vista, 2008, 7, 2008 R2, 8, 2012, 8.1, and 2012 R2. See Also: http://www.nessus.org/u?6533d970 http://www.nessus.org/u?b8230f41 CVSS Base Score: NULL CVSS Vector: CVSS2#AV:A/AC:L/Au:N/C:C/I:C/A:C Plugin Output: KB 3000483 or a related, subsequent update was successfully installed, but the GPO setting "Hardened UNC Paths" has not been enabled. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2015-02-10 Patch Publication Date: 2015-02-10 Plugin Modification Date: 2022-06-20 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Report Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE2 |  |  |  |  | VHAXXXAPPROCHE2.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2015-02-10 | 2019-02-01 21:38 | 2023-05-09 07:17 | 51
|  |  | 104896 |  | Security Updates for Internet Explorer (September 2017) | Microsoft has released security updates for the affected versions of Internet Explorer. | High | Synopsis: The Internet Explorer installation on the remote host is affected by multiple vulnerabilities. Description: The Internet Explorer installation on the remote host is missing security updates. It is, therefore, affected by multiple vulnerabilities : - An information disclosure vulnerability exists when affected Microsoft scripting engines do not properly handle objects in memory. The vulnerability could allow an attacker to detect specific files on the user's computer. (CVE-2017-8529) - A remote code execution vulnerability exists when Microsoft browsers improperly access objects in memory. The vulnerability could corrupt memory in such a way that an attacker could execute arbitrary code in the context of the current user. (CVE-2017-8750) - A spoofing vulnerability exists when Internet Explorer improperly handles specific HTML content. An attacker who successfully exploited this vulnerability could trick a user into believing that the user was visiting a legitimate website. The specially crafted website could either spoof content or serve as a pivot to chain an attack with other vulnerabilities in web services. (CVE-2017-8733) - A remote code execution vulnerability exists when Internet Explorer improperly accesses objects in memory. The vulnerability could corrupt memory in such a way that an attacker could execute arbitrary code in the context of the current user. (CVE-2017-8747, CVE-2017-8749) - A remote code execution vulnerability exists in the way that Microsoft browser JavaScript engines render content when handling objects in memory. The vulnerability could corrupt memory in such a way that an attacker could execute arbitrary code in the context of the current user. In a web-based attack scenario, an attacker could host a specially crafted website that is designed to exploit the vulnerability through Microsoft browsers and then convince a user to view the website. An attacker could also embed an ActiveX control marked &quot;safe for initialization&quot; in an application or Microsoft Office document that hosts the related rendering engine. The attacker could also take advantage of compromised websites, and websites that accept or host user-provided content or advertisements. These websites could contain specially crafted content that could exploit the vulnerability. An attacker who successfully exploited the vulnerability could gain the same user rights as the current user. (CVE-2017-8741, CVE-2017-8748) - An information disclosure vulnerability exists in Microsoft browsers due to improper parent domain verification in certain functionality. An attacker who successfully exploited the vulnerability could obtain specific information that is used in the parent domain. (CVE-2017-8736) Solution: Microsoft has released security updates for the affected versions of Internet Explorer. See Also: http://www.nessus.org/u?26b484bb http://www.nessus.org/u?085e4d22 http://www.nessus.org/u?35364720 http://www.nessus.org/u?1dbb18cc CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: The following registry key is required to enable the fix for CVE-2017-8529 and is missing. HKLM\SOFTWARE\Microsoft\Internet Explorer\Main\FeatureControl\FEATURE_ENABLE_PRINT_INFO_DISCLOSURE_FIX\iexplore.exe The following registry key is required to enable the fix for CVE-2017-8529 and is missing. HKLM\SOFTWARE\WOW6432Node\Microsoft\Internet Explorer\Main\FeatureControl\FEATURE_ENABLE_PRINT_INFO_DISCLOSURE_FIX\iexplore.exe CPE: cpe:/o:microsoft:windows cpe:/a:microsoft:ie CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2017-11-23 Patch Publication Date: 2017-11-23 Plugin Modification Date: 2020-07-17 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Report Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE2 |  |  |  |  | VHAXXXAPPROCHE2.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2017-11-30 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 166030 |  | KB5018476: Windows Server 2012 R2 Security Update (October 2022) | Apply Security Update 5018476 or Cumulative Update 5018474 | High | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5018476. It is, therefore, affected by multiple vulnerabilities - Server Service Remote Protocol Elevation of Privilege Vulnerability (CVE-2022-38045) - Microsoft ODBC Driver Remote Code Execution Vulnerability (CVE-2022-38040) - Microsoft WDAC OLE DB provider for SQL Server Remote Code Execution Vulnerability (CVE-2022-37982, CVE-2022-38031) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5018476 or Cumulative Update 5018474 See Also: https://support.microsoft.com/en-us/help/5018474 https://support.microsoft.com/en-us/help/5018476 https://support.microsoft.com/help/5018474 https://support.microsoft.com/help/5018476 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5018476 - 5018474 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20625 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-10-11 Patch Publication Date: 2022-10-11 Plugin Modification Date: 2023-02-09 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Report Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE2 |  |  |  |  | VHAXXXAPPROCHE2.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-10-11 | 2022-11-08 03:32 | 2023-05-09 07:17 | 6
|  |  | 167109 |  | KB5020010: Windows Server 2012 R2 Security Update (November 2022) | Apply Security Update 5020010 or Cumulative Update 5020023 | High | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5020010. It is, therefore, affected by multiple vulnerabilities - AMD: CVE-2022-23824 IBPB and Return Address Predictor Interactions (CVE-2022-23824) - Windows Kerberos RC4-HMAC Elevation of Privilege Vulnerability (CVE-2022-37966) - Windows Kerberos Elevation of Privilege Vulnerability (CVE-2022-37967) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5020010 or Cumulative Update 5020023 See Also: https://support.microsoft.com/help/5020010 https://support.microsoft.com/help/5020023 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5020023 - 5020010 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20666 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-11-08 Patch Publication Date: 2022-11-08 Plugin Modification Date: 2023-01-03 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Report Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE2 |  |  |  |  | VHAXXXAPPROCHE2.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-11-08 | 2022-12-05 20:44 | 2023-05-09 07:17 | 5
|  |  | 168687 |  | KB5021296: Windows Server 2012 R2 Security Update (December 2022) | Apply Security Update 5021296 or Cumulative Update 5021294 | High | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5021296. It is, therefore, affected by multiple vulnerabilities - PowerShell Remote Code Execution Vulnerability (CVE-2022-41076) - A remote code execution vulnerability. An attacker can exploit this to bypass authentication and execute unauthorized arbitrary commands. (CVE-2022-44670, CVE-2022-44676) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5021296 or Cumulative Update 5021294 See Also: https://support.microsoft.com/en-us/help/5021294 https://support.microsoft.com/en-us/help/5021296 https://support.microsoft.com/help/5021294 https://support.microsoft.com/help/5021296 CVSS Base Score: 8.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5021296 - 5021294 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20716 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-12-13 Patch Publication Date: 2022-12-13 Plugin Modification Date: 2023-01-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Report Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE2 |  |  |  |  | VHAXXXAPPROCHE2.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-12-13 | 2023-01-09 04:57 | 2023-05-09 07:17 | 4
|  |  | 168745 |  | Security Updates for Microsoft .NET Framework (December 2022) | Microsoft has released security updates for Microsoft .NET Framework. | High | Synopsis: The Microsoft .NET Framework installation on the remote host is missing a security update. Description: The Microsoft .NET Framework installation on the remote host is missing a security update. It is, therefore, affected by a remote code execution vulnerability in the handling of XPS files. Solution: Microsoft has released security updates for Microsoft .NET Framework. See Also: http://www.nessus.org/u?0d29de7c http://www.nessus.org/u?e40dadbd https://support.microsoft.com/en-us/help/5020859 https://support.microsoft.com/en-us/help/5020860 https://support.microsoft.com/en-us/help/5020861 https://support.microsoft.com/en-us/help/5020862 https://support.microsoft.com/en-us/help/5020866 https://support.microsoft.com/en-us/help/5020867 https://support.microsoft.com/en-us/help/5020868 https://support.microsoft.com/en-us/help/5020869 https://support.microsoft.com/en-us/help/5020872 https://support.microsoft.com/en-us/help/5020873 https://support.microsoft.com/en-us/help/5020874 https://support.microsoft.com/en-us/help/5020875 https://support.microsoft.com/en-us/help/5020876 https://support.microsoft.com/en-us/help/5020877 https://support.microsoft.com/en-us/help/5020878 https://support.microsoft.com/en-us/help/5020879 https://support.microsoft.com/en-us/help/5020880 https://support.microsoft.com/en-us/help/5020881 https://support.microsoft.com/en-us/help/5020882 https://support.microsoft.com/en-us/help/5020883 https://support.microsoft.com/en-us/help/5020894 https://support.microsoft.com/en-us/help/5020895 https://support.microsoft.com/en-us/help/5020896 https://support.microsoft.com/en-us/help/5020897 https://support.microsoft.com/en-us/help/5020898 https://support.microsoft.com/en-us/help/5020899 https://support.microsoft.com/en-us/help/5020900 https://support.microsoft.com/en-us/help/5020901 https://support.microsoft.com/en-us/help/5020902 https://support.microsoft.com/en-us/help/5020903 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Microsoft .NET Framework 3.5 The remote host is missing one of the following rollup KBs : Security Only - 5020897 Cumulative - 5020862 None of the versions of 'System.Printing.dll' under C:\windows\WinSxS have been patched. Fixed version : 3.0.6920.8953 CPE: cpe:/a:microsoft:.net_framework CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-12-13 Patch Publication Date: 2022-12-13 Plugin Modification Date: 2023-02-16 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Report Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE2 |  |  |  |  | VHAXXXAPPROCHE2.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-12-15 | 2023-01-09 04:57 | 2023-05-09 07:17 | 4
|  |  | 171598 |  | Security Updates for Microsoft .NET Framework (February 2023) | Microsoft has released security updates for Microsoft .NET Framework. | High | Synopsis: The Microsoft .NET Framework installation on the remote host is missing a security update. Description: The Microsoft .NET Framework installation on the remote host is missing a security update. It is, therefore, affected by multiple vulnerabilities, as follows: - A denial of service (DoS) vulnerability. (CVE-2023-21722) - A remote code execution vulnerability. (CVE-2023-21808) Solution: Microsoft has released security updates for Microsoft .NET Framework. See Also: http://www.nessus.org/u?5bd7d30c http://www.nessus.org/u?42dae88f http://www.nessus.org/u?db0b1765 https://support.microsoft.com/en-us/help/5022497 https://support.microsoft.com/en-us/help/5022498 https://support.microsoft.com/en-us/help/5022499 https://support.microsoft.com/en-us/help/5022501 https://support.microsoft.com/en-us/help/5022502 https://support.microsoft.com/en-us/help/5022503 https://support.microsoft.com/en-us/help/5022504 https://support.microsoft.com/en-us/help/5022505 https://support.microsoft.com/en-us/help/5022506 https://support.microsoft.com/en-us/help/5022507 https://support.microsoft.com/en-us/help/5022508 https://support.microsoft.com/en-us/help/5022509 https://support.microsoft.com/en-us/help/5022511 https://support.microsoft.com/en-us/help/5022512 https://support.microsoft.com/en-us/help/5022513 https://support.microsoft.com/en-us/help/5022514 https://support.microsoft.com/en-us/help/5022515 https://support.microsoft.com/en-us/help/5022516 https://support.microsoft.com/en-us/help/5022520 https://support.microsoft.com/en-us/help/5022521 https://support.microsoft.com/en-us/help/5022522 https://support.microsoft.com/en-us/help/5022523 https://support.microsoft.com/en-us/help/5022524 https://support.microsoft.com/en-us/help/5022525 https://support.microsoft.com/en-us/help/5022526 https://support.microsoft.com/en-us/help/5022529 https://support.microsoft.com/en-us/help/5022530 https://support.microsoft.com/en-us/help/5022531 https://support.microsoft.com/en-us/help/5022574 https://support.microsoft.com/en-us/help/5022575 CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Microsoft .NET Framework 3.5 The remote host is missing one of the following rollup KBs : Security Only - 5022531 Cumulative - 5022525 C:\windows\Microsoft.NET\Framework\v2.0.50727\mscorlib.dll has not been patched. Remote version : 2.0.50727.8964 Should be : 2.0.50727.8966 CPE: cpe:/a:microsoft:.net_framework CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-02-14 Patch Publication Date: 2023-02-14 Plugin Modification Date: 2023-02-20 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Report Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE2 |  |  |  |  | VHAXXXAPPROCHE2.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2023-02-17 | 2023-03-07 01:19 | 2023-05-09 07:17 | 2
|  |  | 63155 |  | Microsoft Windows Unquoted Service Path Enumeration | Ensure that any services that contain a space in the path enclose the path in quotes. | High | Synopsis: The remote Windows host has at least one service installed that uses an unquoted service path. Description: The remote Windows host has at least one service installed that uses an unquoted service path, which contains at least one whitespace. A local attacker can gain elevated privileges by inserting an executable file in the path of the affected service. Note that this is a generic test that will flag any application affected by the described vulnerability. Solution: Ensure that any services that contain a space in the path enclose the path in quotes. See Also: http://www.nessus.org/u?84a4cc1c http://cwe.mitre.org/data/definitions/428.html https://www.commonexploits.com/unquoted-service-paths/ http://www.nessus.org/u?4aa6acbc CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:C/A:C Plugin Output: Nessus found the following services with an untrusted path : Roche.DP.RSS.DataDistributionClient.StandAloneWinService : C:\Program Files (x86)\RVA\DDC\Roche.DP.RSS.DataDistributionClient.StandAloneWinService\Roche.DP.RSS.DataDistributionClient.StandAloneWinService.exe AxedaGateway : C:\Program Files (x86)\RVA\Gateway\xGate.exe AxedaWatchDog : C:\Program Files (x86)\RVA\Gateway\xWatchDog.exe CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2012-09-15 Patch Publication Date: NULL Plugin Modification Date: 2020-06-12 Exploitability Ease: Exploits are available Plugin Type: local Source File: | COBAS Infinity Production Server (v3.03) | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXCCACOBASP |  |  |  |  | VHAXXXCCACOBASP.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2012-12-05 | 2022-09-04 23:24 | 2023-05-09 07:18 | 8
|  |  | 65057 |  | Insecure Windows Service Permissions | Ensure the groups listed above do not have permissions to modify or write service executables. Additionally, ensure these groups do not have Full Control permission to any directories that contain service executables. | High | Synopsis: At least one improperly configured Windows service may have a privilege escalation vulnerability. Description: At least one Windows service executable with insecure permissions was detected on the remote host. Services configured to use an executable with weak permissions are vulnerable to privilege escalation attacks. An unprivileged user could modify or overwrite the executable with arbitrary code, which would be executed the next time the service is started. Depending on the user that the service runs as, this could result in privilege escalation. This plugin checks if any of the following groups have permissions to modify executable files that are started by Windows services : - Everyone - Users - Domain Users - Authenticated Users Solution: Ensure the groups listed above do not have permissions to modify or write service executables. Additionally, ensure these groups do not have Full Control permission to any directories that contain service executables. See Also: http://www.nessus.org/u?e4e766b2 CVSS Base Score: 8.4 CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : c:\roche\cobasinfinity\eppr\bin\rits-eppr_service.exe Used by services : EPPR_RocheService File write allowed for groups : Authenticated Users (S-1-5-11) Full control of directory allowed for groups : Authenticated Users (S-1-5-11) Path : c:\roche\cobasinfinity\healthshare\bin\cservice.exe Used by services : Cache_c-_roche_cobasinfinity_healthshare File write allowed for groups : Authenticated Users (S-1-5-11) Full control of directory allowed for groups : Authenticated Users (S-1-5-11) Path : c:\roche\cobasinfinity\healthshare\httpd\bin\httpd.exe Used by services : HSFoundationhttpd File write allowed for groups : Authenticated Users (S-1-5-11) Full control of directory allowed for groups : Authenticated Users (S-1-5-11) Path : c:\roche\cobasinfinity\imagegenerator\bin\rits-igen_service.exe Used by services : RITs-IGen_RocheService File write allowed for groups : Authenticated Users (S-1-5-11) Full control of directory allowed for groups : Authenticated Users (S-1-5-11) Path : c:\roche\cobasinfinity\qc\bin\rits-qc_service.exe Used by services : DarwinQC_RocheService File write allowed for groups : Authenticated Users (S-1-5-11) Full control of directory allowed for groups : Authenticated Users (S-1-5-11) Path : c:\roche\cobasinfinity\reportsserver\bin\rits-reportsserver_service.exe Used by services : ReportsServer_RocheService File write allowed for groups : Authenticated Users (S-1-5-11) Full control of directory allowed for groups : Authenticated Users (S-1-5-11) CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2022-08-19 Exploitability Ease: NULL Plugin Type: local Source File: | COBAS Infinity Production Server (v3.03) | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXCCACOBASP |  |  |  |  | VHAXXXCCACOBASP.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2013-03-06 | 2022-09-04 23:24 | 2023-05-09 07:18 | 8
|  |  | 153585 |  | Apache >= 2.4.17 < 2.4.49 mod_http2 | Upgrade to Apache version 2.4.49 or later. | High | Synopsis: The remote web server is affected by a vulnerability. Description: The version of Apache httpd installed on the remote host is greater than 2.4.17 and prior to 2.4.49. It is, therefore, affected by a vulnerability as referenced in the 2.4.49 changelog. A crafted method sent through HTTP/2 will bypass validation and be forwarded by mod_proxy, which can lead to request splitting or cache poisoning. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.49 or later. See Also: https://downloads.apache.org/httpd/CHANGES_2.4 https://httpd.apache.org/security/vulnerabilities_24.html CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:P/A:N Plugin Output: Path : C:\Roche\CobasInfinity\Apache24\bin\httpd.exe Installed version : 2.4.46.0 Fixed version : 2.4.49 CPE: cpe:/a:apache:httpd cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-09-16 Patch Publication Date: 2021-09-16 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | COBAS Infinity Production Server (v3.03) | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXCCACOBASP |  |  |  |  | VHAXXXCCACOBASP.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2021-09-23 | 2022-09-04 23:24 | 2023-05-09 07:18 | 8
|  |  | 153586 |  | Apache >= 2.4.30 < 2.4.49 mod_proxy_uwsgi | Upgrade to Apache version 2.4.49 or later. | High | Synopsis: The remote web server is affected by a vulnerability. Description: The version of Apache httpd installed on the remote host greater than 2.4.30 and is prior to 2.4.49. It is, therefore, affected by a vulnerability as referenced in the 2.4.49 changelog. A carefully crafted request uri-path can cause mod_proxy_uwsgi to read above the allocated memory and crash (DoS). Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.49 or later. See Also: https://downloads.apache.org/httpd/CHANGES_2.4 https://httpd.apache.org/security/vulnerabilities_24.html CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Path : C:\Roche\CobasInfinity\Apache24\bin\httpd.exe Installed version : 2.4.46.0 Fixed version : 2.4.49 CPE: cpe:/a:apache:httpd cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-09-16 Patch Publication Date: 2021-09-16 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | COBAS Infinity Production Server (v3.03) | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXCCACOBASP |  |  |  |  | VHAXXXCCACOBASP.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2021-09-23 | 2022-09-04 23:24 | 2023-05-09 07:18 | 8
|  |  | 12085 |  | Apache Tomcat Default Files | Delete the default index page and remove the example JSP and servlets. Follow the Tomcat or OWASP instructions to replace or modify the default error page. | Medium | Synopsis: The remote web server contains default files. Description: The default error page, default index page, example JSPs and/or example servlets are installed on the remote Apache Tomcat server. These files should be removed as they may help an attacker uncover information about the remote Tomcat install or host itself. Solution: Delete the default index page and remove the example JSP and servlets. Follow the Tomcat or OWASP instructions to replace or modify the default error page. See Also: http://www.nessus.org/u?4cb3b4dd https://www.owasp.org/index.php/Securing_tomcat CVSS Base Score: 5.3 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:N/A:N Plugin Output: The server is not configured to return a custom page in the event of a client requesting a non-existent resource. This may result in a potential disclosure of sensitive information about the server to attackers. CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2019-08-12 Exploitability Ease: NULL Plugin Type: remote Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 80 | Web Servers | Repeat Vulnerability | 2004-03-02 | 2020-06-11 18:17 | 2023-05-07 06:58 | 35
|  |  | 12085 |  | Apache Tomcat Default Files | Delete the default index page and remove the example JSP and servlets. Follow the Tomcat or OWASP instructions to replace or modify the default error page. | Medium | Synopsis: The remote web server contains default files. Description: The default error page, default index page, example JSPs and/or example servlets are installed on the remote Apache Tomcat server. These files should be removed as they may help an attacker uncover information about the remote Tomcat install or host itself. Solution: Delete the default index page and remove the example JSP and servlets. Follow the Tomcat or OWASP instructions to replace or modify the default error page. See Also: http://www.nessus.org/u?4cb3b4dd https://www.owasp.org/index.php/Securing_tomcat CVSS Base Score: 5.3 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:N/A:N Plugin Output: The server is not configured to return a custom page in the event of a client requesting a non-existent resource. This may result in a potential disclosure of sensitive information about the server to attackers. CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2019-08-12 Exploitability Ease: NULL Plugin Type: remote Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 8080 | Web Servers | Repeat Vulnerability | 2004-03-02 | 2020-07-13 18:42 | 2023-05-07 06:58 | 34
|  |  | 15901 |  | SSL Certificate Expiry | Purchase or generate a new SSL certificate to replace the existing one. | Medium | Synopsis: The remote server's SSL certificate has already expired. Description: This plugin checks expiry dates of certificates associated with SSL- enabled services on the target and reports whether any have already expired. Solution: Purchase or generate a new SSL certificate to replace the existing one. See Also: NULL CVSS Base Score: 5.3 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:P/A:N Plugin Output: The SSL certificate has already expired : Subject : CN=HP Jetdirect A04F5C4E, O=Hewlett-Packard Co., OU=7446A04F5C4E, OU=J8011E Issuer : CN=HP Jetdirect A04F5C4E, O=Hewlett-Packard Co., OU=7446A04F5C4E, OU=J8011E Not valid before : Feb 1 00:00:00 2013 GMT Not valid after : Feb 1 00:00:00 2018 GMT CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2021-02-03 Exploitability Ease: NULL Plugin Type: remote Source File: | HP COLORJET PRINTER | HEWLETT-PACKARD CO |  |  | Region X | VISN X | VAMC | [7 empty] | HP JetDirect | tcp | 443 | General | Repeat Vulnerability | 2004-12-03 | 2022-12-05 05:27 | 2023-05-09 03:44 | 5
|  |  | 51192 |  | SSL Certificate Cannot Be Trusted | Purchase or generate a proper SSL certificate for this service. | Medium | Synopsis: The SSL certificate for this service cannot be trusted. Description: The server's X.509 certificate cannot be trusted. This situation can occur in three different ways, in which the chain of trust can be broken, as stated below : - First, the top of the certificate chain sent by the server might not be descended from a known public certificate authority. This can occur either when the top of the chain is an unrecognized, self-signed certificate, or when intermediate certificates are missing that would connect the top of the certificate chain to a known public certificate authority. - Second, the certificate chain may contain a certificate that is not valid at the time of the scan. This can occur either when the scan occurs before one of the certificate's 'notBefore' dates, or after one of the certificate's 'notAfter' dates. - Third, the certificate chain may contain a signature that either didn't match the certificate's information or could not be verified. Bad signatures can be fixed by getting the certificate with the bad signature to be re-signed by its issuer. Signatures that could not be verified are the result of the certificate's issuer using a signing algorithm that Nessus either does not support or does not recognize. If the remote host is a public host in production, any break in the chain makes it more difficult for users to verify the authenticity and identity of the web server. This could make it easier to carry out man-in-the-middle attacks against the remote host. Solution: Purchase or generate a proper SSL certificate for this service. See Also: https://www.itu.int/rec/T-REC-X.509/en https://en.wikipedia.org/wiki/X.509 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:N Plugin Output: The following certificate was at the top of the certificate chain sent by the remote host, but it is signed by an unknown certificate authority : ¦-Subject : CN=localhost ¦-Issuer : CN=localhost CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2020-04-27 Exploitability Ease: NULL Plugin Type: remote Source File: | X-Ray Room | SIEMENS HEALTHCARE |  |  | Region X | VISN X | VAMC | [7 empty] | Microsoft Windows 10 | tcp | 443 | General | Repeat Vulnerability | 2010-12-15 | 2023-02-05 05:39 | 2023-05-07 07:51 | 3
|  |  | 118036 |  | Apache Tomcat 8.5.x < 8.5.34 Open Redirect Weakness | Upgrade to Apache Tomcat version 8.5.34 or later. | Medium | Synopsis: The remote Apache Tomcat server is affected by a open redirect vulnerability. Description: The version of Apache Tomcat installed on the remote host is 8.5.x prior to 8.5.34. It is, therefore, affected by a open redirect vulnerability. Solution: Upgrade to Apache Tomcat version 8.5.34 or later. See Also: http://www.nessus.org/u?1bddf0bb CVSS Base Score: 4.3 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:N/I:P/A:N Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.34 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-10-03 Patch Publication Date: 2018-09-19 Plugin Modification Date: 2022-04-11 Exploitability Ease: Exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 80 | Web Servers | Repeat Vulnerability | 2018-10-10 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 118036 |  | Apache Tomcat 8.5.x < 8.5.34 Open Redirect Weakness | Upgrade to Apache Tomcat version 8.5.34 or later. | Medium | Synopsis: The remote Apache Tomcat server is affected by a open redirect vulnerability. Description: The version of Apache Tomcat installed on the remote host is 8.5.x prior to 8.5.34. It is, therefore, affected by a open redirect vulnerability. Solution: Upgrade to Apache Tomcat version 8.5.34 or later. See Also: http://www.nessus.org/u?1bddf0bb CVSS Base Score: 4.3 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:N/I:P/A:N Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.34 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-10-03 Patch Publication Date: 2018-09-19 Plugin Modification Date: 2022-04-11 Exploitability Ease: Exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 8080 | Web Servers | Repeat Vulnerability | 2018-10-10 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 141446 |  | Apache Tomcat 8.5.x < 8.5.58 / 9.0.x < 9.0.38 HTTP/2 Request Mix-Up | Upgrade to Apache Tomcat version 8.5.58, 9.0.38 or later. | Medium | Synopsis: The remote Apache Tomcat server is affected by a vulnerability Description: The version of Tomcat installed on the remote host is 8.5.x prior to 8.5.58 or 9.0.x prior to 9.0.38. It is, therefore, affected by a vulnerability. If an HTTP/2 client exceeds the agreed maximum number of concurrent streams for a connection (in violation of the HTTP/2 protocol), it is possible that a subsequent request made on that connection could contain HTTP headers - including HTTP/2 pseudo headers - from a previous request rather than the intended headers. This can lead to users seeing responses for unexpected resources. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.58, 9.0.38 or later. See Also: http://www.nessus.org/u?0656cf04 http://www.nessus.org/u?771617a1 CVSS Base Score: 4.3 CVSS Vector: CVSS2#AV:N/AC:L/Au:S/C:P/I:N/A:N Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.58 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-09-15 Patch Publication Date: 2020-09-15 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 80 | Web Servers | Repeat Vulnerability | 2020-10-14 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 141446 |  | Apache Tomcat 8.5.x < 8.5.58 / 9.0.x < 9.0.38 HTTP/2 Request Mix-Up | Upgrade to Apache Tomcat version 8.5.58, 9.0.38 or later. | Medium | Synopsis: The remote Apache Tomcat server is affected by a vulnerability Description: The version of Tomcat installed on the remote host is 8.5.x prior to 8.5.58 or 9.0.x prior to 9.0.38. It is, therefore, affected by a vulnerability. If an HTTP/2 client exceeds the agreed maximum number of concurrent streams for a connection (in violation of the HTTP/2 protocol), it is possible that a subsequent request made on that connection could contain HTTP headers - including HTTP/2 pseudo headers - from a previous request rather than the intended headers. This can lead to users seeing responses for unexpected resources. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.58, 9.0.38 or later. See Also: http://www.nessus.org/u?0656cf04 http://www.nessus.org/u?771617a1 CVSS Base Score: 4.3 CVSS Vector: CVSS2#AV:N/AC:L/Au:S/C:P/I:N/A:N Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.58 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-09-15 Patch Publication Date: 2020-09-15 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 8080 | Web Servers | Repeat Vulnerability | 2020-10-14 | 2021-02-08 22:51 | 2023-05-07 06:58 | 27
|  |  | 151502 |  | Apache Tomcat 7.0.x <= 7.0.108 / 8.5.x <= 8.5.65 / 9.0.x <= 9.0.45 / 10.0.x <= 10.0.5 vulnerability | Upgrade to Apache Tomcat version 7.0.109, 8.5.66, 9.0.46, 10.0.6 or later. | Medium | Synopsis: The remote Apache Tomcat server is affected by a vulnerability Description: The version of Tomcat installed on the remote host is 7.0.x <= 7.0.108 / 8.5.x <= 8.5.65 / 9.0.x <= 9.0.45 / 10.0.x <= 10.0.5. It is, therefore, affected by a vulnerability as referenced in the fixed_in_apache_tomcat_10.0.6_security-10 advisory. - Queries made by the JNDI Realm did not always correctly escape parameters. Parameter values could be sourced from user provided data (eg user names) as well as configuration data provided by an administrator. In limited circumstances it was possible for users to authenticate using variations of their user name and/or to bypass some of the protection provided by the LockOut Realm. (CVE-2021-30640) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 7.0.109, 8.5.66, 9.0.46, 10.0.6 or later. See Also: http://www.nessus.org/u?d3fb2d8e http://www.nessus.org/u?0fb6f5ab http://www.nessus.org/u?0d761c19 http://www.nessus.org/u?ddfa2b5e http://www.nessus.org/u?95156892 http://www.nessus.org/u?ed08487c http://www.nessus.org/u?806274b5 http://www.nessus.org/u?f104a57d https://bz.apache.org/bugzilla/show_bug.cgi?id=65224 http://www.nessus.org/u?837a9443 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:P/A:N Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.66 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 80 | Web Servers | Repeat Vulnerability | 2021-07-12 | 2022-03-06 01:21 | 2023-05-07 06:58 | 14
|  |  | 151502 |  | Apache Tomcat 7.0.x <= 7.0.108 / 8.5.x <= 8.5.65 / 9.0.x <= 9.0.45 / 10.0.x <= 10.0.5 vulnerability | Upgrade to Apache Tomcat version 7.0.109, 8.5.66, 9.0.46, 10.0.6 or later. | Medium | Synopsis: The remote Apache Tomcat server is affected by a vulnerability Description: The version of Tomcat installed on the remote host is 7.0.x <= 7.0.108 / 8.5.x <= 8.5.65 / 9.0.x <= 9.0.45 / 10.0.x <= 10.0.5. It is, therefore, affected by a vulnerability as referenced in the fixed_in_apache_tomcat_10.0.6_security-10 advisory. - Queries made by the JNDI Realm did not always correctly escape parameters. Parameter values could be sourced from user provided data (eg user names) as well as configuration data provided by an administrator. In limited circumstances it was possible for users to authenticate using variations of their user name and/or to bypass some of the protection provided by the LockOut Realm. (CVE-2021-30640) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 7.0.109, 8.5.66, 9.0.46, 10.0.6 or later. See Also: http://www.nessus.org/u?d3fb2d8e http://www.nessus.org/u?0fb6f5ab http://www.nessus.org/u?0d761c19 http://www.nessus.org/u?ddfa2b5e http://www.nessus.org/u?95156892 http://www.nessus.org/u?ed08487c http://www.nessus.org/u?806274b5 http://www.nessus.org/u?f104a57d https://bz.apache.org/bugzilla/show_bug.cgi?id=65224 http://www.nessus.org/u?837a9443 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:P/A:N Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.66 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 8080 | Web Servers | Repeat Vulnerability | 2021-07-12 | 2022-03-06 01:21 | 2023-05-07 06:58 | 14
|  |  | 152183 |  | Apache Tomcat 8.5.0 < 8.5.68 vulnerability | Upgrade to Apache Tomcat version 8.5.68 or later. | Medium | Synopsis: The remote Apache Tomcat server is affected by a vulnerability Description: The version of Tomcat installed on the remote host is prior to 8.5.68. It is, therefore, affected by a vulnerability as referenced in the fixed_in_apache_tomcat_8.5.68_security-8 advisory. - Apache Tomcat 10.0.0-M1 to 10.0.6, 9.0.0.M1 to 9.0.46 and 8.5.0 to 8.5.66 did not correctly parse the HTTP transfer-encoding request header in some circumstances leading to the possibility to request smuggling when used with a reverse proxy. Specifically: - Tomcat incorrectly ignored the transfer encoding header if the client declared it would only accept an HTTP/1.0 response; - Tomcat honoured the identify encoding; and - Tomcat did not ensure that, if present, the chunked encoding was the final encoding. (CVE-2021-33037) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.68 or later. See Also: http://www.nessus.org/u?ca148f18 http://www.nessus.org/u?9e0e6b06 http://www.nessus.org/u?bea8fba1 http://www.nessus.org/u?836aea5f CVSS Base Score: 5.3 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:P/A:N Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.68 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-06-15 Patch Publication Date: 2021-06-15 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 80 | Web Servers | Repeat Vulnerability | 2021-08-03 | 2021-09-05 06:23 | 2023-05-07 06:58 | 20
|  |  | 152183 |  | Apache Tomcat 8.5.0 < 8.5.68 vulnerability | Upgrade to Apache Tomcat version 8.5.68 or later. | Medium | Synopsis: The remote Apache Tomcat server is affected by a vulnerability Description: The version of Tomcat installed on the remote host is prior to 8.5.68. It is, therefore, affected by a vulnerability as referenced in the fixed_in_apache_tomcat_8.5.68_security-8 advisory. - Apache Tomcat 10.0.0-M1 to 10.0.6, 9.0.0.M1 to 9.0.46 and 8.5.0 to 8.5.66 did not correctly parse the HTTP transfer-encoding request header in some circumstances leading to the possibility to request smuggling when used with a reverse proxy. Specifically: - Tomcat incorrectly ignored the transfer encoding header if the client declared it would only accept an HTTP/1.0 response; - Tomcat honoured the identify encoding; and - Tomcat did not ensure that, if present, the chunked encoding was the final encoding. (CVE-2021-33037) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.68 or later. See Also: http://www.nessus.org/u?ca148f18 http://www.nessus.org/u?9e0e6b06 http://www.nessus.org/u?bea8fba1 http://www.nessus.org/u?836aea5f CVSS Base Score: 5.3 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:P/A:N Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.68 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-06-15 Patch Publication Date: 2021-06-15 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 8080 | Web Servers | Repeat Vulnerability | 2021-08-03 | 2021-09-05 06:23 | 2023-05-07 06:58 | 20
|  |  | 157288 |  | TLS Version 1.1 Protocol Deprecated | Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. | Medium | Synopsis: The remote service encrypts traffic using an older version of TLS. Description: The remote service accepts connections encrypted using TLS 1.1. TLS 1.1 lacks support for current and recommended cipher suites. Ciphers that support encryption before MAC computation, and authenticated encryption modes such as GCM cannot be used with TLS 1.1 As of March 31, 2020, Endpoints that are not enabled for TLS 1.2 and higher will no longer function properly with major web browsers and major vendors. Solution: Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. See Also: https://datatracker.ietf.org/doc/html/rfc8996 http://www.nessus.org/u?c8ae820d CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:P/A:N Plugin Output: TLSv1.1 is enabled and the server supports at least one cipher. CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-04-19 Exploitability Ease: NULL Plugin Type: remote Source File: | X-Ray Room | SIEMENS HEALTHCARE |  |  | Region X | VISN X | VAMC | [7 empty] | Microsoft Windows 10 | tcp | 443 | Service detection | Repeat Vulnerability | 2022-04-04 | 2023-02-05 05:39 | 2023-05-07 07:51 | 3
|  |  | 173256 |  | Apache Tomcat 8.5.0 < 8.5.86 | Upgrade to Apache Tomcat version 8.5.86 or later. | Medium | Synopsis: The remote Apache Tomcat server is affected by a vulnerability Description: The version of Tomcat installed on the remote host is prior to 8.5.86. It is, therefore, affected by a vulnerability as referenced in the fixed_in_apache_tomcat_8.5.86_security-8 advisory. - When using the RemoteIpFilter with requests received from a reverse proxy via HTTP that include the X-Forwarded-Proto header set to https, session cookies created by Apache Tomcat 11.0.0-M1 to 11.0.0.-M2, 10.1.0-M1 to 10.1.5, 9.0.0-M1 to 9.0.71 and 8.5.0 to 8.5.85 did not include the secure attribute. This could result in the user agent transmitting the session cookie over an insecure channel. (CVE-2023-28708) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.86 or later. See Also: http://www.nessus.org/u?bff3d148 https://bz.apache.org/bugzilla/show_bug.cgi?id=66471 http://www.nessus.org/u?e6187e29 CVSS Base Score: 4.3 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:N/A:N Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.86 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-03-22 Patch Publication Date: NULL Plugin Modification Date: 2023-03-28 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 80 | Web Servers | Repeat Vulnerability | 2023-03-22 | 2023-04-02 06:58 | 2023-05-07 06:58 | 1
|  |  | 173256 |  | Apache Tomcat 8.5.0 < 8.5.86 | Upgrade to Apache Tomcat version 8.5.86 or later. | Medium | Synopsis: The remote Apache Tomcat server is affected by a vulnerability Description: The version of Tomcat installed on the remote host is prior to 8.5.86. It is, therefore, affected by a vulnerability as referenced in the fixed_in_apache_tomcat_8.5.86_security-8 advisory. - When using the RemoteIpFilter with requests received from a reverse proxy via HTTP that include the X-Forwarded-Proto header set to https, session cookies created by Apache Tomcat 11.0.0-M1 to 11.0.0.-M2, 10.1.0-M1 to 10.1.5, 9.0.0-M1 to 9.0.71 and 8.5.0 to 8.5.85 did not include the secure attribute. This could result in the user agent transmitting the session cookie over an insecure channel. (CVE-2023-28708) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache Tomcat version 8.5.86 or later. See Also: http://www.nessus.org/u?bff3d148 https://bz.apache.org/bugzilla/show_bug.cgi?id=66471 http://www.nessus.org/u?e6187e29 CVSS Base Score: 4.3 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:N/A:N Plugin Output: Installed version : 8.5.29 Fixed version : 8.5.86 CPE: cpe:/a:apache:tomcat CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-03-22 Patch Publication Date: NULL Plugin Modification Date: 2023-03-28 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | RAPID SERVER | VM |  |  | Region X | VISN X | VAMC | [7 empty] | Ubuntu 16.04 Linux Kernel 4.4 | tcp | 8080 | Web Servers | Repeat Vulnerability | 2023-03-22 | 2023-04-02 06:58 | 2023-05-07 06:58 | 1
|  |  | 142960 |  | HSTS Missing From HTTPS Server (RFC 6797) | Configure the remote web server to use HSTS. | Medium | Synopsis: The remote web server is not enforcing HSTS, as defined by RFC 6797. Description: The remote web server is not enforcing HSTS, as defined by RFC 6797. HSTS is an optional response header that can be configured on the server to instruct the browser to only communicate via HTTPS. The lack of HSTS allows downgrade attacks, SSL-stripping man-in-the-middle attacks, and weakens cookie-hijacking protections. Solution: Configure the remote web server to use HSTS. See Also: https://tools.ietf.org/html/rfc6797 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:P/A:N Plugin Output: The remote HTTPS server does not send the HTTP "Strict-Transport-Security" header. CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-04-21 Exploitability Ease: NULL Plugin Type: remote Source File: | Clinicomp Server | CLINICOMP INTERNATIONAL |  |  | Region X | VISN X | VAMC | [7 empty] | Linux Kernel 3.12 | tcp | 8443 | Web Servers | Repeat Vulnerability | 2020-11-17 | 2021-05-02 18:41 | 2023-05-09 04:02 | 24
|  |  | 142960 |  | HSTS Missing From HTTPS Server (RFC 6797) | Configure the remote web server to use HSTS. | Medium | Synopsis: The remote web server is not enforcing HSTS, as defined by RFC 6797. Description: The remote web server is not enforcing HSTS, as defined by RFC 6797. HSTS is an optional response header that can be configured on the server to instruct the browser to only communicate via HTTPS. The lack of HSTS allows downgrade attacks, SSL-stripping man-in-the-middle attacks, and weakens cookie-hijacking protections. Solution: Configure the remote web server to use HSTS. See Also: https://tools.ietf.org/html/rfc6797 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:P/A:N Plugin Output: The remote HTTPS server does not send the HTTP "Strict-Transport-Security" header. CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-04-21 Exploitability Ease: NULL Plugin Type: remote Source File: | Clinicomp Server | CLINICOMP INTERNATIONAL |  |  | Region X | VISN X | VAMC | [7 empty] | Linux Kernel 3.12 | tcp | 8443 | Web Servers | Repeat Vulnerability | 2020-11-17 | 2021-07-11 07:59 | 2023-05-09 04:02 | 22
|  |  | 38654 |  | McAfee Antivirus ZIP / RAR Scan Evasion | Update the McAfee DAT file to version 5600 or later. | Medium | Synopsis: An antivirus application installed on the remote host is affected by a scan evasion vulnerability. Description: The McAfee antivirus application installed on the remote host is affected by a scan evasion vulnerability due to the virus definitions being out of date. In this case, the DAT file version of the installed antivirus product is prior to 5600. An attacker can exploit this, by embedding malicious code in a specially crafted ZIP or RAR file, to evade detection by the scanning engine. Solution: Update the McAfee DAT file to version 5600 or later. See Also: http://www.nessus.org/u?ccdf87f9 https://seclists.org/fulldisclosure/2009/Apr/309 http://www.nessus.org/u?24888ca6 CVSS Base Score: NULL CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:P/A:N Plugin Output: DAT file version 1111.0000 is installed on the remote system. CPE: cpe:/a:mcafee:virusscan_plus cpe:/a:mcafee:mcafee:total_protection cpe:/a:mcafee:internet_security_suite cpe:/a:mcafee:virusscan_usb cpe:/a:mcafee:virusscan_enterprise cpe:/a:mcafee:virusscan_commandline cpe:/a:mcafee:securityshield_for_microsoft_isa_server cpe:/a:mcafee:securityshield_for_microsoft_sharepoint cpe:/a:mcafee:securityshield_for_email_servers cpe:/a:mcafee:email_gateway cpe:/a:mcafee:total_protection_for_endpoint cpe:/a:mcafee:active_virus_defense cpe:/a:mcafee:active_virusscan CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2009-04-30 Patch Publication Date: 2009-04-29 Plugin Modification Date: 2018-11-15 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2009-05-01 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 42290 |  | McAfee Antivirus TAR / PDF Scan Evasion (SB10003) | Update the McAfee DAT file to version 5693 or later. | Medium | Synopsis: An antivirus application installed on the remote host is affected by a scan evasion vulnerability. Description: The McAfee antivirus application installed on the remote host is affected by a scan evasion vulnerability due to the virus definitions being out of date. In this case, the DAT file version of the installed antivirus product is prior to 5693. An attacker can exploit this, by embedding malicious code in a specially crafted TAR or PDF file, to evade detection by the scanning engine. Solution: Update the McAfee DAT file to version 5693 or later. See Also: http://www.g-sec.lu/mcafee-pdf-bypass.html https://seclists.org/bugtraq/2009/Oct/273 https://kc.mcafee.com/corporate/index?page=content&id=SB10003 CVSS Base Score: NULL CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:P/A:N Plugin Output: DAT file version 1111.0000 is installed on the remote system. CPE: cpe:/a:mcafee:virusscan_plus cpe:/a:mcafee:mcafee:total_protection cpe:/a:mcafee:internet_security_suite cpe:/a:mcafee:virusscan_usb cpe:/a:mcafee:virusscan_enterprise cpe:/a:mcafee:virusscan_commandline cpe:/a:mcafee:securityshield_for_microsoft_isa_server cpe:/a:mcafee:securityshield_for_microsoft_sharepoint cpe:/a:mcafee:securityshield_for_email_servers cpe:/a:mcafee:email_gateway cpe:/a:mcafee:total_protection_for_endpoint cpe:/a:mcafee:active_virus_defense cpe:/a:mcafee:active_virusscan cpe:/a:mcafee:linuxshield cpe:/a:mcafee:groupshield cpe:/a:mcafee:netshield cpe:/a:mcafee:portalshield cpe:/a:mcafee:virex CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2009-10-27 Patch Publication Date: 2009-10-20 Plugin Modification Date: 2019-01-02 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2009-10-28 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 81126 |  | Apache 2.4.x < 2.4.12 Multiple Vulnerabilities | Upgrade to Apache version 2.4.12 or later. Alternatively, ensure that the affected modules are not in use. | Medium | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: According to its banner, the version of Apache 2.4.x running on the remote host is prior to 2.4.12. It is, therefore, affected by the following vulnerabilities : - A flaw exists in module mod_headers that can allow HTTP trailers to replace HTTP headers late during request processing, which a remote attacker can exploit to inject arbitrary headers. This can also cause some modules to function incorrectly or appear to function incorrectly. (CVE-2013-5704) - A NULL pointer dereference flaw exists in module mod_cache. A remote attacker, using an empty HTTP Content-Type header, can exploit this vulnerability to crash a caching forward proxy configuration, resulting in a denial of service if using a threaded MPM. (CVE-2014-3581) - A out-of-bounds memory read flaw exists in module mod_proxy_fcgi. An attacker, using a remote FastCGI server to send long response headers, can exploit this vulnerability to cause a denial of service by causing a buffer over-read. (CVE-2014-3583) - A flaw exists in module mod_lua when handling a LuaAuthzProvider used in multiple Require directives with different arguments. An attacker can exploit this vulnerability to bypass intended access restrictions. (CVE-2014-8109) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.12 or later. Alternatively, ensure that the affected modules are not in use. See Also: https://archive.apache.org/dist/httpd/CHANGES_2.4.12 http://httpd.apache.org/security/vulnerabilities_24.html CVSS Base Score: 5.3 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:P/A:N Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.12 CPE: cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2013-10-19 Patch Publication Date: 2015-01-30 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2015-02-02 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 84959 |  | Apache 2.4.x < 2.4.16 Multiple Vulnerabilities | Upgrade to Apache version 2.4.16 or later. Alternatively, ensure that the affected modules are not in use. | Medium | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: According to its banner, the version of Apache 2.4.x installed on the remote host is prior to 2.4.16. It is, therefore, affected by the following vulnerabilities : - A flaw exists in the lua_websocket_read() function in the 'mod_lua' module due to incorrect handling of WebSocket PING frames. A remote attacker can exploit this, by sending a crafted WebSocket PING frame after a Lua script has called the wsupgrade() function, to crash a child process, resulting in a denial of service condition. (CVE-2015-0228) - A NULL pointer dereference flaw exists in the read_request_line() function due to a failure to initialize the protocol structure member. A remote attacker can exploit this flaw, on installations that enable the INCLUDES filter and has an ErrorDocument 400 directive specifying a local URI, by sending a request that lacks a method, to cause a denial of service condition. (CVE-2015-0253) - A flaw exists in the chunked transfer coding implementation due to a failure to properly parse chunk headers. A remote attacker can exploit this to conduct HTTP request smuggling attacks. (CVE-2015-3183) - A flaw exists in the ap_some_auth_required() function due to a failure to consider that a Require directive may be associated with an authorization setting rather than an authentication setting. A remote attacker can exploit this, if a module that relies on the 2.2 API behavior exists, to bypass intended access restrictions. (CVE-2015-3185) - A flaw exists in the RC4 algorithm due to an initial double-byte bias in the keystream generation. An attacker can exploit this, via Bayesian analysis that combines an a priori plaintext distribution with keystream distribution statistics, to conduct a plaintext recovery of the ciphertext. Note that RC4 cipher suites are prohibited per RFC 7465. This issue was fixed in Apache version 2.4.13; however, 2.4.13, 2.4.14, and 2.4.15 were never publicly released. Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.16 or later. Alternatively, ensure that the affected modules are not in use. See Also: https://archive.apache.org/dist/httpd/CHANGES_2.4.16 http://httpd.apache.org/security/vulnerabilities_24.html http://www.nessus.org/u?7ec9a07a https://tools.ietf.org/html/rfc7465 CVSS Base Score: 5.3 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:P/A:N Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.16 CPE: cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2015-02-03 Patch Publication Date: 2015-07-15 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2015-07-23 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 117807 |  | Apache 2.4.x < 2.4.35 DoS | Upgrade to Apache version 2.4.35 or later. | Medium | Synopsis: The remote web server is affected by a denial of service vulnerability. Description: According to its banner, the version of Apache running on the remote host is 2.4.x prior to 2.4.35. It is, therefore, affected by the following vulnerability: - By sending continuous SETTINGS frames of maximum size an ongoing HTTP/2 connection could be kept busy and would never time out. This can be abused for a DoS on the server. This only affect a server that has enabled the h2 protocol. Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.35 or later. See Also: https://archive.apache.org/dist/httpd/CHANGES_2.4.35 https://httpd.apache.org/security/vulnerabilities_24.html#2.4.35 CVSS Base Score: 5.9 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:N/I:N/A:P Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.35 CPE: cpe:/a:apache:http_server cpe:/a:apache:httpd CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-09-25 Patch Publication Date: 2018-09-25 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2018-09-27 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 125781 |  | McAfee Agent 5.0.x < 5.0.6 HF1267994 / 5.5.x < 5.5.1 HF1267991 / 5.6.x < 5.6.1 UDP DoS (SB10272) | Upgrade to McAfee Agent version 5.0.6 HF1267994, 5.5.1 HF1267991, 5.6.1, or later. | Medium | Synopsis: A security management agent installed on the remote host is affected by a denial of service vulnerability. Description: The version of McAfee Agent, formerly McAfee ePolicy Orchestrator (ePO) Agent, installed on the remote host is 5.0.x prior to 5.0.6 HF1267994, 5.5.x prior to 5.5.1 HF1267991, or 5.6.x prior to 5.6.1. It is, therefore, affected by a denial of service vulnerability. An unauthenticated, remote attacker can exploit this issue, via specifically crafted UDP packets, to cause the application to stop responding. Solution: Upgrade to McAfee Agent version 5.0.6 HF1267994, 5.5.1 HF1267991, 5.6.1, or later. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10272 CVSS Base Score: 5.3 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Path : C:\Program Files (x86)\McAfee\Common Framework\ Installed version : 5.0.6.220 Fixed version : 5.0.6 HF1267994 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2019-02-14 Patch Publication Date: 2019-02-14 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Misc. | Repeat Vulnerability | 2019-06-07 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 126821 |  | Oracle Java SE 1.7.0_231 / 1.8.0_221 / 1.11.0_4 / 1.12.0_2 Multiple Vulnerabilities (Jul 2019 CPU) | Upgrade to Oracle JDK / JRE 12 Update 2 , 11 Update 4, 8 Update 221 / 7 Update 231 or later. If necessary, remove any affected versions. | Medium | Synopsis: The remote Windows host contains a programming platform that is affected by multiple vulnerabilities. Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 7 Update 231, 8 Update 221, 11 Update 4, or 12 Update 2. It is, therefore, affected by multiple vulnerabilities: - Unspecified vulnerabilities in the utilities and JCE subcomponents of Oracle Java SE, which could allow an unauthenticated remote attacker to cause a partial denial of service. (CVE-2019-2762, CVE-2019-2769, CVE-2019-2842) - An unspecified vulnerability in the security subcomponent of Oracle Java SE, which could allow an unauthenticated local attacker to gain unauthorized access to critical Java SE data. (CVE-2019-2745) - Unspecified vulnerabilities in the networking and security subcomponents of Oracle Java SE, which could allow an unauthenticated remote attacker to gain unauthorized access to Java SE data. Exploitation of this vulnerability requires user interaction. (CVE-2019-2766, CVE-2019-2786, CVE-2019-2818) - An unspecified vulnerability in the networking subcomponent of Oracle Java SE, which could allow an unauthenticated remote attacker unauthorized read, update, insert or delete access to Java SE data. (CVE-2019-2816) - An unspecified vulnerability in the JSSE subcomponent of Oracle Java SE, which could allow an unauthenticated, remote attacker to gain unauthorized access to critical Java SE data. Exploitation of this vulnerability requires user interaction. (CVE-2019-2821) - A use after free vulnerability exists in the libpng subcomponent of Oracle Java SE. An unauthenticated, remote attacker can exploit this to cause a complete denial of service condition in Java SE. Exploitation of this vulnerability requires user interaction. (CVE-2019-7317) Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Oracle JDK / JRE 12 Update 2 , 11 Update 4, 8 Update 221 / 7 Update 231 or later. If necessary, remove any affected versions. See Also: http://www.nessus.org/u?9aa2b901 CVSS Base Score: 5.3 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:P/A:N Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.7.0_231 / 1.8.0_221 / 1.11.0_4 / 1.12.0_2 CPE: cpe:/a:oracle:jre cpe:/a:oracle:jdk CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2019-07-16 Patch Publication Date: 2019-07-16 Plugin Modification Date: 2022-12-06 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2019-07-19 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 132101 |  | Windows Speculative Execution Configuration Check | Apply vendor recommended settings. | Medium | Synopsis: The remote host has not properly mitigated a series of speculative execution vulnerabilities. Description: The remote host has not properly mitigated a series of known speculative execution vulnerabilities. It, therefore, may be affected by : - Branch Target Injection (BTI) (CVE-2017-5715) - Bounds Check Bypass (BCB) (CVE-2017-5753) - Rogue Data Cache Load (RDCL) (CVE-2017-5754) - Rogue System Register Read (RSRE) (CVE-2018-3640) - Speculative Store Bypass (SSB) (CVE-2018-3639) - L1 Terminal Fault (L1TF) (CVE-2018-3615, CVE-2018-3620, CVE-2018-3646) - Microarchitectural Data Sampling Uncacheable Memory (MDSUM) (CVE-2019-11091) - Microarchitectural Store Buffer Data Sampling (MSBDS) (CVE-2018-12126) - Microarchitectural Load Port Data Sampling (MLPDS) (CVE-2018-12127) - Microarchitectural Fill Buffer Data Sampling (MFBDS) (CVE-2018-12130) - TSX Asynchronous Abort (TAA) (CVE-2019-11135) Solution: Apply vendor recommended settings. See Also: http://www.nessus.org/u?8902cebb http://www.nessus.org/u?6a005ed4 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:P/A:N Plugin Output: Current Settings: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: Not Set - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: Not Set ----------------------------------- Recommended Settings 1: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00000048 (72) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading enabled. ----------------------------------- Recommended Settings 2: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00002048 (8264) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading disabled. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-01-04 Patch Publication Date: NULL Plugin Modification Date: 2023-05-03 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2019-12-18 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 135290 |  | Apache 2.4.x < 2.4.42 Multiple Vulnerabilities | Upgrade to Apache version 2.4.42 or later. | Medium | Synopsis: The remote web server is affected by multiple vulnerabilities. Description: The version of Apache httpd installed on the remote host is prior to 2.4.42. It is, therefore, affected by multiple vulnerabilities as referenced in the 2.4.42 advisory. - In Apache HTTP Server 2.4.0 to 2.4.41, mod_proxy_ftp may use uninitialized memory when proxying to a malicious FTP server. (CVE-2020-1934) - In Apache HTTP Server 2.4.0 to 2.4.41, redirects configured with mod_rewrite that were intended to be self-referential might be fooled by encoded newlines and redirect instead to an unexpected URL within the request URL. (CVE-2020-1927) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Apache version 2.4.42 or later. See Also: NULL CVSS Base Score: 6.1 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:P/A:N Plugin Output: Path : C:\Program Files (x86)\Apache Software Foundation\Apache2.4\bin\httpd.exe Installed version : 2.4.10.0 Fixed version : 2.4.42 CPE: cpe:/a:apache:httpd cpe:/a:apache:http_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-04-01 Patch Publication Date: 2020-04-01 Plugin Modification Date: 2022-12-05 Exploitability Ease: No known exploits are available Plugin Type: combined Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp |  | Web Servers | Repeat Vulnerability | 2020-04-10 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 141800 |  | Oracle Java SE 1.7.0_281 / 1.8.0_271 / 1.11.0_9 / 1.15.0_1 Multiple Vulnerabilities (Oct 2020 CPU) | Apply the appropriate patch according to the October 2020 Oracle Critical Patch Update advisory. | Medium | Synopsis: The remote host is affected by multiple vulnerabilities Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 7 Update 281, 8 Update 271, 11 Update 9, or 15 Update 1. It is, therefore, affected by multiple vulnerabilities related to the following components as referenced in the October 2020 CPU advisory: - Vulnerability in the Oracle GraalVM Enterprise Edition product of Oracle GraalVM (component: Java). Supported versions that are affected are 19.3.3 and 20.2.0. Easily exploitable vulnerability allows unauthenticated attacker with network access via multiple protocols to compromise Oracle GraalVM Enterprise Edition. Successful attacks of this vulnerability can result in unauthorized read access to a subset of Oracle GraalVM Enterprise Edition accessible data. (CVE-2020-14803) - Vulnerability in the Java SE, Java SE Embedded product of Oracle Java SE (component: Hotspot). Supported versions that are affected are Java SE: 7u271, 8u261, 11.0.8 and 15; Java SE Embedded: 8u261. Difficult to exploit vulnerability allows unauthenticated attacker with network access via multiple protocols to compromise Java SE, Java SE Embedded. Successful attacks require human interaction from a person other than the attacker. Successful attacks of this vulnerability can result in unauthorized update, insert or delete access to some of Java SE, Java SE Embedded accessible data as well as unauthorized read access to a subset of Java SE, Java SE Embedded accessible data. Note: Applies to client and server deployment of Java. This vulnerability can be exploited through sandboxed Java Web Start applications and sandboxed Java applets. It can also be exploited by supplying data to APIs in the specified Component without using sandboxed Java Web Start applications or sandboxed Java applets, such as through a web service. (CVE-2020-14792) - Vulnerability in the Java SE, Java SE Embedded product of Oracle Java SE (component: JNDI). Supported versions that are affected are Java SE: 7u271, 8u261, 11.0.8 and 15; Java SE Embedded: 8u261. Difficult to exploit vulnerability allows unauthenticated attacker with network access via multiple protocols to compromise Java SE, Java SE Embedded. Successful attacks of this vulnerability can result in unauthorized read access to a subset of Java SE, Java SE Embedded accessible data. Note: Applies to client and server deployment of Java. This vulnerability can be exploited through sandboxed Java Web Start applications and sandboxed Java applets. It can also be exploited by supplying data to APIs in the specified Component without using sandboxed Java Web Start applications or sandboxed Java applets, such as through a web service. (CVE-2020-14781) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Apply the appropriate patch according to the October 2020 Oracle Critical Patch Update advisory. See Also: https://www.oracle.com/a/tech/docs/cpuoct2020cvrf.xml https://www.oracle.com/security-alerts/cpuoct2020.html CVSS Base Score: 5.3 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:P/A:N Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.7.0_281 / 1.8.0_271 / 1.11.0_9 / 1.15.0_1 CPE: cpe:/a:oracle:jre cpe:/a:oracle:jdk CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-10-20 Patch Publication Date: 2020-10-20 Plugin Modification Date: 2022-12-05 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2020-10-22 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 144450 |  | McAfee VirusScan Enterprise < 8.8 Patch 16 Access Control Bypass Vulnerability (SB10338) | Upgrade to McAfee VirusScan Enterprise version 8.8 Patch 16. | Medium | Synopsis: The antivirus application installed on the remote Windows host is affected by an access control bypass vulnerability. Description: The version of McAfee VirusScan Enterprise (VSE) installed on the remote Windows host is prior to 8.8 Patch 16. It is, therefore, affected by an access control bypass vulnerability which allows local administrators the ability to bypass local security protections by carefully manipulating the code integrity checks associated with the windows defender integration. Solution: Upgrade to McAfee VirusScan Enterprise version 8.8 Patch 16. See Also: http://www.nessus.org/u?0ef336c3 CVSS Base Score: 6.7 CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Installed version : 8.8.0.2024 Fixed version : 8.8.0.2300 CPE: cpe:/a:mcafee:virusscan_enterprise CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-12-08 Patch Publication Date: 2020-12-08 Plugin Modification Date: 2021-01-19 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2020-12-18 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 145218 |  | Oracle Java SE 1.7.0_291 / 1.8.0_281 / 1.11.0_10 / 1.15.0_2 Information Disclosure (Windows Jan 2021 CPU) | Apply the appropriate patch according to the January 2021 Oracle Critical Patch Update advisory. | Medium | Synopsis: The remote host is affected by an information disclosure vulnerability. Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 7 Update 291, 8 Update 281, 11 Update 10, or 15 Update 2. It is, therefore, affected by an information disclosure vulnerability as referenced in the January 2021 CPU advisory. Specifically, an unauthenticated, remote attacker can gain unauthorized read access to some data accessible to Java SE and Java SE Embedded. Only Java deployments that load and run untrusted code and rely on the Java sandbox for security are affected. This vulnerability does not apply to Java deployments, typically in servers, that load and run only trusted code (e.g., code installed by an administrator). Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Apply the appropriate patch according to the January 2021 Oracle Critical Patch Update advisory. See Also: https://www.oracle.com/a/tech/docs/cpujan2021cvrf.xml https://www.oracle.com/security-alerts/cpujan2021.html#AppendixJAVA CVSS Base Score: 5.3 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:N/A:N Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.7.0_291 / 1.8.0_281 / 1.11.0_10 / 1.15.0_2 CPE: cpe:/a:oracle:jre cpe:/a:oracle:jdk CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-10-21 Patch Publication Date: 2021-01-19 Plugin Modification Date: 2022-12-07 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2021-01-20 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 148960 |  | Oracle Java SE 1.7.0_301 / 1.8.0_291 / 1.11.0_11 / 1.16.0_1 Multiple Vulnerabilities (Apr 2021 CPU) | Apply the appropriate patch according to the April 2021 Oracle Critical Patch Update advisory. | Medium | Synopsis: The remote host is affected by multiple vulnerabilities. Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is prior to 7 Update 301, 8 Update 291, 11 Update 11, or 16 Update 1. It is, therefore, affected by multiple vulnerabilities as referenced in the April 2021 CPU advisory: - A vulnerability in Java SE, SE Embedded and Oracle GraalVM Enterprise Edition allows unauthenticated remote attacker to compromise the system which can result in an unauthorized creation, deletion or modification access to critical data. (CVE-2021-2161) - A vulnerability in Java SE, SE Embedded and Oracle GraalVM Enterprise Edition allows unauthenticated remote attacker with a human interaction from a person other than the attacker to compromise the system which can result in an unauthorized creation, deletion or modification access to critical data. (CVE-2021-2163) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Apply the appropriate patch according to the April 2021 Oracle Critical Patch Update advisory. See Also: https://www.oracle.com/a/tech/docs/cpuapr2021cvrf.xml https://www.oracle.com/security-alerts/cpuapr2021.html#AppendixJAVA CVSS Base Score: 5.9 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:N/I:P/A:N Plugin Output: The following vulnerable instance of Java is installed on the remote host : Path : C:\Program Files\Java\jre1.8.0_101 Installed version : 1.8.0_101 Fixed version : 1.7.0_301 / 1.8.0_291 / 1.11.0_11 / 1.16.0_1 CPE: cpe:/a:oracle:jre cpe:/a:oracle:jdk CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-04-20 Patch Publication Date: 2021-04-20 Plugin Modification Date: 2022-12-05 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2021-04-23 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 156887 |  | Oracle Java SE 1.7.0_331 / 1.8.0_321 / 1.11.0_14 / 1.17.0_2 Multiple Vulnerabilities (January 2022 CPU) | Apply the appropriate patch according to the January 2022 Oracle Critical Patch Update advisory. | Medium | Synopsis: The remote host is affected by multiple vulnerabilities. Description: The version of Oracle (formerly Sun) Java SE or Java for Business installed on the remote host is affected by multiple vulnerabilities as referenced in the January 2022 CPU advisory: - Vulnerability in the Oracle Java SE, Oracle GraalVM Enterprise Edition product of Oracle Java SE (component: 2D). Supported versions that are affected are Oracle Java SE: 7u321, 8u311; Oracle GraalVM Enterprise Edition: 20.3.4 and 21.3.0. Easily exploitable vulnerability allows unauthenticated attacker with network access via multiple protocols to compromise Oracle Java SE, Oracle GraalVM Enterprise Edition. Successful attacks of this vulnerability can result in unauthorized ability to cause a partial denial of service (partial DOS) of Oracle Java SE, Oracle GraalVM Enterprise Edition. Note: This vulnerability applies to Java deployments, typically in clients running sandboxed Java Web Start applications or sandboxed Java applets, that load and run untrusted code (e.g., code that comes from the internet) and rely on the Java sandbox for security. This vulnerability can also be exploited by using APIs in the specified Component, e.g., through a web service which supplies data to the APIs. (CVE-2022-21349) - Vulnerability in the Oracle Java SE, Oracle GraalVM Enterprise Edition product of Oracle Java SE (component: Hotspot). Supported versions that are affected are Oracle Java SE: 7u321, 8u311, 11.0.13, 17.01; Oracle GraalVM Enterprise Edition: 20.3.4 and 21.3.0. Easily exploitable vulnerability allows unauthenticated attacker with network access via multiple protocols to compromise Oracle Java SE, Oracle GraalVM Enterprise Edition. Successful attacks of this vulnerability can result in unauthorized update, insert or delete access to some of Oracle Java SE, Oracle GraalVM Enterprise Edition accessible data. Note: This vulnerability applies to Java deployments, typically in clients running sandboxed Java Web Start applications or sandboxed Java applets, that load and run untrusted code (e.g., code that comes from the internet) and rely on the Java sandbox for security. This vulnerability can also be exploited by using APIs in the specified Component, e.g., through a web service which supplies data to the APIs. (CVE-2022-21291) - Vulnerability in the Oracle Java SE, Oracle GraalVM Enterprise Edition product of Oracle Java SE (component: Hotspot). Supported versions that are affected are Oracle Java SE: 7u321, 8u311, 11.0.13, 17.01; Oracle GraalVM Enterprise Edition: 20.3.4 and 21.3.0. Easily exploitable vulnerability allows unauthenticated attacker with network access via multiple protocols to compromise Oracle Java SE, Oracle GraalVM Enterprise Edition. Successful attacks of this vulnerability can result in unauthorized update, insert or delete access to some of Oracle Java SE, Oracle GraalVM Enterprise Edition accessible data. Note: This vulnerability applies to Java deployments, typically in clients running sandboxed Java Web Start applications or sandboxed Java applets, that load and run untrusted code (e.g., code that comes from the internet) and rely on the Java sandbox for security. This vulnerability can also be exploited by using APIs in the specified Component, e.g., through a web service which supplies data to the APIs. (CVE-2022-21305) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Apply the appropriate patch according to the January 2022 Oracle Critical Patch Update advisory. See Also: https://www.oracle.com/a/tech/docs/cpujan2022cvrf.xml https://www.oracle.com/security-alerts/cpujan2022.html#AppendixJAVA CVSS Base Score: 5.3 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:P/A:N Plugin Output: The following vulnerable instance of Java is installed on the remot | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2022-01-20 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 169456 |  | Trellix Agent < 5.7.8 Privilege Escalation (SB10391) | Upgrade to McAfee Agent version 5.7.8 or later. | Medium | Synopsis: A security management agent installed on the remote host is affected by a privilege escalation vulnerability. Description: The version of Trellix Agent, formerly McAfee Agent or McAfee Policy Orchestrator (ePO) Agent, installed on the remote host is prior to 5.7.8. It is, therefore, affected by an privilege escalation vulnerability due to an uncontrolled search path. An attacker with admin access to the affected host can place a malicious DLL in the restricted Windows System folder to elevate their privileges to System. Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to McAfee Agent version 5.7.8 or later. See Also: https://kcm.trellix.com/corporate/index?page=content&id=SB10391 CVSS Base Score: 6.7 CVSS Vector: CVSS2#AV:L/AC:L/Au:M/C:C/I:C/A:C Plugin Output: Path : C:\Program Files (x86)\McAfee\Common Framework\ Installed version : 5.0.6.220 Fixed version : 5.7.8 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent cpe:/a:trellix:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-11-29 Patch Publication Date: 2022-11-29 Plugin Modification Date: 2023-05-04 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2023-01-03 | 2023-01-09 04:57 | 2023-05-09 07:17 | 4
|  |  | 132101 |  | Windows Speculative Execution Configuration Check | Apply vendor recommended settings. | Medium | Synopsis: The remote host has not properly mitigated a series of speculative execution vulnerabilities. Description: The remote host has not properly mitigated a series of known speculative execution vulnerabilities. It, therefore, may be affected by : - Branch Target Injection (BTI) (CVE-2017-5715) - Bounds Check Bypass (BCB) (CVE-2017-5753) - Rogue Data Cache Load (RDCL) (CVE-2017-5754) - Rogue System Register Read (RSRE) (CVE-2018-3640) - Speculative Store Bypass (SSB) (CVE-2018-3639) - L1 Terminal Fault (L1TF) (CVE-2018-3615, CVE-2018-3620, CVE-2018-3646) - Microarchitectural Data Sampling Uncacheable Memory (MDSUM) (CVE-2019-11091) - Microarchitectural Store Buffer Data Sampling (MSBDS) (CVE-2018-12126) - Microarchitectural Load Port Data Sampling (MLPDS) (CVE-2018-12127) - Microarchitectural Fill Buffer Data Sampling (MFBDS) (CVE-2018-12130) - TSX Asynchronous Abort (TAA) (CVE-2019-11135) Solution: Apply vendor recommended settings. See Also: http://www.nessus.org/u?8902cebb http://www.nessus.org/u?6a005ed4 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:P/A:N Plugin Output: Current Settings: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: Not Set - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: Not Set ----------------------------------- Recommended Settings 1: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00000048 (72) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading enabled. ----------------------------------- Recommended Settings 2: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00002048 (8264) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading disabled. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-01-04 Patch Publication Date: NULL Plugin Modification Date: 2023-05-03 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Compass Router - Prod | LAUREL BRIDGE SOFTWARE INC |  |  | Region X | VISN X | VAMC | VHAXXXAPPCROUTP |  |  |  |  | VHAXXXAPPCROUTP.vxx.med.va.gov |  | Microsoft Windows Server 2019 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2019-12-18 | 2021-11-07 02:01 | 2023-05-08 00:45 | 18
|  |  | 132101 |  | Windows Speculative Execution Configuration Check | Apply vendor recommended settings. | Medium | Synopsis: The remote host has not properly mitigated a series of speculative execution vulnerabilities. Description: The remote host has not properly mitigated a series of known speculative execution vulnerabilities. It, therefore, may be affected by : - Branch Target Injection (BTI) (CVE-2017-5715) - Bounds Check Bypass (BCB) (CVE-2017-5753) - Rogue Data Cache Load (RDCL) (CVE-2017-5754) - Rogue System Register Read (RSRE) (CVE-2018-3640) - Speculative Store Bypass (SSB) (CVE-2018-3639) - L1 Terminal Fault (L1TF) (CVE-2018-3615, CVE-2018-3620, CVE-2018-3646) - Microarchitectural Data Sampling Uncacheable Memory (MDSUM) (CVE-2019-11091) - Microarchitectural Store Buffer Data Sampling (MSBDS) (CVE-2018-12126) - Microarchitectural Load Port Data Sampling (MLPDS) (CVE-2018-12127) - Microarchitectural Fill Buffer Data Sampling (MFBDS) (CVE-2018-12130) - TSX Asynchronous Abort (TAA) (CVE-2019-11135) Solution: Apply vendor recommended settings. See Also: http://www.nessus.org/u?8902cebb http://www.nessus.org/u?6a005ed4 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:P/A:N Plugin Output: Current Settings: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: Not Set - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: Not Set ----------------------------------- Recommended Settings 1: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00000048 (72) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading enabled. ----------------------------------- Recommended Settings 2: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00002048 (8264) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading disabled. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-01-04 Patch Publication Date: NULL Plugin Modification Date: 2023-05-03 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Compass Router - Hot Spare | LAUREL BRIDGE SOFTWARE INC |  |  | Region X | VISN X | VAMC | VHAXXXAPPCROUTS |  |  |  |  | VHAXXXAPPCROUTS.vxx.med.va.gov |  | Microsoft Windows Server 2019 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2019-12-18 | 2021-11-07 02:01 | 2023-05-08 00:46 | 18
|  |  | 132101 |  | Windows Speculative Execution Configuration Check | Apply vendor recommended settings. | Medium | Synopsis: The remote host has not properly mitigated a series of speculative execution vulnerabilities. Description: The remote host has not properly mitigated a series of known speculative execution vulnerabilities. It, therefore, may be affected by : - Branch Target Injection (BTI) (CVE-2017-5715) - Bounds Check Bypass (BCB) (CVE-2017-5753) - Rogue Data Cache Load (RDCL) (CVE-2017-5754) - Rogue System Register Read (RSRE) (CVE-2018-3640) - Speculative Store Bypass (SSB) (CVE-2018-3639) - L1 Terminal Fault (L1TF) (CVE-2018-3615, CVE-2018-3620, CVE-2018-3646) - Microarchitectural Data Sampling Uncacheable Memory (MDSUM) (CVE-2019-11091) - Microarchitectural Store Buffer Data Sampling (MSBDS) (CVE-2018-12126) - Microarchitectural Load Port Data Sampling (MLPDS) (CVE-2018-12127) - Microarchitectural Fill Buffer Data Sampling (MFBDS) (CVE-2018-12130) - TSX Asynchronous Abort (TAA) (CVE-2019-11135) Solution: Apply vendor recommended settings. See Also: http://www.nessus.org/u?8902cebb http://www.nessus.org/u?6a005ed4 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:P/A:N Plugin Output: Current Settings: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: Not Set - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: Not Set ----------------------------------- Recommended Settings 1: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00000048 (72) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading enabled. ----------------------------------- Recommended Settings 2: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00002048 (8264) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading disabled. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-01-04 Patch Publication Date: NULL Plugin Modification Date: 2023-05-03 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Compass Router - Test | LAUREL BRIDGE SOFTWARE INC |  |  | Region X | VISN X | VAMC | VHAXXXAPPCROUTT |  |  |  |  | VHAXXXAPPCROUTT.vxx.med.va.gov |  | Microsoft Windows Server 2019 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2019-12-18 | 2021-11-07 02:01 | 2023-05-08 00:44 | 18
|  |  | 125781 |  | McAfee Agent 5.0.x < 5.0.6 HF1267994 / 5.5.x < 5.5.1 HF1267991 / 5.6.x < 5.6.1 UDP DoS (SB10272) | Upgrade to McAfee Agent version 5.0.6 HF1267994, 5.5.1 HF1267991, 5.6.1, or later. | Medium | Synopsis: A security management agent installed on the remote host is affected by a denial of service vulnerability. Description: The version of McAfee Agent, formerly McAfee ePolicy Orchestrator (ePO) Agent, installed on the remote host is 5.0.x prior to 5.0.6 HF1267994, 5.5.x prior to 5.5.1 HF1267991, or 5.6.x prior to 5.6.1. It is, therefore, affected by a denial of service vulnerability. An unauthenticated, remote attacker can exploit this issue, via specifically crafted UDP packets, to cause the application to stop responding. Solution: Upgrade to McAfee Agent version 5.0.6 HF1267994, 5.5.1 HF1267991, 5.6.1, or later. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10272 CVSS Base Score: 5.3 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Path : C:\Program Files\McAfee\Agent\ Installed version : 5.5.1.342 Fixed version : 5.5.1 HF1267991 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2019-02-14 Patch Publication Date: 2019-02-14 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Misc. | Repeat Vulnerability | 2019-06-07 | 2022-02-06 21:50 | 2023-05-09 07:10 | 15
|  |  | 132101 |  | Windows Speculative Execution Configuration Check | Apply vendor recommended settings. | Medium | Synopsis: The remote host has not properly mitigated a series of speculative execution vulnerabilities. Description: The remote host has not properly mitigated a series of known speculative execution vulnerabilities. It, therefore, may be affected by : - Branch Target Injection (BTI) (CVE-2017-5715) - Bounds Check Bypass (BCB) (CVE-2017-5753) - Rogue Data Cache Load (RDCL) (CVE-2017-5754) - Rogue System Register Read (RSRE) (CVE-2018-3640) - Speculative Store Bypass (SSB) (CVE-2018-3639) - L1 Terminal Fault (L1TF) (CVE-2018-3615, CVE-2018-3620, CVE-2018-3646) - Microarchitectural Data Sampling Uncacheable Memory (MDSUM) (CVE-2019-11091) - Microarchitectural Store Buffer Data Sampling (MSBDS) (CVE-2018-12126) - Microarchitectural Load Port Data Sampling (MLPDS) (CVE-2018-12127) - Microarchitectural Fill Buffer Data Sampling (MFBDS) (CVE-2018-12130) - TSX Asynchronous Abort (TAA) (CVE-2019-11135) Solution: Apply vendor recommended settings. See Also: http://www.nessus.org/u?8902cebb http://www.nessus.org/u?6a005ed4 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:P/A:N Plugin Output: Current Settings: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: Not Set - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: Not Set ----------------------------------- Recommended Settings 1: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00000048 (72) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading enabled. ----------------------------------- Recommended Settings 2: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00002048 (8264) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading disabled. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-01-04 Patch Publication Date: NULL Plugin Modification Date: 2023-05-03 Exploitability Ease: Exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2019-12-18 | 2022-02-06 21:50 | 2023-05-09 07:10 | 15
|  |  | 134762 |  | McAfee Agent 5.5.x prior to 5.5.4 / 5.6.x prior to 5.6.4 Improper Access Control Vulnerability (SB10312) | Upgrade to McAfee Agent version 5.5.4, 5.6.4 or later. | Medium | Synopsis: A security management agent installed on the remote host is affected by improper access control vulnerability. Description: The version of McAfee Agent, formerly McAfee ePolicy Orchestrator (ePO) Agent, installed on the remote host is 5.5.x prior to 5.5.4 / 5.6.x prior to 5.6.4. It is, therefore, affected by improper access control vulnerability. It allows local users with administrator privileges to disable self-protection via a McAfee supplied command-line utility. Solution: Upgrade to McAfee Agent version 5.5.4, 5.6.4 or later. See Also: https://kc.mcafee.com/corporate/index?page=content&id=SB10312 CVSS Base Score: 4.4 CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:N/I:P/A:N Plugin Output: Path : C:\Program Files\McAfee\Agent\ Installed version : 5.5.1.342 Fixed version : 5.5.4 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-03-10 Patch Publication Date: 2020-03-10 Plugin Modification Date: 2021-09-27 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2020-03-20 | 2022-09-04 23:24 | 2023-05-09 07:10 | 8
|  |  | 169456 |  | Trellix Agent < 5.7.8 Privilege Escalation (SB10391) | Upgrade to McAfee Agent version 5.7.8 or later. | Medium | Synopsis: A security management agent installed on the remote host is affected by a privilege escalation vulnerability. Description: The version of Trellix Agent, formerly McAfee Agent or McAfee Policy Orchestrator (ePO) Agent, installed on the remote host is prior to 5.7.8. It is, therefore, affected by an privilege escalation vulnerability due to an uncontrolled search path. An attacker with admin access to the affected host can place a malicious DLL in the restricted Windows System folder to elevate their privileges to System. Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to McAfee Agent version 5.7.8 or later. See Also: https://kcm.trellix.com/corporate/index?page=content&id=SB10391 CVSS Base Score: 6.7 CVSS Vector: CVSS2#AV:L/AC:L/Au:M/C:C/I:C/A:C Plugin Output: Path : C:\Program Files\McAfee\Agent\ Installed version : 5.5.1.342 Fixed version : 5.7.8 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent cpe:/a:trellix:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-11-29 Patch Publication Date: 2022-11-29 Plugin Modification Date: 2023-05-04 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2023-01-03 | 2023-01-09 04:57 | 2023-05-09 07:10 | 4
|  |  | 169783 |  | Security Updates for Windows Malicious Software Removal Tool (January 2023) | Microsoft has released version 5.109 to address this issue. | Medium | Synopsis: The remote Windows host has an antimalware application that is affected by a privilege escalation vulnerability. Description: The Windows Malicious Software Removal Tool installation on the remote host is missing a security update. It is, therefore, affected by the following vulnerability: - An elevation of privilege vulnerability. An attacker can exploit this to gain elevated privileges. (CVE-2023-21725) Solution: Microsoft has released version 5.109 to address this issue. See Also: http://www.nessus.org/u?867b0b4e CVSS Base Score: 6.3 CVSS Vector: CVSS2#AV:L/AC:H/Au:S/C:N/I:C/A:C Plugin Output: Product : Microsoft Malicious Software Removal Tool Installed version : 5.105.19638.1 Fixed version : 5.109.19957.1 CPE: cpe:/o:microsoft:windows cpe:/a:microsoft:malware_protection_engine CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-01-10 Patch Publication Date: 2023-01-10 Plugin Modification Date: 2023-01-11 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2023-01-10 | 2023-02-06 05:03 | 2023-05-09 07:10 | 3
|  |  | 132101 |  | Windows Speculative Execution Configuration Check | Apply vendor recommended settings. | Medium | Synopsis: The remote host has not properly mitigated a series of speculative execution vulnerabilities. Description: The remote host has not properly mitigated a series of known speculative execution vulnerabilities. It, therefore, may be affected by : - Branch Target Injection (BTI) (CVE-2017-5715) - Bounds Check Bypass (BCB) (CVE-2017-5753) - Rogue Data Cache Load (RDCL) (CVE-2017-5754) - Rogue System Register Read (RSRE) (CVE-2018-3640) - Speculative Store Bypass (SSB) (CVE-2018-3639) - L1 Terminal Fault (L1TF) (CVE-2018-3615, CVE-2018-3620, CVE-2018-3646) - Microarchitectural Data Sampling Uncacheable Memory (MDSUM) (CVE-2019-11091) - Microarchitectural Store Buffer Data Sampling (MSBDS) (CVE-2018-12126) - Microarchitectural Load Port Data Sampling (MLPDS) (CVE-2018-12127) - Microarchitectural Fill Buffer Data Sampling (MFBDS) (CVE-2018-12130) - TSX Asynchronous Abort (TAA) (CVE-2019-11135) Solution: Apply vendor recommended settings. See Also: http://www.nessus.org/u?8902cebb http://www.nessus.org/u?6a005ed4 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:P/A:N Plugin Output: Current Settings: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: Not Set - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: Not Set ----------------------------------- Recommended Settings 1: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00000048 (72) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading enabled. ----------------------------------- Recommended Settings 2: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00002048 (8264) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading disabled. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-01-04 Patch Publication Date: NULL Plugin Modification Date: 2023-05-03 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Report Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE2 |  |  |  |  | VHAXXXAPPROCHE2.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2019-12-18 | 2022-02-06 21:50 | 2023-05-09 07:17 | 15
|  |  | 169783 |  | Security Updates for Windows Malicious Software Removal Tool (January 2023) | Microsoft has released version 5.109 to address this issue. | Medium | Synopsis: The remote Windows host has an antimalware application that is affected by a privilege escalation vulnerability. Description: The Windows Malicious Software Removal Tool installation on the remote host is missing a security update. It is, therefore, affected by the following vulnerability: - An elevation of privilege vulnerability. An attacker can exploit this to gain elevated privileges. (CVE-2023-21725) Solution: Microsoft has released version 5.109 to address this issue. See Also: http://www.nessus.org/u?867b0b4e CVSS Base Score: 6.3 CVSS Vector: CVSS2#AV:L/AC:H/Au:S/C:N/I:C/A:C Plugin Output: Product : Microsoft Malicious Software Removal Tool Installed version : 5.105.19638.1 Fixed version : 5.109.19957.1 CPE: cpe:/o:microsoft:windows cpe:/a:microsoft:malware_protection_engine CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-01-10 Patch Publication Date: 2023-01-10 Plugin Modification Date: 2023-01-11 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Report Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE2 |  |  |  |  | VHAXXXAPPROCHE2.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2023-01-10 | 2023-02-06 05:03 | 2023-05-09 07:17 | 3
|  |  | 132101 |  | Windows Speculative Execution Configuration Check | Apply vendor recommended settings. | Medium | Synopsis: The remote host has not properly mitigated a series of speculative execution vulnerabilities. Description: The remote host has not properly mitigated a series of known speculative execution vulnerabilities. It, therefore, may be affected by : - Branch Target Injection (BTI) (CVE-2017-5715) - Bounds Check Bypass (BCB) (CVE-2017-5753) - Rogue Data Cache Load (RDCL) (CVE-2017-5754) - Rogue System Register Read (RSRE) (CVE-2018-3640) - Speculative Store Bypass (SSB) (CVE-2018-3639) - L1 Terminal Fault (L1TF) (CVE-2018-3615, CVE-2018-3620, CVE-2018-3646) - Microarchitectural Data Sampling Uncacheable Memory (MDSUM) (CVE-2019-11091) - Microarchitectural Store Buffer Data Sampling (MSBDS) (CVE-2018-12126) - Microarchitectural Load Port Data Sampling (MLPDS) (CVE-2018-12127) - Microarchitectural Fill Buffer Data Sampling (MFBDS) (CVE-2018-12130) - TSX Asynchronous Abort (TAA) (CVE-2019-11135) Solution: Apply vendor recommended settings. See Also: http://www.nessus.org/u?8902cebb http://www.nessus.org/u?6a005ed4 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:P/A:N Plugin Output: Current Settings: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: Not Set - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: Not Set ----------------------------------- Recommended Settings 1: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00000048 (72) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading enabled. ----------------------------------- Recommended Settings 2: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00002048 (8264) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading disabled. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-01-04 Patch Publication Date: NULL Plugin Modification Date: 2023-05-03 Exploitability Ease: Exploits are available Plugin Type: local Source File: | COBAS Infinity Production Server (v3.03) | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXCCACOBASP |  |  |  |  | VHAXXXCCACOBASP.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2019-12-18 | 2022-08-07 23:57 | 2023-05-09 07:18 | 9
|  |  | 136946 |  | Windows 10 / Windows Server 2016 September 2017 Information Disclosure Vulnerability (CVE-2017-8529) | Refer to the Microsoft CVE article for additional information. | Medium | Synopsis: The remote Windows host is affected by an information disclosure vulnerability. Description: The remote Windows host is missing a security update or a registry setting required to enable protections for CVE-2017-8529. It is, therefore, affected by an information disclosure vulnerability: - An information disclosure vulnerability exists when affected Microsoft scripting engines do not properly handle objects in memory. The vulnerability could allow an attacker to detect specific files on the user's computer. In a web-based attack scenario, an attacker could host a website that is used to attempt to exploit the vulnerability. Solution: Refer to the Microsoft CVE article for additional information. See Also: http://www.nessus.org/u?1f6a3c24 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:N/A:N Plugin Output: The following registry key is required to enable the fix for CVE-2017-8529 and is missing. HKLM\SOFTWARE\Microsoft\Internet Explorer\Main\FeatureControl\FEATURE_ENABLE_PRINT_INFO_DISCLOSURE_FIX\iexplore.exe The following registry key is required to enable the fix for CVE-2017-8529 and is missing. HKLM\SOFTWARE\WOW6432Node\Microsoft\Internet Explorer\Main\FeatureControl\FEATURE_ENABLE_PRINT_INFO_DISCLOSURE_FIX\iexplore.exe CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2017-06-13 Patch Publication Date: 2017-06-13 Plugin Modification Date: 2020-06-04 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | COBAS Infinity Production Server (v3.03) | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXCCACOBASP |  |  |  |  | VHAXXXCCACOBASP.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2020-05-28 | 2022-08-07 23:57 | 2023-05-09 07:18 | 9
|  |  | 132101 |  | Windows Speculative Execution Configuration Check | Apply vendor recommended settings. | Medium | Synopsis: The remote host has not properly mitigated a series of speculative execution vulnerabilities. Description: The remote host has not properly mitigated a series of known speculative execution vulnerabilities. It, therefore, may be affected by : - Branch Target Injection (BTI) (CVE-2017-5715) - Bounds Check Bypass (BCB) (CVE-2017-5753) - Rogue Data Cache Load (RDCL) (CVE-2017-5754) - Rogue System Register Read (RSRE) (CVE-2018-3640) - Speculative Store Bypass (SSB) (CVE-2018-3639) - L1 Terminal Fault (L1TF) (CVE-2018-3615, CVE-2018-3620, CVE-2018-3646) - Microarchitectural Data Sampling Uncacheable Memory (MDSUM) (CVE-2019-11091) - Microarchitectural Store Buffer Data Sampling (MSBDS) (CVE-2018-12126) - Microarchitectural Load Port Data Sampling (MLPDS) (CVE-2018-12127) - Microarchitectural Fill Buffer Data Sampling (MFBDS) (CVE-2018-12130) - TSX Asynchronous Abort (TAA) (CVE-2019-11135) Solution: Apply vendor recommended settings. See Also: http://www.nessus.org/u?8902cebb http://www.nessus.org/u?6a005ed4 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:P/A:N Plugin Output: Current Settings: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: Not Set - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: Not Set ----------------------------------- Recommended Settings 1: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00000048 (72) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading enabled. ----------------------------------- Recommended Settings 2: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00002048 (8264) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading disabled. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-01-04 Patch Publication Date: NULL Plugin Modification Date: 2023-05-03 Exploitability Ease: Exploits are available Plugin Type: local Source File: | VIRTUAL SERVER |  |  |  | Region X | VISN X | VAMC | VHAXXXCCAMIM |  |  |  |  | VHAXXXCCAMIM.vxx.med.va.gov |  | Microsoft Windows Server 2019 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2019-12-18 | 2022-04-02 23:19 | 2023-05-07 06:50 | 13
|  |  | 42873 |  | SSL Medium Strength Cipher Suites Supported (SWEET32) | Reconfigure the affected application if possible to avoid use of medium strength ciphers. | High | Synopsis: The remote service supports the use of medium strength SSL ciphers. Description: The remote host supports the use of SSL ciphers that offer medium strength encryption. Nessus regards medium strength as any encryption that uses key lengths at least 64 bits and less than 112 bits, or else that uses the 3DES encryption suite. Note that it is considerably easier to circumvent medium strength encryption if the attacker is on the same physical network. Solution: Reconfigure the affected application if possible to avoid use of medium strength ciphers. See Also: https://www.openssl.org/blog/blog/2016/08/24/sweet32/ https://sweet32.info CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:N/A:N Plugin Output: Medium Strength Ciphers (> 64-bit and < 112-bit key, or 3DES) Name Code KEX Auth Encryption MAC ---------------------- ---------- --- ---- --------------------- --- DES-CBC3-SHA 0x00, 0x0A RSA RSA 3DES-CBC(168) SHA1 The fields above are : {Tenable ciphername} {Cipher ID code} Kex={key exchange} Auth={authentication} Encrypt={symmetric encryption method} MAC={message authentication code} {export flag} CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2016-08-24 Patch Publication Date: NULL Plugin Modification Date: 2021-02-03 Exploitability Ease: NULL Plugin Type: remote Source File: | Nexodose Server | BRACCO DIAGNOSTICS INC (US) |  |  | Region X | VISN X | VAMC | VHAXXXCCANEXOPR |  |  |  |  | VHAXXXCCANEXOPR.vxx.med.va.gov |  | Microsoft Windows 10 | tcp | 3389 | General | Repeat Vulnerability | 2009-11-23 | 2022-11-07 22:23 | 2023-05-09 02:05 | 6
|  |  | 51192 |  | SSL Certificate Cannot Be Trusted | Purchase or generate a proper SSL certificate for this service. | Medium | Synopsis: The SSL certificate for this service cannot be trusted. Description: The server's X.509 certificate cannot be trusted. This situation can occur in three different ways, in which the chain of trust can be broken, as stated below : - First, the top of the certificate chain sent by the server might not be descended from a known public certificate authority. This can occur either when the top of the chain is an unrecognized, self-signed certificate, or when intermediate certificates are missing that would connect the top of the certificate chain to a known public certificate authority. - Second, the certificate chain may contain a certificate that is not valid at the time of the scan. This can occur either when the scan occurs before one of the certificate's 'notBefore' dates, or after one of the certificate's 'notAfter' dates. - Third, the certificate chain may contain a signature that either didn't match the certificate's information or could not be verified. Bad signatures can be fixed by getting the certificate with the bad signature to be re-signed by its issuer. Signatures that could not be verified are the result of the certificate's issuer using a signing algorithm that Nessus either does not support or does not recognize. If the remote host is a public host in production, any break in the chain makes it more difficult for users to verify the authenticity and identity of the web server. This could make it easier to carry out man-in-the-middle attacks against the remote host. Solution: Purchase or generate a proper SSL certificate for this service. See Also: https://www.itu.int/rec/T-REC-X.509/en https://en.wikipedia.org/wiki/X.509 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:N Plugin Output: The following certificate was at the top of the certificate chain sent by the remote host, but it is signed by an unknown certificate authority : ¦-Subject : CN=VHADETCCANEXOPR.v11.med.va.gov ¦-Issuer : CN=VHADETCCANEXOPR.v11.med.va.gov CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2020-04-27 Exploitability Ease: NULL Plugin Type: remote Source File: | Nexodose Server | BRACCO DIAGNOSTICS INC (US) |  |  | Region X | VISN X | VAMC | VHAXXXCCANEXOPR |  |  |  |  | VHAXXXCCANEXOPR.vxx.med.va.gov |  | Microsoft Windows 10 | tcp | 3389 | General | Repeat Vulnerability | 2010-12-15 | 2022-11-07 22:23 | 2023-05-09 02:05 | 6
|  |  | 57582 |  | SSL Self-Signed Certificate | Purchase or generate a proper SSL certificate for this service. | Medium | Synopsis: The SSL certificate chain for this service ends in an unrecognized self-signed certificate. Description: The X.509 certificate chain for this service is not signed by a recognized certificate authority. If the remote host is a public host in production, this nullifies the use of SSL as anyone could establish a man-in-the-middle attack against the remote host. Note that this plugin does not check for certificate chains that end in a certificate that is not self-signed, but is signed by an unrecognized certificate authority. Solution: Purchase or generate a proper SSL certificate for this service. See Also: NULL CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:N Plugin Output: The following certificate was found at the top of the certificate chain sent by the remote host, but is self-signed and was not found in the list of known certificate authorities : ¦-Subject : CN=VHADETCCANEXOPR.v11.med.va.gov CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2022-06-14 Exploitability Ease: NULL Plugin Type: remote Source File: | Nexodose Server | BRACCO DIAGNOSTICS INC (US) |  |  | Region X | VISN X | VAMC | VHAXXXCCANEXOPR |  |  |  |  | VHAXXXCCANEXOPR.vxx.med.va.gov |  | Microsoft Windows 10 | tcp | 3389 | General | Repeat Vulnerability | 2012-01-17 | 2022-11-07 22:23 | 2023-05-09 02:05 | 6
|  |  | 104743 |  | TLS Version 1.0 Protocol Detection | Enable support for TLS 1.2 and 1.3, and disable support for TLS 1.0. | Medium | Synopsis: The remote service encrypts traffic using an older version of TLS. Description: The remote service accepts connections encrypted using TLS 1.0. TLS 1.0 has a number of cryptographic design flaws. Modern implementations of TLS 1.0 mitigate these problems, but newer versions of TLS like 1.2 and 1.3 are designed against these flaws and should be used whenever possible. As of March 31, 2020, Endpoints that aren’t enabled for TLS 1.2 and higher will no longer function properly with major web browsers and major vendors. PCI DSS v3.2 requires that TLS 1.0 be disabled entirely by June 30, 2018, except for POS POI terminals (and the SSL/TLS termination points to which they connect) that can be verified as not being susceptible to any known exploits. Solution: Enable support for TLS 1.2 and 1.3, and disable support for TLS 1.0. See Also: https://tools.ietf.org/html/draft-ietf-tls-oldversions-deprecate-00 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:P/A:N Plugin Output: TLSv1 is enabled and the server supports at least one cipher. CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-04-19 Exploitability Ease: NULL Plugin Type: remote Source File: | Nexodose Server | BRACCO DIAGNOSTICS INC (US) |  |  | Region X | VISN X | VAMC | VHAXXXCCANEXOPR |  |  |  |  | VHAXXXCCANEXOPR.vxx.med.va.gov |  | Microsoft Windows 10 | tcp | 3389 | Service detection | Repeat Vulnerability | 2017-11-22 | 2022-11-07 22:23 | 2023-05-09 02:05 | 6
|  |  | 157288 |  | TLS Version 1.1 Protocol Deprecated | Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. | Medium | Synopsis: The remote service encrypts traffic using an older version of TLS. Description: The remote service accepts connections encrypted using TLS 1.1. TLS 1.1 lacks support for current and recommended cipher suites. Ciphers that support encryption before MAC computation, and authenticated encryption modes such as GCM cannot be used with TLS 1.1 As of March 31, 2020, Endpoints that are not enabled for TLS 1.2 and higher will no longer function properly with major web browsers and major vendors. Solution: Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. See Also: https://datatracker.ietf.org/doc/html/rfc8996 http://www.nessus.org/u?c8ae820d CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:P/A:N Plugin Output: TLSv1.1 is enabled and the server supports at least one cipher. CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-04-19 Exploitability Ease: NULL Plugin Type: remote Source File: | Nexodose Server | BRACCO DIAGNOSTICS INC (US) |  |  | Region X | VISN X | VAMC | VHAXXXCCANEXOPR |  |  |  |  | VHAXXXCCANEXOPR.vxx.med.va.gov |  | Microsoft Windows 10 | tcp | 3389 | Service detection | Repeat Vulnerability | 2022-04-04 | 2022-11-07 22:23 | 2023-05-09 02:05 | 6
|  |  | 42873 |  | SSL Medium Strength Cipher Suites Supported (SWEET32) | Reconfigure the affected application if possible to avoid use of medium strength ciphers. | High | Synopsis: The remote service supports the use of medium strength SSL ciphers. Description: The remote host supports the use of SSL ciphers that offer medium strength encryption. Nessus regards medium strength as any encryption that uses key lengths at least 64 bits and less than 112 bits, or else that uses the 3DES encryption suite. Note that it is considerably easier to circumvent medium strength encryption if the attacker is on the same physical network. Solution: Reconfigure the affected application if possible to avoid use of medium strength ciphers. See Also: https://www.openssl.org/blog/blog/2016/08/24/sweet32/ https://sweet32.info CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:N/A:N Plugin Output: Medium Strength Ciphers (> 64-bit and < 112-bit key, or 3DES) Name Code KEX Auth Encryption MAC ---------------------- ---------- --- ---- --------------------- --- DES-CBC3-SHA 0x00, 0x0A RSA RSA 3DES-CBC(168) SHA1 The fields above are : {Tenable ciphername} {Cipher ID code} Kex={key exchange} Auth={authentication} Encrypt={symmetric encryption method} MAC={message authentication code} {export flag} CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2016-08-24 Patch Publication Date: NULL Plugin Modification Date: 2021-02-03 Exploitability Ease: NULL Plugin Type: remote Source File: | cyberREN Core Server | Cybernius Medical |  |  | Region X | VISN X | VAMC | VHAXXXAPPCYBER2 |  |  |  |  | VHAXXXAPPCYBER2.vxx.med.va.gov |  | Linux Kernel 3.2 | tcp | 3389 | General | Repeat Vulnerability | 2009-11-23 | 2020-09-03 02:23 | 2023-05-09 02:36 | 32
|  |  | 51192 |  | SSL Certificate Cannot Be Trusted | Purchase or generate a proper SSL certificate for this service. | Medium | Synopsis: The SSL certificate for this service cannot be trusted. Description: The server's X.509 certificate cannot be trusted. This situation can occur in three different ways, in which the chain of trust can be broken, as stated below : - First, the top of the certificate chain sent by the server might not be descended from a known public certificate authority. This can occur either when the top of the chain is an unrecognized, self-signed certificate, or when intermediate certificates are missing that would connect the top of the certificate chain to a known public certificate authority. - Second, the certificate chain may contain a certificate that is not valid at the time of the scan. This can occur either when the scan occurs before one of the certificate's 'notBefore' dates, or after one of the certificate's 'notAfter' dates. - Third, the certificate chain may contain a signature that either didn't match the certificate's information or could not be verified. Bad signatures can be fixed by getting the certificate with the bad signature to be re-signed by its issuer. Signatures that could not be verified are the result of the certificate's issuer using a signing algorithm that Nessus either does not support or does not recognize. If the remote host is a public host in production, any break in the chain makes it more difficult for users to verify the authenticity and identity of the web server. This could make it easier to carry out man-in-the-middle attacks against the remote host. Solution: Purchase or generate a proper SSL certificate for this service. See Also: https://www.itu.int/rec/T-REC-X.509/en https://en.wikipedia.org/wiki/X.509 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:N Plugin Output: The following certificate was at the top of the certificate chain sent by the remote host, but it is signed by an unknown certificate authority : ¦-Subject : CN=VHADETAPPCYBER2.v11.med.va.gov ¦-Issuer : CN=VHADETAPPCYBER2.v11.med.va.gov CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2020-04-27 Exploitability Ease: NULL Plugin Type: remote Source File: | cyberREN Core Server | Cybernius Medical |  |  | Region X | VISN X | VAMC | VHAXXXAPPCYBER2 |  |  |  |  | VHAXXXAPPCYBER2.vxx.med.va.gov |  | Linux Kernel 3.2 | tcp | 3389 | General | Repeat Vulnerability | 2010-12-15 | 2020-09-03 02:23 | 2023-05-09 02:36 | 32
|  |  | 57582 |  | SSL Self-Signed Certificate | Purchase or generate a proper SSL certificate for this service. | Medium | Synopsis: The SSL certificate chain for this service ends in an unrecognized self-signed certificate. Description: The X.509 certificate chain for this service is not signed by a recognized certificate authority. If the remote host is a public host in production, this nullifies the use of SSL as anyone could establish a man-in-the-middle attack against the remote host. Note that this plugin does not check for certificate chains that end in a certificate that is not self-signed, but is signed by an unrecognized certificate authority. Solution: Purchase or generate a proper SSL certificate for this service. See Also: NULL CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:N Plugin Output: The following certificate was found at the top of the certificate chain sent by the remote host, but is self-signed and was not found in the list of known certificate authorities : ¦-Subject : CN=VHADETAPPCYBER2.v11.med.va.gov CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2022-06-14 Exploitability Ease: NULL Plugin Type: remote Source File: | cyberREN Core Server | Cybernius Medical |  |  | Region X | VISN X | VAMC | VHAXXXAPPCYBER2 |  |  |  |  | VHAXXXAPPCYBER2.vxx.med.va.gov |  | Linux Kernel 3.2 | tcp | 3389 | General | Repeat Vulnerability | 2012-01-17 | 2020-09-03 02:23 | 2023-05-09 02:36 | 32
|  |  | 104743 |  | TLS Version 1.0 Protocol Detection | Enable support for TLS 1.2 and 1.3, and disable support for TLS 1.0. | Medium | Synopsis: The remote service encrypts traffic using an older version of TLS. Description: The remote service accepts connections encrypted using TLS 1.0. TLS 1.0 has a number of cryptographic design flaws. Modern implementations of TLS 1.0 mitigate these problems, but newer versions of TLS like 1.2 and 1.3 are designed against these flaws and should be used whenever possible. As of March 31, 2020, Endpoints that aren’t enabled for TLS 1.2 and higher will no longer function properly with major web browsers and major vendors. PCI DSS v3.2 requires that TLS 1.0 be disabled entirely by June 30, 2018, except for POS POI terminals (and the SSL/TLS termination points to which they connect) that can be verified as not being susceptible to any known exploits. Solution: Enable support for TLS 1.2 and 1.3, and disable support for TLS 1.0. See Also: https://tools.ietf.org/html/draft-ietf-tls-oldversions-deprecate-00 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:P/A:N Plugin Output: TLSv1 is enabled and the server supports at least one cipher. CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-04-19 Exploitability Ease: NULL Plugin Type: remote Source File: | cyberREN Core Server | Cybernius Medical |  |  | Region X | VISN X | VAMC | VHAXXXAPPCYBER2 |  |  |  |  | VHAXXXAPPCYBER2.vxx.med.va.gov |  | Linux Kernel 3.2 | tcp | 3389 | Service detection | Repeat Vulnerability | 2017-11-22 | 2020-09-03 02:23 | 2023-05-09 02:36 | 32
|  |  | 157288 |  | TLS Version 1.1 Protocol Deprecated | Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. | Medium | Synopsis: The remote service encrypts traffic using an older version of TLS. Description: The remote service accepts connections encrypted using TLS 1.1. TLS 1.1 lacks support for current and recommended cipher suites. Ciphers that support encryption before MAC computation, and authenticated encryption modes such as GCM cannot be used with TLS 1.1 As of March 31, 2020, Endpoints that are not enabled for TLS 1.2 and higher will no longer function properly with major web browsers and major vendors. Solution: Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. See Also: https://datatracker.ietf.org/doc/html/rfc8996 http://www.nessus.org/u?c8ae820d CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:P/A:N Plugin Output: TLSv1.1 is enabled and the server supports at least one cipher. CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-04-19 Exploitability Ease: NULL Plugin Type: remote Source File: | cyberREN Core Server | Cybernius Medical |  |  | Region X | VISN X | VAMC | VHAXXXAPPCYBER2 |  |  |  |  | VHAXXXAPPCYBER2.vxx.med.va.gov |  | Linux Kernel 3.2 | tcp | 3389 | Service detection | Repeat Vulnerability | 2022-04-04 | 2022-05-02 06:33 | 2023-05-09 02:36 | 12
|  |  | 172178 |  | ASP.NET Core SEoL | Upgrade to a version of ASP.NET Core that is currently supported. | Critical | Synopsis: An unsupported version of ASP.NET Core is installed on the remote host. Description: According to its version, the ASP.NET Core installed on the remote host is no longer maintained by its vendor or provider. Lack of support implies that no new security patches for the product will be released by the vendor. As a result, it may contain security vulnerabilities. Solution: Upgrade to a version of ASP.NET Core that is currently supported. See Also: http://www.nessus.org/u?89faa62b CVSS Base Score: 10 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.AspNetCore.App\3.1.16 Installed version : 3.1.16 Security End of Life : December 13, 2022 Time since Security End of Life (Est.) : 0 Years, 4 Months, 23 Days ¦ 143 Total Days CPE: cpe:/a:microsoft:asp.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-03-07 Exploitability Ease: NULL Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp |  | Misc. | Repeat Vulnerability | 2023-03-07 | 2023-04-04 03:53 | 2023-05-09 06:58 | 1
|  |  | 172179 |  | Microsoft .NET Core SEoL | Upgrade to a version of Microsoft .NET Core that is currently supported. | Critical | Synopsis: An unsupported version of Microsoft .NET Core is installed on the remote host. Description: According to its version, the Microsoft .NET Core installed on the remote host is no longer maintained by its vendor or provider. Lack of support implies that no new security patches for the product will be released by the vendor. As a result, it may contain security vulnerabilities. Solution: Upgrade to a version of Microsoft .NET Core that is currently supported. See Also: http://www.nessus.org/u?89faa62b CVSS Base Score: 10 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.NetCore.App\3.1.16\ Installed version : 3.1.16.30112 Security End of Life : December 13, 2022 Time since Security End of Life (Est.) : 0 Years, 4 Months, 23 Days ¦ 143 Total Days CPE: cpe:/a:microsoft:.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-03-07 Exploitability Ease: NULL Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp |  | Misc. | Repeat Vulnerability | 2023-03-07 | 2023-04-04 03:53 | 2023-05-09 06:58 | 1
|  |  | 172532 |  | KB5023697: Windows 10 Version 1607 and Windows Server 2016 Security Update (March 2023) | Apply Security Update 5023697 | Critical | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5023697. It is, therefore, affected by multiple vulnerabilities - An out-of-bounds write vulnerability exists in TPM2.0's Module Library allowing writing of a 2-byte data past the end of TPM2.0 command in the CryptParameterDecryption routine. An attacker who can successfully exploit this vulnerability can lead to denial of service (crashing the TPM chip/process or rendering it unusable) and/or arbitrary code execution in the TPM context. (CVE-2023-1017) - An out-of-bounds read vulnerability exists in TPM2.0's Module Library allowing a 2-byte read past the end of a TPM2.0 command in the CryptParameterDecryption routine. An attacker who can successfully exploit this vulnerability can read or access sensitive data stored in the TPM. (CVE-2023-1018) - Remote Procedure Call Runtime Remote Code Execution Vulnerability (CVE-2023-21708, CVE-2023-23405, CVE-2023-24869, CVE-2023-24908) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5023697 See Also: https://support.microsoft.com/help/5023697 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5023697 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 10.0.14393.5717 Should be : 10.0.14393.5786 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-02-28 Patch Publication Date: 2023-03-14 Plugin Modification Date: 2023-04-13 Exploitability Ease: Exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2023-03-14 | 2023-04-04 03:53 | 2023-05-09 06:58 | 1
|  |  | 172178 |  | ASP.NET Core SEoL | Upgrade to a version of ASP.NET Core that is currently supported. | Critical | Synopsis: An unsupported version of ASP.NET Core is installed on the remote host. Description: According to its version, the ASP.NET Core installed on the remote host is no longer maintained by its vendor or provider. Lack of support implies that no new security patches for the product will be released by the vendor. As a result, it may contain security vulnerabilities. Solution: Upgrade to a version of ASP.NET Core that is currently supported. See Also: http://www.nessus.org/u?89faa62b CVSS Base Score: 10 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files (x86)\dotnet\shared\Microsoft.AspNetCore.App\3.1.16 Installed version : 3.1.16 Security End of Life : December 13, 2022 Time since Security End of Life (Est.) : 0 Years, 4 Months, 23 Days ¦ 143 Total Days CPE: cpe:/a:microsoft:asp.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-03-07 Exploitability Ease: NULL Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp |  | Misc. | Repeat Vulnerability | 2023-03-07 | 2023-04-04 03:47 | 2023-05-09 06:45 | 1
|  |  | 172179 |  | Microsoft .NET Core SEoL | Upgrade to a version of Microsoft .NET Core that is currently supported. | Critical | Synopsis: An unsupported version of Microsoft .NET Core is installed on the remote host. Description: According to its version, the Microsoft .NET Core installed on the remote host is no longer maintained by its vendor or provider. Lack of support implies that no new security patches for the product will be released by the vendor. As a result, it may contain security vulnerabilities. Solution: Upgrade to a version of Microsoft .NET Core that is currently supported. See Also: http://www.nessus.org/u?89faa62b CVSS Base Score: 10 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.NetCore.App\3.1.16\ Installed version : 3.1.16.30112 Security End of Life : December 13, 2022 Time since Security End of Life (Est.) : 0 Years, 4 Months, 23 Days ¦ 143 Total Days CPE: cpe:/a:microsoft:.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-03-07 Exploitability Ease: NULL Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp |  | Misc. | Repeat Vulnerability | 2023-03-07 | 2023-04-04 03:47 | 2023-05-09 06:45 | 1
|  |  | 35453 |  | Microsoft Windows Update Reboot Required | Reboot the remote system to put pending changes into effect. | High | Synopsis: The remote Windows host requires a reboot. Description: According to entries in its registry, a reboot is required by Windows Update to complete installation of at least one update. If the pending changes are security-related, the remote host could remain vulnerable to attack until a reboot occurs. Solution: Reboot the remote system to put pending changes into effect. See Also: http://technet.microsoft.com/en-us/library/cc960241.aspx CVSS Base Score: NULL CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Nessus determined a reboot is required based on the following info : One or more applications have 'RebootRequired' flag set. CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2020-09-16 Exploitability Ease: NULL Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp |  | Windows | Repeat Vulnerability | 2009-01-23 | 2022-10-03 05:05 | 2023-05-09 06:58 | 7
|  |  | 87882 |  | MS16-004: Security Update for Microsoft Office to Address Remote Code Execution (3124585) | Microsoft has released a set of patches for Office 2007, 2010, 2013, 2013 RT, 2016, Word, Word Viewer, Excel, Excel Viewer, PowerPoint, Visio, SharePoint Server 2013, SharePoint Foundation 2013, Microsoft Office Compatibility Pack, and Visual Basic 6.0 Runtime. | High | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host has a version of Microsoft Office, Word, Word Viewer, Excel, Excel Viewer, PowerPoint, Visio, SharePoint, Visual Basic, or Microsoft Office Compatibility Pack installed that is affected by multiple vulnerabilities : - Multiple cross-site scripting vulnerabilities exist in Microsoft SharePoint due to improper enforcement of Access Control Policy (ACP) configuration settings. A remote attacker can exploit these vulnerabilities, via a specially crafted request, to execute arbitrary script code in a user's browser session. (CVE-2015-6117, CVE-2016-0011) - Multiple remote code execution vulnerabilities exist in Microsoft Office due to improper handling of objects in memory. An attacker can exploit these vulnerabilities by convincing a user to open a specially crafted file in Microsoft Office, resulting in execution of arbitrary code in the context of the current user. (CVE-2016-0010, CVE-2016-0035) - An information disclosure vulnerability exists in Microsoft Office due to a failure to use the Address Space Layout Randomization (ASLR) security feature. An attacker can exploit this to predict memory offsets of specific instructions in a call stack. (CVE-2016-0012) Solution: Microsoft has released a set of patches for Office 2007, 2010, 2013, 2013 RT, 2016, Word, Word Viewer, Excel, Excel Viewer, PowerPoint, Visio, SharePoint Server 2013, SharePoint Foundation 2013, Microsoft Office Compatibility Pack, and Visual Basic 6.0 Runtime. See Also: https://technet.microsoft.com/library/security/ms16-004 CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:C/I:C/A:C Plugin Output: Product : Microsoft Visual Basic 6.0 KB : 3096896 - C:\windows\SysWOW64\mscomctl.ocx has not been patched. Remote version : 6.0.88.62 Should be : 6.1.98.46 CPE: cpe:/a:microsoft:office cpe:/a:microsoft:word cpe:/a:microsoft:word_viewer cpe:/a:microsoft:excel cpe:/a:microsoft:excel_viewer cpe:/a:microsoft:powerpoint cpe:/a:microsoft:visio cpe:/a:microsoft:sharepoint_server cpe:/a:microsoft:sharepoint_foundation cpe:/a:microsoft:visual_basic cpe:/a:microsoft:office_compatibility_pack CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2016-01-12 Patch Publication Date: 2016-01-12 Plugin Modification Date: 2023-02-17 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2016-01-12 | 2022-08-07 23:43 | 2023-05-09 06:58 | 9
|  |  | 152488 |  | Security Update for .NET Core (August 2021) | Update .NET Core, remove vulnerable packages and refer to vendor advisory. | High | Synopsis: The remote Windows host is affected by a .NET Core denial of service (DoS) vulnerability. Description: The Microsoft .NET Core installation on the remote host is version 2.1.x prior to 2.1.29, 3.1.x prior to 3.1.18, or 5.x prior to 5.0.9. It is, therefore affected by a denial of service (DoS) vulnerability, as server applications providing WebSocket endpoints can be tricked into endlessly looping while trying to read a single WebSocket frame. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update .NET Core, remove vulnerable packages and refer to vendor advisory. See Also: https://dotnet.microsoft.com/download/dotnet/2.1 https://dotnet.microsoft.com/download/dotnet-core/3.1 https://dotnet.microsoft.com/download/dotnet/5.0 http://www.nessus.org/u?8ff12246 http://www.nessus.org/u?0933ffe1 http://www.nessus.org/u?6242d65f https://devblogs.microsoft.com/dotnet/net-august-2021/ CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.NetCore.App\3.1.16\ Installed version : 3.1.16.30112 Fixed version : 3.1.18 CPE: cpe:/a:microsoft:.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-08-10 Patch Publication Date: 2021-08-10 Plugin Modification Date: 2021-08-20 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2021-08-11 | 2022-02-06 21:22 | 2023-05-09 06:58 | 15
|  |  | 156227 |  | Security Updates for Microsoft ASP.NET Core (December 2021) | Update ASP.NET Core, remove vulnerable packages and refer to vendor advisory. | High | Synopsis: The Microsoft ASP.NET Core installations on the remote host are missing a security update. Description: The Microsoft ASP.NET Core installations on the remote host are missing a security update. It is, therefore, affected by an elevation of privilege vulnerability. An attacker can exploit this to gain elevated privileges. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update ASP.NET Core, remove vulnerable packages and refer to vendor advisory. See Also: https://github.com/dotnet/announcements/issues/206 https://github.com/dotnet/aspnetcore/issues/39028 CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files (x86)\dotnet\shared\Microsoft.AspNetCore.App\3.1.16 Installed version : 3.1.16 Fixed version : 3.1.22 CPE: cpe:/a:microsoft:asp.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-12-14 Patch Publication Date: 2021-12-14 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2021-12-21 | 2022-02-06 21:22 | 2023-05-09 06:58 | 15
|  |  | 161167 |  | Security Updates for Microsoft .NET core (May 2022) | Update .NET Core Runtime to version 3.1.25, 5.0.17 or 6.0.5. | High | Synopsis: The Microsoft .NET core installations on the remote host are affected by multiple vulnerabilities. Description: The Microsoft .NET core installations on the remote host are missing security updates. It is, therefore, affected by multiple denial of service vulnerabilities: - A vulnerability where a malicious client can cause a denial of service via excess memory allocations through HttpClient. (CVE-2022-23267) - A vulnerability where a malicious client can manipulate cookies and cause a denial of service. (CVE-2022-29117) - A vulnerability where a malicious client can cause a denial of service when HTML forms are parsed. (CVE-2022-29145) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update .NET Core Runtime to version 3.1.25, 5.0.17 or 6.0.5. See Also: https://dotnet.microsoft.com/download/dotnet/3.1 https://dotnet.microsoft.com/download/dotnet/5.0 https://dotnet.microsoft.com/download/dotnet/6.0 https://github.com/dotnet/announcements/issues/219 http://www.nessus.org/u?3b99f604 http://www.nessus.org/u?b1b0aff4 http://www.nessus.org/u?39d07c32 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.NetCore.App\3.1.16\ Installed version : 3.1.16.30112 Fixed version : 3.1.25 CPE: cpe:/a:microsoft:.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-05-10 Patch Publication Date: 2022-05-10 Plugin Modification Date: 2022-11-04 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Misc. | Repeat Vulnerability | 2022-05-13 | 2022-07-10 08:29 | 2023-05-09 06:58 | 10
|  |  | 162393 |  | Security Updates for Microsoft SQL Server (June 2022) | Microsoft has released the following security updates to address this issue: -KB5015371 -KB5014553 -KB5014351 -KB5014353 -KB5014354 -KB5014356 -KB5014365 -KB5014355 -KB5014165 -KB5014164 | High | Synopsis: The Microsoft SQL Server installation on the remote host is missing a security update. Description: The Microsoft SQL Server installation on the remote host is missing a security update. It is, therefore, affected by the following vulnerability: - A remote code execution vulnerability. An attacker can exploit this to bypass authentication and execute unauthorized arbitrary commands. (CVE-2022-29143) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Microsoft has released the following security updates to address this issue: -KB5015371 -KB5014553 -KB5014351 -KB5014353 -KB5014354 -KB5014356 -KB5014365 -KB5014355 -KB5014165 -KB5014164 See Also: https://support.microsoft.com/en-us/help/5015371 https://support.microsoft.com/en-us/help/5014553 https://support.microsoft.com/en-us/help/5014351 https://support.microsoft.com/en-us/help/5014353 https://support.microsoft.com/en-us/help/5014354 https://support.microsoft.com/en-us/help/5014356 https://support.microsoft.com/en-us/help/5014365 https://support.microsoft.com/en-us/help/5014355 https://support.microsoft.com/en-us/help/5014165 https://support.microsoft.com/en-us/help/5014164 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:M/Au:S/C:P/I:P/A:P Plugin Output: KB : 5014553 - C:\Program Files\Microsoft SQL Server\MSSQL14.RALS_PLUS\MSSQL\Binn\sqlservr.exe has not been patched. Remote version : 2017.140.3391.2 Should be : 2017.140.3445.2 SQL Server Version : 14.0.3391.2 Standard Edition SQL Server Instance : RALS_PLUS CPE: cpe:/a:microsoft:sql_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-06-14 Patch Publication Date: 2022-06-14 Plugin Modification Date: 2023-02-16 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-06-17 | 2022-07-10 08:29 | 2023-05-09 06:58 | 10
|  |  | 165076 |  | Security Updates for Microsoft ASP.NET Core (September 2022) | Update ASP.NET Core Runtime to version 3.1.29 or 6.0.9. | High | Synopsis: The Microsoft ASP.NET Core installations on the remote host are missing a security update. Description: A denial of service vulnerability exists in ASP.NET core 6.0 < 6.0.9 and ASP.NET Core 3.1 < 3.1.29. An unauthenticated, remote attacker can exploit this, by sending a customized payload that is parsed during model binding, to cause a stack overflow, which may cause the application to stop responding. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update ASP.NET Core Runtime to version 3.1.29 or 6.0.9. See Also: https://github.com/dotnet/announcements/issues/234 http://www.nessus.org/u?c76821a3 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:C Plugin Output: Path : C:\Program Files (x86)\dotnet\shared\Microsoft.AspNetCore.App\3.1.16 Installed version : 3.1.16 Fixed version : 3.1.29 CPE: cpe:/a:microsoft:asp.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-09-13 Patch Publication Date: 2022-09-13 Plugin Modification Date: 2022-11-30 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-09-14 | 2022-10-03 05:05 | 2023-05-09 06:58 | 7
|  |  | 165077 |  | Security Updates for Microsoft .NET Core (September 2022) | Update .NET Core Runtime to version 3.1.29 or 6.0.9. | High | Synopsis: The Microsoft .NET core installations on the remote host are affected by a denial of service vulnerability. Description: A denial of service vulnerability exists in .NET core 6.0 < 6.0.9 and .NET Core 3.1 < 3.1.29. An unauthenticated, remote attacker can exploit this, by sending a customized payload that is parsed during model binding, to cause a stack overflow, which may cause the application to stop responding. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update .NET Core Runtime to version 3.1.29 or 6.0.9. See Also: https://support.microsoft.com/help/5017903 https://support.microsoft.com/help/5017915 https://dotnet.microsoft.com/download/dotnet/3.1 https://dotnet.microsoft.com/download/dotnet/6.0 http://www.nessus.org/u?cf2fdae6 http://www.nessus.org/u?775af4a9 https://github.com/dotnet/core/issues/7791 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:C Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.NetCore.App\3.1.16\ Installed version : 3.1.16.30112 Fixed version : 3.1.29 CPE: cpe:/a:microsoft:.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-09-13 Patch Publication Date: 2022-09-13 Plugin Modification Date: 2022-10-18 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-09-14 | 2022-10-03 05:05 | 2023-05-09 06:58 | 7
|  |  | 166054 |  | Security Updates for Microsoft .NET Core (October 2022) | Update .NET Core Runtime to version 3.1.30 or 6.0.10. | High | Synopsis: The Microsoft .NET core installations on the remote host are affected by a privilege escalation vulnerability. Description: A privilege escalation vulnerability exists in .NET core 6.0 < 6.0.10 and .NET Core 3.1 < 3.1.30. An authenticated, local attacker can exploit this, via the NuGet client, to cause the user to execute arbitrary code. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update .NET Core Runtime to version 3.1.30 or 6.0.10. See Also: https://support.microsoft.com/help/5019349 https://support.microsoft.com/help/5019351 https://dotnet.microsoft.com/download/dotnet/3.1 https://dotnet.microsoft.com/download/dotnet/6.0 http://www.nessus.org/u?1a5250e3 http://www.nessus.org/u?0eafd070 https://github.com/dotnet/core/issues/7864 CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:L/Au:S/C:C/I:C/A:C Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.NetCore.App\3.1.16\ Installed version : 3.1.16.30112 Fixed version : 3.1.30 CPE: cpe:/a:microsoft:.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-10-11 Patch Publication Date: 2022-10-11 Plugin Modification Date: 2022-12-15 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-10-12 | 2022-11-08 03:12 | 2023-05-09 06:58 | 6
|  |  | 168747 |  | Security Updates for Microsoft .NET Core (December 2022) | Update .NET Core Runtime to version 3.1.32 or 6.0.12 or 7.0.1. | High | Synopsis: The Microsoft .NET core installations on the remote host are affected by remote code execution vulnerability. Description: A remote code execution vulnerability exists in .NET Core 3.1, .NET 6.0, and .NET 7.0, where a malicious actor could cause a user to run arbitrary code as a result of parsing maliciously crafted xps files. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update .NET Core Runtime to version 3.1.32 or 6.0.12 or 7.0.1. See Also: https://support.microsoft.com/en-us/help/5021953 https://support.microsoft.com/en-us/help/5021954 https://support.microsoft.com/en-us/help/5021955 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.NetCore.App\3.1.16\ Installed version : 3.1.16.30112 Fixed version : 3.1.32 CPE: cpe:/a:microsoft:.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-12-13 Patch Publication Date: 2022-12-13 Plugin Modification Date: 2022-12-19 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-12-15 | 2023-01-09 04:38 | 2023-05-09 06:58 | 4
|  |  | 168826 |  | Security Updates for Microsoft ASP.NET Core (December 2022) | Update ASP.NET Core Runtime to version 3.1.32 or 6.0.12 or 7.0.1. | High | Synopsis: The Microsoft ASP.NET core installations on the remote host are affected by remote code execution vulnerability. Description: A remote code execution vulnerability exists in ASP.NET core 3.1, ASP.NET 6.0, and ASP.NET 7.0, where a malicious actor could cause a user to run arbitrary code as a result of parsing maliciously crafted xps files. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update ASP.NET Core Runtime to version 3.1.32 or 6.0.12 or 7.0.1. See Also: https://support.microsoft.com/en-us/help/5021953 https://support.microsoft.com/en-us/help/5021954 https://support.microsoft.com/en-us/help/5021955 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.AspNetCore.App\3.1.16 Installed version : 3.1.16 Fixed version : 3.1.32 CPE: cpe:/a:microsoft:asp.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-12-13 Patch Publication Date: 2022-12-13 Plugin Modification Date: 2022-12-16 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-12-15 | 2023-01-09 04:38 | 2023-05-09 06:58 | 4
|  |  | 171268 |  | Microsoft Edge (Chromium) < 110.0.1587.41 Multiple Vulnerabilities | Upgrade to Microsoft Edge version 110.0.1587.41 or later. | High | Synopsis: The remote host has an web browser installed that is affected by multiple vulnerabilities. Description: The version of Microsoft Edge installed on the remote Windows host is prior to 110.0.1587.41. It is, therefore, affected by multiple vulnerabilities as referenced in the February 9, 2023 advisory. - Type confusion in V8 in Google Chrome prior to 110.0.5481.77 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-0696) - Inappropriate implementation in Full screen mode in Google Chrome on Android prior to 110.0.5481.77 allowed a remote attacker to spoof the contents of the security UI via a crafted HTML page. (Chromium security severity: High) (CVE-2023-0697) - Out of bounds read in WebRTC in Google Chrome prior to 110.0.5481.77 allowed a remote attacker to perform an out of bounds memory read via a crafted HTML page. (Chromium security severity: High) (CVE-2023-0698) - Use after free in GPU in Google Chrome prior to 110.0.5481.77 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page and browser shutdown. (Chromium security severity: Medium) (CVE-2023-0699) - Inappropriate implementation in Download in Google Chrome prior to 110.0.5481.77 allowed a remote attacker to potentially spoof the contents of the Omnibox (URL bar) via a crafted HTML page. (Chromium security severity: Medium) (CVE-2023-0700) - Heap buffer overflow in WebUI in Google Chrome prior to 110.0.5481.77 allowed a remote attacker who convinced a user to engage in specific UI interactions to potentially exploit heap corruption via UI interaction . (Chromium security severity: Medium) (CVE-2023-0701) - Type confusion in Data Transfer in Google Chrome prior to 110.0.5481.77 allowed a remote attacker who convinced a user to engage in specific UI interactions to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: Medium) (CVE-2023-0702) - Type confusion in DevTools in Google Chrome prior to 110.0.5481.77 allowed a remote attacker who convinced a user to engage in specific UI interactions to potentially exploit heap corruption via UI interactions. (Chromium security severity: Medium) (CVE-2023-0703) - Insufficient policy enforcement in DevTools in Google Chrome prior to 110.0.5481.77 allowed a remote attacker to bypass same origin policy and proxy settings via a crafted HTML page. (Chromium security severity: Low) (CVE-2023-0704) - Integer overflow in Core in Google Chrome prior to 110.0.5481.77 allowed a remote attacker who had one a race condition to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: Low) (CVE-2023-0705) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Microsoft Edge version 110.0.1587.41 or later. See Also: http://www.nessus.org/u?245dfb65 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0696 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0697 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0698 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0699 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0700 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0701 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0702 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0703 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0704 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0705 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-21794 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files (x86)\Microsoft\Edge\Application Installed versi | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2023-02-09 | 2023-03-07 01:04 | 2023-05-09 06:58 | 2
|  |  | 171332 |  | Microsoft Edge (Chromium) < 109.0.1518.70 / 108.0.1462.95 Multiple Vulnerabilities | Upgrade to Microsoft Edge version 109.0.1518.70 / 108.0.1462.95 or later. | High | Synopsis: The remote host has an web browser installed that is affected by multiple vulnerabilities. Description: The version of Microsoft Edge installed on the remote Windows host is prior to 109.0.1518.70 / 108.0.1462.95. It is, therefore, affected by multiple vulnerabilities as referenced in the January 26, 2023 advisory. - Use after free in WebTransport in Google Chrome prior to 109.0.5414.119 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-0471) - Use after free in WebRTC in Google Chrome prior to 109.0.5414.119 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-0472) - Type Confusion in ServiceWorker API in Google Chrome prior to 109.0.5414.119 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: Medium) (CVE-2023-0473) - Use after free in GuestView in Google Chrome prior to 109.0.5414.119 allowed an attacker who convinced a user to install a malicious extension to potentially exploit heap corruption via a Chrome web app. (Chromium security severity: Medium) (CVE-2023-0474) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Microsoft Edge version 109.0.1518.70 / 108.0.1462.95 or later. See Also: http://www.nessus.org/u?245dfb65 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0471 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0472 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0473 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0474 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files (x86)\Microsoft\Edge\Application Installed version : 109.0.1518.55 Fixed version : 109.0.1518.70 CPE: cpe:/a:microsoft:edge CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-01-24 Patch Publication Date: 2023-01-26 Plugin Modification Date: 2023-02-10 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2023-02-10 | 2023-03-07 01:04 | 2023-05-09 06:58 | 2
|  |  | 171604 |  | Security Updates for Microsoft SQL Server (February 2023) | Microsoft has released the following security updates to address this issue: -KB5021126 -KB5021129 -KB5021522 -KB5021127 -KB5021045 -KB5021037 -KB5021128 -KB5021124 -KB5021125 -KB5020863 -KB5021112 -KB5021123 | High | Synopsis: The Microsoft SQL Server installation on the remote host is affected by multiple vulnerabilities. Description: The Microsoft SQL Server installation on the remote host is missing security updates. It is, therefore, affected by multiple vulnerabilities: - A remote code execution vulnerability. An attacker can exploit this to bypass authentication and execute unauthorized arbitrary commands. (CVE-2023-21528, CVE-2023-21568, CVE-2023-21704, CVE-2023-21705, CVE-2023-21713, CVE-2023-21718) Solution: Microsoft has released the following security updates to address this issue: -KB5021126 -KB5021129 -KB5021522 -KB5021127 -KB5021045 -KB5021037 -KB5021128 -KB5021124 -KB5021125 -KB5020863 -KB5021112 -KB5021123 See Also: https://support.microsoft.com/en-us/help/5020863 https://support.microsoft.com/en-us/help/5021112 https://support.microsoft.com/en-us/help/5021126 https://support.microsoft.com/en-us/help/5021129 https://support.microsoft.com/en-us/help/5021522 https://support.microsoft.com/en-us/help/5021127 https://support.microsoft.com/en-us/help/5021045 https://support.microsoft.com/en-us/help/5021037 https://support.microsoft.com/en-us/help/5021128 https://support.microsoft.com/en-us/help/5021123 https://support.microsoft.com/en-us/help/5021124 https://support.microsoft.com/en-us/help/5021125 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:S/C:C/I:C/A:C Plugin Output: KB : 5021126 - C:\Program Files\Microsoft SQL Server\MSSQL14.RALS_PLUS\MSSQL\Binn\sqlservr.exe has not been patched. Remote version : 2017.140.3391.2 Should be : 2017.140.3460.9 SQL Server Version : 14.0.3391.2 Standard Edition SQL Server Instance : RALS_PLUS CPE: cpe:/a:microsoft:sql_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-02-14 Patch Publication Date: 2023-02-14 Plugin Modification Date: 2023-03-08 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2023-02-17 | 2023-03-07 01:04 | 2023-05-09 06:58 | 2
|  |  | 171927 |  | Microsoft Edge (Chromium) < 110.0.1587.56 Multiple Vulnerabilities | Upgrade to Microsoft Edge version 110.0.1587.56 or later. | High | Synopsis: The remote host has an web browser installed that is affected by multiple vulnerabilities. Description: The version of Microsoft Edge installed on the remote Windows host is prior to 110.0.1587.56. It is, therefore, affected by multiple vulnerabilities as referenced in the February 25, 2023 advisory. - Use after free in Web Payments API in Google Chrome on Android prior to 110.0.5481.177 allowed a remote attacker who had compromised the renderer process to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-0927) - Use after free in SwiftShader in Google Chrome prior to 110.0.5481.177 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-0928) - Use after free in Vulkan in Google Chrome prior to 110.0.5481.177 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-0929) - Heap buffer overflow in Video in Google Chrome prior to 110.0.5481.177 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-0930) - Use after free in Video in Google Chrome prior to 110.0.5481.177 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-0931) - Use after free in WebRTC in Google Chrome on Windows prior to 110.0.5481.177 allowed a remote attacker who convinced the user to engage in specific UI interactions to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-0932) - Integer overflow in PDF in Google Chrome prior to 110.0.5481.177 allowed a remote attacker to potentially exploit heap corruption via a crafted PDF file. (Chromium security severity: Medium) (CVE-2023-0933) - Use after free in Prompts in Google Chrome prior to 110.0.5481.177 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: Critical) (CVE-2023-0941) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Microsoft Edge version 110.0.1587.56 or later. See Also: http://www.nessus.org/u?245dfb65 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0927 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0928 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0929 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0930 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0931 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0932 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0933 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-0941 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files (x86)\Microsoft\Edge\Application Installed version : 109.0.1518.55 Fixed version : 110.0.1587.56 CPE: cpe:/a:microsoft:edge CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-02-22 Patch Publication Date: 2023-02-25 Plugin Modification Date: 2023-03-16 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2023-02-27 | 2023-03-07 01:04 | 2023-05-09 06:58 | 2
|  |  | 63155 |  | Microsoft Windows Unquoted Service Path Enumeration | Ensure that any services that contain a space in the path enclose the path in quotes. | High | Synopsis: The remote Windows host has at least one service installed that uses an unquoted service path. Description: The remote Windows host has at least one service installed that uses an unquoted service path, which contains at least one whitespace. A local attacker can gain elevated privileges by inserting an executable file in the path of the affected service. Note that this is a generic test that will flag any application affected by the described vulnerability. Solution: Ensure that any services that contain a space in the path enclose the path in quotes. See Also: http://www.nessus.org/u?84a4cc1c http://cwe.mitre.org/data/definitions/428.html https://www.commonexploits.com/unquoted-service-paths/ http://www.nessus.org/u?4aa6acbc CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:C/A:C Plugin Output: Nessus found the following service with an untrusted path : SQLANYs_iSTATDEDB : C:\Program Files (x86)\SQL Anywhere 17\bin32\dbsrv17.exe CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2012-09-15 Patch Publication Date: NULL Plugin Modification Date: 2020-06-12 Exploitability Ease: Exploits are available Plugin Type: local Source File: | RALS I-Stat Data Exchange (DE) Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSD |  |  |  |  | VHAXXXCCARALSD.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2012-12-05 | 2022-02-06 21:14 | 2023-05-09 06:44 | 15
|  |  | 152488 |  | Security Update for .NET Core (August 2021) | Update .NET Core, remove vulnerable packages and refer to vendor advisory. | High | Synopsis: The remote Windows host is affected by a .NET Core denial of service (DoS) vulnerability. Description: The Microsoft .NET Core installation on the remote host is version 2.1.x prior to 2.1.29, 3.1.x prior to 3.1.18, or 5.x prior to 5.0.9. It is, therefore affected by a denial of service (DoS) vulnerability, as server applications providing WebSocket endpoints can be tricked into endlessly looping while trying to read a single WebSocket frame. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update .NET Core, remove vulnerable packages and refer to vendor advisory. See Also: https://dotnet.microsoft.com/download/dotnet/2.1 https://dotnet.microsoft.com/download/dotnet-core/3.1 https://dotnet.microsoft.com/download/dotnet/5.0 http://www.nessus.org/u?8ff12246 http://www.nessus.org/u?0933ffe1 http://www.nessus.org/u?6242d65f https://devblogs.microsoft.com/dotnet/net-august-2021/ CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.NetCore.App\3.1.16\ Installed version : 3.1.16.30112 Fixed version : 3.1.18 CPE: cpe:/a:microsoft:.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-08-10 Patch Publication Date: 2021-08-10 Plugin Modification Date: 2021-08-20 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2021-08-11 | 2022-02-06 21:14 | 2023-05-09 06:45 | 15
|  |  | 156227 |  | Security Updates for Microsoft ASP.NET Core (December 2021) | Update ASP.NET Core, remove vulnerable packages and refer to vendor advisory. | High | Synopsis: The Microsoft ASP.NET Core installations on the remote host are missing a security update. Description: The Microsoft ASP.NET Core installations on the remote host are missing a security update. It is, therefore, affected by an elevation of privilege vulnerability. An attacker can exploit this to gain elevated privileges. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update ASP.NET Core, remove vulnerable packages and refer to vendor advisory. See Also: https://github.com/dotnet/announcements/issues/206 https://github.com/dotnet/aspnetcore/issues/39028 CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.AspNetCore.App\3.1.16 Installed version : 3.1.16 Fixed version : 3.1.22 CPE: cpe:/a:microsoft:asp.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-12-14 Patch Publication Date: 2021-12-14 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2021-12-21 | 2022-02-06 21:14 | 2023-05-09 06:45 | 15
|  |  | 161167 |  | Security Updates for Microsoft .NET core (May 2022) | Update .NET Core Runtime to version 3.1.25, 5.0.17 or 6.0.5. | High | Synopsis: The Microsoft .NET core installations on the remote host are affected by multiple vulnerabilities. Description: The Microsoft .NET core installations on the remote host are missing security updates. It is, therefore, affected by multiple denial of service vulnerabilities: - A vulnerability where a malicious client can cause a denial of service via excess memory allocations through HttpClient. (CVE-2022-23267) - A vulnerability where a malicious client can manipulate cookies and cause a denial of service. (CVE-2022-29117) - A vulnerability where a malicious client can cause a denial of service when HTML forms are parsed. (CVE-2022-29145) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update .NET Core Runtime to version 3.1.25, 5.0.17 or 6.0.5. See Also: https://dotnet.microsoft.com/download/dotnet/3.1 https://dotnet.microsoft.com/download/dotnet/5.0 https://dotnet.microsoft.com/download/dotnet/6.0 https://github.com/dotnet/announcements/issues/219 http://www.nessus.org/u?3b99f604 http://www.nessus.org/u?b1b0aff4 http://www.nessus.org/u?39d07c32 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.NetCore.App\3.1.16\ Installed version : 3.1.16.30112 Fixed version : 3.1.25 CPE: cpe:/a:microsoft:.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-05-10 Patch Publication Date: 2022-05-10 Plugin Modification Date: 2022-11-04 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Misc. | Repeat Vulnerability | 2022-05-13 | 2022-07-10 08:10 | 2023-05-09 06:45 | 10
|  |  | 162393 |  | Security Updates for Microsoft SQL Server (June 2022) | Microsoft has released the following security updates to address this issue: -KB5015371 -KB5014553 -KB5014351 -KB5014353 -KB5014354 -KB5014356 -KB5014365 -KB5014355 -KB5014165 -KB5014164 | High | Synopsis: The Microsoft SQL Server installation on the remote host is missing a security update. Description: The Microsoft SQL Server installation on the remote host is missing a security update. It is, therefore, affected by the following vulnerability: - A remote code execution vulnerability. An attacker can exploit this to bypass authentication and execute unauthorized arbitrary commands. (CVE-2022-29143) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Microsoft has released the following security updates to address this issue: -KB5015371 -KB5014553 -KB5014351 -KB5014353 -KB5014354 -KB5014356 -KB5014365 -KB5014355 -KB5014165 -KB5014164 See Also: https://support.microsoft.com/en-us/help/5015371 https://support.microsoft.com/en-us/help/5014553 https://support.microsoft.com/en-us/help/5014351 https://support.microsoft.com/en-us/help/5014353 https://support.microsoft.com/en-us/help/5014354 https://support.microsoft.com/en-us/help/5014356 https://support.microsoft.com/en-us/help/5014365 https://support.microsoft.com/en-us/help/5014355 https://support.microsoft.com/en-us/help/5014165 https://support.microsoft.com/en-us/help/5014164 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:M/Au:S/C:P/I:P/A:P Plugin Output: KB : 5014553 - C:\Program Files\Microsoft SQL Server\MSSQL14.RALS_PLUS\MSSQL\Binn\sqlservr.exe has not been patched. Remote version : 2017.140.3391.2 Should be : 2017.140.3445.2 SQL Server Version : 14.0.3391.2 Express Edition SQL Server Instance : RALS_PLUS CPE: cpe:/a:microsoft:sql_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-06-14 Patch Publication Date: 2022-06-14 Plugin Modification Date: 2023-02-16 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-06-17 | 2022-07-10 08:10 | 2023-05-09 06:45 | 10
|  |  | 165076 |  | Security Updates for Microsoft ASP.NET Core (September 2022) | Update ASP.NET Core Runtime to version 3.1.29 or 6.0.9. | High | Synopsis: The Microsoft ASP.NET Core installations on the remote host are missing a security update. Description: A denial of service vulnerability exists in ASP.NET core 6.0 < 6.0.9 and ASP.NET Core 3.1 < 3.1.29. An unauthenticated, remote attacker can exploit this, by sending a customized payload that is parsed during model binding, to cause a stack overflow, which may cause the application to stop responding. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update ASP.NET Core Runtime to version 3.1.29 or 6.0.9. See Also: https://github.com/dotnet/announcements/issues/234 http://www.nessus.org/u?c76821a3 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:C Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.AspNetCore.App\3.1.16 Installed version : 3.1.16 Fixed version : 3.1.29 CPE: cpe:/a:microsoft:asp.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-09-13 Patch Publication Date: 2022-09-13 Plugin Modification Date: 2022-11-30 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-09-14 | 2022-10-03 04:58 | 2023-05-09 06:45 | 7
|  |  | 165077 |  | Security Updates for Microsoft .NET Core (September 2022) | Update .NET Core Runtime to version 3.1.29 or 6.0.9. | High | Synopsis: The Microsoft .NET core installations on the remote host are affected by a denial of service vulnerability. Description: A denial of service vulnerability exists in .NET core 6.0 < 6.0.9 and .NET Core 3.1 < 3.1.29. An unauthenticated, remote attacker can exploit this, by sending a customized payload that is parsed during model binding, to cause a stack overflow, which may cause the application to stop responding. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update .NET Core Runtime to version 3.1.29 or 6.0.9. See Also: https://support.microsoft.com/help/5017903 https://support.microsoft.com/help/5017915 https://dotnet.microsoft.com/download/dotnet/3.1 https://dotnet.microsoft.com/download/dotnet/6.0 http://www.nessus.org/u?cf2fdae6 http://www.nessus.org/u?775af4a9 https://github.com/dotnet/core/issues/7791 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:C Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.NetCore.App\3.1.16\ Installed version : 3.1.16.30112 Fixed version : 3.1.29 CPE: cpe:/a:microsoft:.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-09-13 Patch Publication Date: 2022-09-13 Plugin Modification Date: 2022-10-18 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-09-14 | 2022-10-03 04:58 | 2023-05-09 06:45 | 7
|  |  | 166054 |  | Security Updates for Microsoft .NET Core (October 2022) | Update .NET Core Runtime to version 3.1.30 or 6.0.10. | High | Synopsis: The Microsoft .NET core installations on the remote host are affected by a privilege escalation vulnerability. Description: A privilege escalation vulnerability exists in .NET core 6.0 < 6.0.10 and .NET Core 3.1 < 3.1.30. An authenticated, local attacker can exploit this, via the NuGet client, to cause the user to execute arbitrary code. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update .NET Core Runtime to version 3.1.30 or 6.0.10. See Also: https://support.microsoft.com/help/5019349 https://support.microsoft.com/help/5019351 https://dotnet.microsoft.com/download/dotnet/3.1 https://dotnet.microsoft.com/download/dotnet/6.0 http://www.nessus.org/u?1a5250e3 http://www.nessus.org/u?0eafd070 https://github.com/dotnet/core/issues/7864 CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:L/Au:S/C:C/I:C/A:C Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.NetCore.App\3.1.16\ Installed version : 3.1.16.30112 Fixed version : 3.1.30 CPE: cpe:/a:microsoft:.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-10-11 Patch Publication Date: 2022-10-11 Plugin Modification Date: 2022-12-15 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-10-12 | 2022-11-08 03:06 | 2023-05-09 06:45 | 6
|  |  | 168747 |  | Security Updates for Microsoft .NET Core (December 2022) | Update .NET Core Runtime to version 3.1.32 or 6.0.12 or 7.0.1. | High | Synopsis: The Microsoft .NET core installations on the remote host are affected by remote code execution vulnerability. Description: A remote code execution vulnerability exists in .NET Core 3.1, .NET 6.0, and .NET 7.0, where a malicious actor could cause a user to run arbitrary code as a result of parsing maliciously crafted xps files. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update .NET Core Runtime to version 3.1.32 or 6.0.12 or 7.0.1. See Also: https://support.microsoft.com/en-us/help/5021953 https://support.microsoft.com/en-us/help/5021954 https://support.microsoft.com/en-us/help/5021955 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.NetCore.App\3.1.16\ Installed version : 3.1.16.30112 Fixed version : 3.1.32 CPE: cpe:/a:microsoft:.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-12-13 Patch Publication Date: 2022-12-13 Plugin Modification Date: 2022-12-19 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-12-15 | 2023-01-09 04:28 | 2023-05-09 06:45 | 4
|  |  | 168826 |  | Security Updates for Microsoft ASP.NET Core (December 2022) | Update ASP.NET Core Runtime to version 3.1.32 or 6.0.12 or 7.0.1. | High | Synopsis: The Microsoft ASP.NET core installations on the remote host are affected by remote code execution vulnerability. Description: A remote code execution vulnerability exists in ASP.NET core 3.1, ASP.NET 6.0, and ASP.NET 7.0, where a malicious actor could cause a user to run arbitrary code as a result of parsing maliciously crafted xps files. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update ASP.NET Core Runtime to version 3.1.32 or 6.0.12 or 7.0.1. See Also: https://support.microsoft.com/en-us/help/5021953 https://support.microsoft.com/en-us/help/5021954 https://support.microsoft.com/en-us/help/5021955 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.AspNetCore.App\3.1.16 Installed version : 3.1.16 Fixed version : 3.1.32 CPE: cpe:/a:microsoft:asp.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-12-13 Patch Publication Date: 2022-12-13 Plugin Modification Date: 2022-12-16 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-12-15 | 2023-01-09 04:28 | 2023-05-09 06:45 | 4
|  |  | 171604 |  | Security Updates for Microsoft SQL Server (February 2023) | Microsoft has released the following security updates to address this issue: -KB5021126 -KB5021129 -KB5021522 -KB5021127 -KB5021045 -KB5021037 -KB5021128 -KB5021124 -KB5021125 -KB5020863 -KB5021112 -KB5021123 | High | Synopsis: The Microsoft SQL Server installation on the remote host is affected by multiple vulnerabilities. Description: The Microsoft SQL Server installation on the remote host is missing security updates. It is, therefore, affected by multiple vulnerabilities: - A remote code execution vulnerability. An attacker can exploit this to bypass authentication and execute unauthorized arbitrary commands. (CVE-2023-21528, CVE-2023-21568, CVE-2023-21704, CVE-2023-21705, CVE-2023-21713, CVE-2023-21718) Solution: Microsoft has released the following security updates to address this issue: -KB5021126 -KB5021129 -KB5021522 -KB5021127 -KB5021045 -KB5021037 -KB5021128 -KB5021124 -KB5021125 -KB5020863 -KB5021112 -KB5021123 See Also: https://support.microsoft.com/en-us/help/5020863 https://support.microsoft.com/en-us/help/5021112 https://support.microsoft.com/en-us/help/5021126 https://support.microsoft.com/en-us/help/5021129 https://support.microsoft.com/en-us/help/5021522 https://support.microsoft.com/en-us/help/5021127 https://support.microsoft.com/en-us/help/5021045 https://support.microsoft.com/en-us/help/5021037 https://support.microsoft.com/en-us/help/5021128 https://support.microsoft.com/en-us/help/5021123 https://support.microsoft.com/en-us/help/5021124 https://support.microsoft.com/en-us/help/5021125 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:S/C:C/I:C/A:C Plugin Output: KB : 5021126 - C:\Program Files\Microsoft SQL Server\MSSQL14.RALS_PLUS\MSSQL\Binn\sqlservr.exe has not been patched. Remote version : 2017.140.3391.2 Should be : 2017.140.3460.9 SQL Server Version : 14.0.3391.2 Express Edition SQL Server Instance : RALS_PLUS CPE: cpe:/a:microsoft:sql_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-02-14 Patch Publication Date: 2023-02-14 Plugin Modification Date: 2023-03-08 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2023-02-17 | 2023-03-07 00:45 | 2023-05-09 06:45 | 2
|  |  | 110291 |  | Bitvise SSH Server < 7.41 Multiple Vulnerabilities (remote) | Upgrade to Bitvise SSH Server 7.41 or later. | Medium | Synopsis: The SSH server running on the remote host is affected by multiple vulnerabilities. Description: According to its banner, the version of Bitvise SSH Server running on the remote host is prior to 7.41. It is, therefore, affected by multiple vulnerabilities. Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Bitvise SSH Server 7.41 or later. See Also: http://www.nessus.org/u?6bf2994b CVSS Base Score: 5.3 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:N/I:N/A:P Plugin Output: Version source : SSH-2.0-5.37 FlowSsh: Bitvise SSH Server (WinSSHD) 6.45 Installed version : 6.45 Fixed version : 7.41 CPE: cpe:/a:bitvise:ssh_server CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-05-18 Patch Publication Date: 2018-05-18 Plugin Modification Date: 2021-06-03 Exploitability Ease: NULL Plugin Type: remote Source File: | ScriptPro Server | SCRIPTPRO |  |  | Region X | VISN X | VAMC | [7 empty] | Microsoft Windows Server 2012 R2 Standard | tcp | 22 | Misc. | Repeat Vulnerability | 2018-06-01 | 2020-04-09 06:05 | 2023-05-09 02:04 | 37
|  |  | 51192 |  | SSL Certificate Cannot Be Trusted | Purchase or generate a proper SSL certificate for this service. | Medium | Synopsis: The SSL certificate for this service cannot be trusted. Description: The server's X.509 certificate cannot be trusted. This situation can occur in three different ways, in which the chain of trust can be broken, as stated below : - First, the top of the certificate chain sent by the server might not be descended from a known public certificate authority. This can occur either when the top of the chain is an unrecognized, self-signed certificate, or when intermediate certificates are missing that would connect the top of the certificate chain to a known public certificate authority. - Second, the certificate chain may contain a certificate that is not valid at the time of the scan. This can occur either when the scan occurs before one of the certificate's 'notBefore' dates, or after one of the certificate's 'notAfter' dates. - Third, the certificate chain may contain a signature that either didn't match the certificate's information or could not be verified. Bad signatures can be fixed by getting the certificate with the bad signature to be re-signed by its issuer. Signatures that could not be verified are the result of the certificate's issuer using a signing algorithm that Nessus either does not support or does not recognize. If the remote host is a public host in production, any break in the chain makes it more difficult for users to verify the authenticity and identity of the web server. This could make it easier to carry out man-in-the-middle attacks against the remote host. Solution: Purchase or generate a proper SSL certificate for this service. See Also: https://www.itu.int/rec/T-REC-X.509/en https://en.wikipedia.org/wiki/X.509 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:N Plugin Output: The following certificate was at the top of the certificate chain sent by the remote host, but it is signed by an unknown certificate authority : ¦-Subject : CN=VHADETAPPRALSC.v11.med.va.gov ¦-Issuer : CN=VHADETAPPRALSC.v11.med.va.gov CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2020-04-27 Exploitability Ease: NULL Plugin Type: remote Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 3389 | General | Repeat Vulnerability | 2010-12-15 | 2021-03-09 07:00 | 2023-05-09 06:58 | 26
|  |  | 132101 |  | Windows Speculative Execution Configuration Check | Apply vendor recommended settings. | Medium | Synopsis: The remote host has not properly mitigated a series of speculative execution vulnerabilities. Description: The remote host has not properly mitigated a series of known speculative execution vulnerabilities. It, therefore, may be affected by : - Branch Target Injection (BTI) (CVE-2017-5715) - Bounds Check Bypass (BCB) (CVE-2017-5753) - Rogue Data Cache Load (RDCL) (CVE-2017-5754) - Rogue System Register Read (RSRE) (CVE-2018-3640) - Speculative Store Bypass (SSB) (CVE-2018-3639) - L1 Terminal Fault (L1TF) (CVE-2018-3615, CVE-2018-3620, CVE-2018-3646) - Microarchitectural Data Sampling Uncacheable Memory (MDSUM) (CVE-2019-11091) - Microarchitectural Store Buffer Data Sampling (MSBDS) (CVE-2018-12126) - Microarchitectural Load Port Data Sampling (MLPDS) (CVE-2018-12127) - Microarchitectural Fill Buffer Data Sampling (MFBDS) (CVE-2018-12130) - TSX Asynchronous Abort (TAA) (CVE-2019-11135) Solution: Apply vendor recommended settings. See Also: http://www.nessus.org/u?8902cebb http://www.nessus.org/u?6a005ed4 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:P/A:N Plugin Output: Current Settings: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: Not Set - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: Not Set ----------------------------------- Recommended Settings 1: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00000048 (72) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading enabled. ----------------------------------- Recommended Settings 2: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00002048 (8264) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading disabled. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-01-04 Patch Publication Date: NULL Plugin Modification Date: 2023-05-03 Exploitability Ease: Exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2019-12-18 | 2022-02-06 21:22 | 2023-05-09 06:58 | 15
|  |  | 136946 |  | Windows 10 / Windows Server 2016 September 2017 Information Disclosure Vulnerability (CVE-2017-8529) | Refer to the Microsoft CVE article for additional information. | Medium | Synopsis: The remote Windows host is affected by an information disclosure vulnerability. Description: The remote Windows host is missing a security update or a registry setting required to enable protections for CVE-2017-8529. It is, therefore, affected by an information disclosure vulnerability: - An information disclosure vulnerability exists when affected Microsoft scripting engines do not properly handle objects in memory. The vulnerability could allow an attacker to detect specific files on the user's computer. In a web-based attack scenario, an attacker could host a website that is used to attempt to exploit the vulnerability. Solution: Refer to the Microsoft CVE article for additional information. See Also: http://www.nessus.org/u?1f6a3c24 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:N/A:N Plugin Output: The following registry key is required to enable the fix for CVE-2017-8529 and is missing. HKLM\SOFTWARE\Microsoft\Internet Explorer\Main\FeatureControl\FEATURE_ENABLE_PRINT_INFO_DISCLOSURE_FIX\iexplore.exe The following registry key is required to enable the fix for CVE-2017-8529 and is missing. HKLM\SOFTWARE\WOW6432Node\Microsoft\Internet Explorer\Main\FeatureControl\FEATURE_ENABLE_PRINT_INFO_DISCLOSURE_FIX\iexplore.exe CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2017-06-13 Patch Publication Date: 2017-06-13 Plugin Modification Date: 2020-06-04 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2020-05-28 | 2022-02-06 21:22 | 2023-05-09 06:58 | 15
|  |  | 152528 |  | Security Update for Microsoft ASP.NET Core (August 2021) | Update ASP.NET Core, remove vulnerable packages and refer to vendor advisory. | Medium | Synopsis: The Microsoft ASP.NET Core installations on the remote host is affected by multiple vulnerabilities. Description: The Microsoft ASP.NET Core installation on the remote host is version 2.1.x prior to 2.1.29, 3.1.x prior to 3.1.18, or 5.x prior to 5.0.9. It is, therefore, affected by multiple vulnerabilities: - A denial of service (DoS) vulnerability. An attacker can exploit this issue to cause the affected component to deny system or application services. (CVE-2021-26423) - An information disclosure vulnerability. An attacker can exploit this to disclose potentially sensitive information. (CVE-2021-34485, CVE-2021-34532) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update ASP.NET Core, remove vulnerable packages and refer to vendor advisory. See Also: https://dotnet.microsoft.com/download/dotnet-core/2.1 https://dotnet.microsoft.com/download/dotnet-core/3.1 https://dotnet.microsoft.com/download/dotnet/5.0 https://devblogs.microsoft.com/dotnet/net-august-2021 https://github.com/dotnet/announcements/issues/194 https://github.com/dotnet/announcements/issues/195 https://github.com/dotnet/announcements/issues/196 http://www.nessus.org/u?bb1ce96e http://www.nessus.org/u?0933ffe1 http://www.nessus.org/u?6242d65f CVSS Base Score: 5.5 CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:P/I:N/A:N Plugin Output: Path : C:\Program Files (x86)\dotnet\shared\Microsoft.AspNetCore.App\3.1.16 Installed version : 3.1.16 Fixed version : 3.1.18 CPE: cpe:/a:microsoft:asp.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-08-10 Patch Publication Date: 2021-08-10 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2021-08-12 | 2022-09-04 23:02 | 2023-05-09 06:58 | 8
|  |  | 157288 |  | TLS Version 1.1 Protocol Deprecated | Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. | Medium | Synopsis: The remote service encrypts traffic using an older version of TLS. Description: The remote service accepts connections encrypted using TLS 1.1. TLS 1.1 lacks support for current and recommended cipher suites. Ciphers that support encryption before MAC computation, and authenticated encryption modes such as GCM cannot be used with TLS 1.1 As of March 31, 2020, Endpoints that are not enabled for TLS 1.2 and higher will no longer function properly with major web browsers and major vendors. Solution: Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. See Also: https://datatracker.ietf.org/doc/html/rfc8996 http://www.nessus.org/u?c8ae820d CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:P/A:N Plugin Output: TLSv1.1 is enabled and the server supports at least one cipher. CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-04-19 Exploitability Ease: NULL Plugin Type: remote Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 3389 | Service detection | Repeat Vulnerability | 2022-04-04 | 2022-05-02 06:37 | 2023-05-09 06:58 | 12
|  |  | 158744 |  | Security Updates for Microsoft .NET core (March 2022) | Update .NET Core Runtime to version 3.1.23, 5.0.15 or 6.0.3. | Medium | Synopsis: The Microsoft .NET core installations on the remote host are affected by multiple vulnerabilities. Description: The Microsoft .NET core installations on the remote host are missing security updates. It is, therefore, affected by multiple vulnerabilities: - A denial of service (DoS) vulnerability. An attacker can exploit this issue to cause the affected component to deny system or application services. (CVE-2022-24464) - A remote code execution vulnerability. An attacker can exploit this to bypass authentication and execute unauthorized arbitrary commands. (CVE-2020-8927, CVE-2022-24512) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update .NET Core Runtime to version 3.1.23, 5.0.15 or 6.0.3. See Also: https://dotnet.microsoft.com/download/dotnet/3.1 https://dotnet.microsoft.com/download/dotnet/5.0 https://dotnet.microsoft.com/download/dotnet/6.0 https://github.com/dotnet/announcements/issues/210 http://www.nessus.org/u?56caba70 http://www.nessus.org/u?95177a8e http://www.nessus.org/u?96c2a71d CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.NetCore.App\3.1.16\ Installed version : 3.1.16.30112 Fixed version : 3.1.23 CPE: cpe:/a:microsoft:.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-08-27 Patch Publication Date: 2022-03-08 Plugin Modification Date: 2023-04-18 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Misc. | Repeat Vulnerability | 2022-03-09 | 2022-04-04 01:54 | 2023-05-09 06:58 | 13
|  |  | 162314 |  | Security Updates for Microsoft .NET core (June 2022) | Update .NET Core Runtime to version 3.1.26 or 6.0.6. | Medium | Synopsis: The Microsoft .NET core installations on the remote host are affected by an information disclosure vulnerability. Description: An information disclosure vulnerability exists in .NET core 6.0 < 6.0.6 and .NET Core 3.1 < 3.1.26. An unauthenticated, local attacker can exploit this, to disclose potentially sensitive information. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update .NET Core Runtime to version 3.1.26 or 6.0.6. See Also: https://dotnet.microsoft.com/download/dotnet/3.1 https://dotnet.microsoft.com/download/dotnet/6.0 https://github.com/dotnet/announcements/issues/225 http://www.nessus.org/u?bfb8ea98 CVSS Base Score: 5.5 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:N/A:N Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.NetCore.App\3.1.16\ Installed version : 3.1.16.30112 Fixed version : 3.1.26 CPE: cpe:/a:microsoft:.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-06-14 Patch Publication Date: 2022-06-14 Plugin Modification Date: 2023-03-23 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-06-16 | 2022-07-10 08:29 | 2023-05-09 06:58 | 10
|  |  | 163974 |  | Security Updates for Microsoft .NET Core (August 2022) | Update .NET Core Runtime to version 3.1.28 or 6.0.8. | Medium | Synopsis: The Microsoft .NET core installations on the remote host are affected by a spoofing vulnerability. Description: A spoofing vulnerability exists in .NET core 6.0 < 6.0.8 and .NET Core 3.1 < 3.1.28. An unauthenticated, remote attacker can exploit this, to perform actions with the privileges of another user. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update .NET Core Runtime to version 3.1.28 or 6.0.8. See Also: https://support.microsoft.com/help/5016987 https://support.microsoft.com/help/5016990 https://dotnet.microsoft.com/download/dotnet/3.1 https://dotnet.microsoft.com/download/dotnet/6.0 http://www.nessus.org/u?327bb1fb http://www.nessus.org/u?7ce182ee https://github.com/dotnet/core/issues/7682 CVSS Base Score: 5.9 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:N/A:N Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.NetCore.App\3.1.16\ Installed version : 3.1.16.30112 Fixed version : 3.1.28 CPE: cpe:/a:microsoft:.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-08-09 Patch Publication Date: 2022-08-09 Plugin Modification Date: 2022-12-07 Exploitability Ease: Exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-08-10 | 2022-09-04 23:02 | 2023-05-09 06:58 | 8
|  |  | 170690 |  | Microsoft Edge (Chromium) < 109.0.1518.61 Security Feature Bypass (CVE-2023-21719) | Upgrade to Microsoft Edge version 109.0.1518.61 or later. | Medium | Synopsis: The remote host has an web browser installed that is affected by security feature bypass. Description: The version of Microsoft Edge installed on the remote Windows host is prior to 109.0.1518.61. It is, therefore, affected by a vulnerability as referenced in the January 19, 2023 advisory. - Microsoft Edge (Chromium-based) Security Feature Bypass Vulnerability. (CVE-2023-21719) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Microsoft Edge version 109.0.1518.61 or later. See Also: http://www.nessus.org/u?245dfb65 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-21719 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:N/A:N Plugin Output: Path : C:\Program Files (x86)\Microsoft\Edge\Application Installed version : 109.0.1518.55 Fixed version : 109.0.1518.61 CPE: cpe:/a:microsoft:edge CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-01-19 Patch Publication Date: 2023-01-19 Plugin Modification Date: 2023-02-10 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2023-01-27 | 2023-03-07 01:04 | 2023-05-09 06:58 | 2
|  |  | 171334 |  | Microsoft Edge (Chromium) < 109.0.1518.78 Tampering (CVE-2023-21720) | Upgrade to Microsoft Edge version 109.0.1518.78 or later. | Medium | Synopsis: The remote host has an web browser installed that is affected by tampering. Description: The version of Microsoft Edge installed on the remote Windows host is prior to 109.0.1518.78. It is, therefore, affected by a vulnerability as referenced in the February 2, 2023 advisory. - Microsoft Edge (Chromium-based) Tampering Vulnerability (CVE-2023-21720) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Microsoft Edge version 109.0.1518.78 or later. See Also: http://www.nessus.org/u?245dfb65 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-21720 CVSS Base Score: 5.3 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:N/I:N/A:C Plugin Output: Path : C:\Program Files (x86)\Microsoft\Edge\Application Installed version : 109.0.1518.55 Fixed version : 109.0.1518.78 CPE: cpe:/a:microsoft:edge CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-02-02 Patch Publication Date: 2023-02-03 Plugin Modification Date: 2023-02-16 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2023-02-10 | 2023-03-07 01:04 | 2023-05-09 06:58 | 2
|  |  | 51192 |  | SSL Certificate Cannot Be Trusted | Purchase or generate a proper SSL certificate for this service. | Medium | Synopsis: The SSL certificate for this service cannot be trusted. Description: The server's X.509 certificate cannot be trusted. This situation can occur in three different ways, in which the chain of trust can be broken, as stated below : - First, the top of the certificate chain sent by the server might not be descended from a known public certificate authority. This can occur either when the top of the chain is an unrecognized, self-signed certificate, or when intermediate certificates are missing that would connect the top of the certificate chain to a known public certificate authority. - Second, the certificate chain may contain a certificate that is not valid at the time of the scan. This can occur either when the scan occurs before one of the certificate's 'notBefore' dates, or after one of the certificate's 'notAfter' dates. - Third, the certificate chain may contain a signature that either didn't match the certificate's information or could not be verified. Bad signatures can be fixed by getting the certificate with the bad signature to be re-signed by its issuer. Signatures that could not be verified are the result of the certificate's issuer using a signing algorithm that Nessus either does not support or does not recognize. If the remote host is a public host in production, any break in the chain makes it more difficult for users to verify the authenticity and identity of the web server. This could make it easier to carry out man-in-the-middle attacks against the remote host. Solution: Purchase or generate a proper SSL certificate for this service. See Also: https://www.itu.int/rec/T-REC-X.509/en https://en.wikipedia.org/wiki/X.509 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:N Plugin Output: The following certificate was at the top of the certificate chain sent by the remote host, but it is signed by an unknown certificate authority : ¦-Subject : CN=VHADETCCAOMNICP.v11.med.va.gov ¦-Issuer : CN=VHADETCCAOMNICP.v11.med.va.gov CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2020-04-27 Exploitability Ease: NULL Plugin Type: remote Source File: | Omnicell Omnicenter Prod Server | OMNICELL INC |  |  | Region X | VISN X | VAMC | VHAXXXCCAOMNICP |  |  |  |  | VHAXXXCCAOMNICP.vxx.med.va.gov |  | Linux Kernel 3.2 | tcp | 3389 | General | Repeat Vulnerability | 2010-12-15 | 2022-12-05 02:04 | 2023-05-09 00:32 | 5
|  |  | 157288 |  | TLS Version 1.1 Protocol Deprecated | Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. | Medium | Synopsis: The remote service encrypts traffic using an older version of TLS. Description: The remote service accepts connections encrypted using TLS 1.1. TLS 1.1 lacks support for current and recommended cipher suites. Ciphers that support encryption before MAC computation, and authenticated encryption modes such as GCM cannot be used with TLS 1.1 As of March 31, 2020, Endpoints that are not enabled for TLS 1.2 and higher will no longer function properly with major web browsers and major vendors. Solution: Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. See Also: https://datatracker.ietf.org/doc/html/rfc8996 http://www.nessus.org/u?c8ae820d CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:P/A:N Plugin Output: TLSv1.1 is enabled and the server supports at least one cipher. CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-04-19 Exploitability Ease: NULL Plugin Type: remote Source File: | Omnicell Omnicenter Prod Server | OMNICELL INC |  |  | Region X | VISN X | VAMC | VHAXXXCCAOMNICP |  |  |  |  | VHAXXXCCAOMNICP.vxx.med.va.gov |  | Linux Kernel 3.2 | tcp | 3389 | Service detection | Repeat Vulnerability | 2022-04-04 | 2022-12-05 02:04 | 2023-05-09 00:32 | 5
|  |  | 51192 |  | SSL Certificate Cannot Be Trusted | Purchase or generate a proper SSL certificate for this service. | Medium | Synopsis: The SSL certificate for this service cannot be trusted. Description: The server's X.509 certificate cannot be trusted. This situation can occur in three different ways, in which the chain of trust can be broken, as stated below : - First, the top of the certificate chain sent by the server might not be descended from a known public certificate authority. This can occur either when the top of the chain is an unrecognized, self-signed certificate, or when intermediate certificates are missing that would connect the top of the certificate chain to a known public certificate authority. - Second, the certificate chain may contain a certificate that is not valid at the time of the scan. This can occur either when the scan occurs before one of the certificate's 'notBefore' dates, or after one of the certificate's 'notAfter' dates. - Third, the certificate chain may contain a signature that either didn't match the certificate's information or could not be verified. Bad signatures can be fixed by getting the certificate with the bad signature to be re-signed by its issuer. Signatures that could not be verified are the result of the certificate's issuer using a signing algorithm that Nessus either does not support or does not recognize. If the remote host is a public host in production, any break in the chain makes it more difficult for users to verify the authenticity and identity of the web server. This could make it easier to carry out man-in-the-middle attacks against the remote host. Solution: Purchase or generate a proper SSL certificate for this service. See Also: https://www.itu.int/rec/T-REC-X.509/en https://en.wikipedia.org/wiki/X.509 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:N Plugin Output: The following certificate was at the top of the certificate chain sent by the remote host, but it is signed by an unknown certificate authority : ¦-Subject : CN=VHADETCCARALSD.v11.med.va.gov ¦-Issuer : CN=VHADETCCARALSD.v11.med.va.gov CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2020-04-27 Exploitability Ease: NULL Plugin Type: remote Source File: | RALS I-Stat Data Exchange (DE) Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSD |  |  |  |  | VHAXXXCCARALSD.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 3389 | General | Repeat Vulnerability | 2010-12-15 | 2021-09-05 13:04 | 2023-05-09 06:44 | 20
|  |  | 132101 |  | Windows Speculative Execution Configuration Check | Apply vendor recommended settings. | Medium | Synopsis: The remote host has not properly mitigated a series of speculative execution vulnerabilities. Description: The remote host has not properly mitigated a series of known speculative execution vulnerabilities. It, therefore, may be affected by : - Branch Target Injection (BTI) (CVE-2017-5715) - Bounds Check Bypass (BCB) (CVE-2017-5753) - Rogue Data Cache Load (RDCL) (CVE-2017-5754) - Rogue System Register Read (RSRE) (CVE-2018-3640) - Speculative Store Bypass (SSB) (CVE-2018-3639) - L1 Terminal Fault (L1TF) (CVE-2018-3615, CVE-2018-3620, CVE-2018-3646) - Microarchitectural Data Sampling Uncacheable Memory (MDSUM) (CVE-2019-11091) - Microarchitectural Store Buffer Data Sampling (MSBDS) (CVE-2018-12126) - Microarchitectural Load Port Data Sampling (MLPDS) (CVE-2018-12127) - Microarchitectural Fill Buffer Data Sampling (MFBDS) (CVE-2018-12130) - TSX Asynchronous Abort (TAA) (CVE-2019-11135) Solution: Apply vendor recommended settings. See Also: http://www.nessus.org/u?8902cebb http://www.nessus.org/u?6a005ed4 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:P/A:N Plugin Output: Current Settings: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: Not Set - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: Not Set ----------------------------------- Recommended Settings 1: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00000048 (72) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading enabled. ----------------------------------- Recommended Settings 2: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00002048 (8264) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading disabled. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-01-04 Patch Publication Date: NULL Plugin Modification Date: 2023-05-03 Exploitability Ease: Exploits are available Plugin Type: local Source File: | RALS I-Stat Data Exchange (DE) Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSD |  |  |  |  | VHAXXXCCARALSD.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2019-12-18 | 2022-02-06 21:14 | 2023-05-09 06:44 | 15
|  |  | 136946 |  | Windows 10 / Windows Server 2016 September 2017 Information Disclosure Vulnerability (CVE-2017-8529) | Refer to the Microsoft CVE article for additional information. | Medium | Synopsis: The remote Windows host is affected by an information disclosure vulnerability. Description: The remote Windows host is missing a security update or a registry setting required to enable protections for CVE-2017-8529. It is, therefore, affected by an information disclosure vulnerability: - An information disclosure vulnerability exists when affected Microsoft scripting engines do not properly handle objects in memory. The vulnerability could allow an attacker to detect specific files on the user's computer. In a web-based attack scenario, an attacker could host a website that is used to attempt to exploit the vulnerability. Solution: Refer to the Microsoft CVE article for additional information. See Also: http://www.nessus.org/u?1f6a3c24 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:N/A:N Plugin Output: The following registry key is required to enable the fix for CVE-2017-8529 and is missing. HKLM\SOFTWARE\Microsoft\Internet Explorer\Main\FeatureControl\FEATURE_ENABLE_PRINT_INFO_DISCLOSURE_FIX\iexplore.exe The following registry key is required to enable the fix for CVE-2017-8529 and is missing. HKLM\SOFTWARE\WOW6432Node\Microsoft\Internet Explorer\Main\FeatureControl\FEATURE_ENABLE_PRINT_INFO_DISCLOSURE_FIX\iexplore.exe CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2017-06-13 Patch Publication Date: 2017-06-13 Plugin Modification Date: 2020-06-04 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS I-Stat Data Exchange (DE) Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSD |  |  |  |  | VHAXXXCCARALSD.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2020-05-28 | 2022-02-06 21:14 | 2023-05-09 06:44 | 15
|  |  | 157288 |  | TLS Version 1.1 Protocol Deprecated | Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. | Medium | Synopsis: The remote service encrypts traffic using an older version of TLS. Description: The remote service accepts connections encrypted using TLS 1.1. TLS 1.1 lacks support for current and recommended cipher suites. Ciphers that support encryption before MAC computation, and authenticated encryption modes such as GCM cannot be used with TLS 1.1 As of March 31, 2020, Endpoints that are not enabled for TLS 1.2 and higher will no longer function properly with major web browsers and major vendors. Solution: Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. See Also: https://datatracker.ietf.org/doc/html/rfc8996 http://www.nessus.org/u?c8ae820d CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:P/A:N Plugin Output: TLSv1.1 is enabled and the server supports at least one cipher. CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-04-19 Exploitability Ease: NULL Plugin Type: remote Source File: | RALS I-Stat Data Exchange (DE) Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSD |  |  |  |  | VHAXXXCCARALSD.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 3389 | Service detection | Repeat Vulnerability | 2022-04-04 | 2022-05-02 06:26 | 2023-05-09 06:44 | 12
|  |  | 51192 |  | SSL Certificate Cannot Be Trusted | Purchase or generate a proper SSL certificate for this service. | Medium | Synopsis: The SSL certificate for this service cannot be trusted. Description: The server's X.509 certificate cannot be trusted. This situation can occur in three different ways, in which the chain of trust can be broken, as stated below : - First, the top of the certificate chain sent by the server might not be descended from a known public certificate authority. This can occur either when the top of the chain is an unrecognized, self-signed certificate, or when intermediate certificates are missing that would connect the top of the certificate chain to a known public certificate authority. - Second, the certificate chain may contain a certificate that is not valid at the time of the scan. This can occur either when the scan occurs before one of the certificate's 'notBefore' dates, or after one of the certificate's 'notAfter' dates. - Third, the certificate chain may contain a signature that either didn't match the certificate's information or could not be verified. Bad signatures can be fixed by getting the certificate with the bad signature to be re-signed by its issuer. Signatures that could not be verified are the result of the certificate's issuer using a signing algorithm that Nessus either does not support or does not recognize. If the remote host is a public host in production, any break in the chain makes it more difficult for users to verify the authenticity and identity of the web server. This could make it easier to carry out man-in-the-middle attacks against the remote host. Solution: Purchase or generate a proper SSL certificate for this service. See Also: https://www.itu.int/rec/T-REC-X.509/en https://en.wikipedia.org/wiki/X.509 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:N Plugin Output: The following certificate was at the top of the certificate chain sent by the remote host, but it is signed by an unknown certificate authority : ¦-Subject : CN=VHADETCCARALST.v11.med.va.gov ¦-Issuer : CN=VHADETCCARALST.v11.med.va.gov CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2020-04-27 Exploitability Ease: NULL Plugin Type: remote Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 3389 | General | Repeat Vulnerability | 2010-12-15 | 2021-09-05 13:04 | 2023-05-09 06:45 | 20
|  |  | 132101 |  | Windows Speculative Execution Configuration Check | Apply vendor recommended settings. | Medium | Synopsis: The remote host has not properly mitigated a series of speculative execution vulnerabilities. Description: The remote host has not properly mitigated a series of known speculative execution vulnerabilities. It, therefore, may be affected by : - Branch Target Injection (BTI) (CVE-2017-5715) - Bounds Check Bypass (BCB) (CVE-2017-5753) - Rogue Data Cache Load (RDCL) (CVE-2017-5754) - Rogue System Register Read (RSRE) (CVE-2018-3640) - Speculative Store Bypass (SSB) (CVE-2018-3639) - L1 Terminal Fault (L1TF) (CVE-2018-3615, CVE-2018-3620, CVE-2018-3646) - Microarchitectural Data Sampling Uncacheable Memory (MDSUM) (CVE-2019-11091) - Microarchitectural Store Buffer Data Sampling (MSBDS) (CVE-2018-12126) - Microarchitectural Load Port Data Sampling (MLPDS) (CVE-2018-12127) - Microarchitectural Fill Buffer Data Sampling (MFBDS) (CVE-2018-12130) - TSX Asynchronous Abort (TAA) (CVE-2019-11135) Solution: Apply vendor recommended settings. See Also: http://www.nessus.org/u?8902cebb http://www.nessus.org/u?6a005ed4 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:P/A:N Plugin Output: Current Settings: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: Not Set - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: Not Set ----------------------------------- Recommended Settings 1: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00000048 (72) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading enabled. ----------------------------------- Recommended Settings 2: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00002048 (8264) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading disabled. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-01-04 Patch Publication Date: NULL Plugin Modification Date: 2023-05-03 Exploitability Ease: Exploits are available Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2019-12-18 | 2022-02-06 21:14 | 2023-05-09 06:45 | 15
|  |  | 136946 |  | Windows 10 / Windows Server 2016 September 2017 Information Disclosure Vulnerability (CVE-2017-8529) | Refer to the Microsoft CVE article for additional information. | Medium | Synopsis: The remote Windows host is affected by an information disclosure vulnerability. Description: The remote Windows host is missing a security update or a registry setting required to enable protections for CVE-2017-8529. It is, therefore, affected by an information disclosure vulnerability: - An information disclosure vulnerability exists when affected Microsoft scripting engines do not properly handle objects in memory. The vulnerability could allow an attacker to detect specific files on the user's computer. In a web-based attack scenario, an attacker could host a website that is used to attempt to exploit the vulnerability. Solution: Refer to the Microsoft CVE article for additional information. See Also: http://www.nessus.org/u?1f6a3c24 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:N/A:N Plugin Output: The following registry key is required to enable the fix for CVE-2017-8529 and is missing. HKLM\SOFTWARE\Microsoft\Internet Explorer\Main\FeatureControl\FEATURE_ENABLE_PRINT_INFO_DISCLOSURE_FIX\iexplore.exe The following registry key is required to enable the fix for CVE-2017-8529 and is missing. HKLM\SOFTWARE\WOW6432Node\Microsoft\Internet Explorer\Main\FeatureControl\FEATURE_ENABLE_PRINT_INFO_DISCLOSURE_FIX\iexplore.exe CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2017-06-13 Patch Publication Date: 2017-06-13 Plugin Modification Date: 2020-06-04 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2020-05-28 | 2022-02-06 21:14 | 2023-05-09 06:45 | 15
|  |  | 152528 |  | Security Update for Microsoft ASP.NET Core (August 2021) | Update ASP.NET Core, remove vulnerable packages and refer to vendor advisory. | Medium | Synopsis: The Microsoft ASP.NET Core installations on the remote host is affected by multiple vulnerabilities. Description: The Microsoft ASP.NET Core installation on the remote host is version 2.1.x prior to 2.1.29, 3.1.x prior to 3.1.18, or 5.x prior to 5.0.9. It is, therefore, affected by multiple vulnerabilities: - A denial of service (DoS) vulnerability. An attacker can exploit this issue to cause the affected component to deny system or application services. (CVE-2021-26423) - An information disclosure vulnerability. An attacker can exploit this to disclose potentially sensitive information. (CVE-2021-34485, CVE-2021-34532) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update ASP.NET Core, remove vulnerable packages and refer to vendor advisory. See Also: https://dotnet.microsoft.com/download/dotnet-core/2.1 https://dotnet.microsoft.com/download/dotnet-core/3.1 https://dotnet.microsoft.com/download/dotnet/5.0 https://devblogs.microsoft.com/dotnet/net-august-2021 https://github.com/dotnet/announcements/issues/194 https://github.com/dotnet/announcements/issues/195 https://github.com/dotnet/announcements/issues/196 http://www.nessus.org/u?bb1ce96e http://www.nessus.org/u?0933ffe1 http://www.nessus.org/u?6242d65f CVSS Base Score: 5.5 CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:P/I:N/A:N Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.AspNetCore.App\3.1.16 Installed version : 3.1.16 Fixed version : 3.1.18 CPE: cpe:/a:microsoft:asp.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2021-08-10 Patch Publication Date: 2021-08-10 Plugin Modification Date: 2022-04-11 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2021-08-12 | 2022-09-04 22:51 | 2023-05-09 06:45 | 8
|  |  | 157288 |  | TLS Version 1.1 Protocol Deprecated | Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. | Medium | Synopsis: The remote service encrypts traffic using an older version of TLS. Description: The remote service accepts connections encrypted using TLS 1.1. TLS 1.1 lacks support for current and recommended cipher suites. Ciphers that support encryption before MAC computation, and authenticated encryption modes such as GCM cannot be used with TLS 1.1 As of March 31, 2020, Endpoints that are not enabled for TLS 1.2 and higher will no longer function properly with major web browsers and major vendors. Solution: Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. See Also: https://datatracker.ietf.org/doc/html/rfc8996 http://www.nessus.org/u?c8ae820d CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:P/A:N Plugin Output: TLSv1.1 is enabled and the server supports at least one cipher. CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-04-19 Exploitability Ease: NULL Plugin Type: remote Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 3389 | Service detection | Repeat Vulnerability | 2022-04-04 | 2022-05-02 06:26 | 2023-05-09 06:45 | 12
|  |  | 158744 |  | Security Updates for Microsoft .NET core (March 2022) | Update .NET Core Runtime to version 3.1.23, 5.0.15 or 6.0.3. | Medium | Synopsis: The Microsoft .NET core installations on the remote host are affected by multiple vulnerabilities. Description: The Microsoft .NET core installations on the remote host are missing security updates. It is, therefore, affected by multiple vulnerabilities: - A denial of service (DoS) vulnerability. An attacker can exploit this issue to cause the affected component to deny system or application services. (CVE-2022-24464) - A remote code execution vulnerability. An attacker can exploit this to bypass authentication and execute unauthorized arbitrary commands. (CVE-2020-8927, CVE-2022-24512) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update .NET Core Runtime to version 3.1.23, 5.0.15 or 6.0.3. See Also: https://dotnet.microsoft.com/download/dotnet/3.1 https://dotnet.microsoft.com/download/dotnet/5.0 https://dotnet.microsoft.com/download/dotnet/6.0 https://github.com/dotnet/announcements/issues/210 http://www.nessus.org/u?56caba70 http://www.nessus.org/u?95177a8e http://www.nessus.org/u?96c2a71d CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.NetCore.App\3.1.16\ Installed version : 3.1.16.30112 Fixed version : 3.1.23 CPE: cpe:/a:microsoft:.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2020-08-27 Patch Publication Date: 2022-03-08 Plugin Modification Date: 2023-04-18 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Misc. | Repeat Vulnerability | 2022-03-09 | 2022-04-04 01:40 | 2023-05-09 06:45 | 13
|  |  | 162314 |  | Security Updates for Microsoft .NET core (June 2022) | Update .NET Core Runtime to version 3.1.26 or 6.0.6. | Medium | Synopsis: The Microsoft .NET core installations on the remote host are affected by an information disclosure vulnerability. Description: An information disclosure vulnerability exists in .NET core 6.0 < 6.0.6 and .NET Core 3.1 < 3.1.26. An unauthenticated, local attacker can exploit this, to disclose potentially sensitive information. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update .NET Core Runtime to version 3.1.26 or 6.0.6. See Also: https://dotnet.microsoft.com/download/dotnet/3.1 https://dotnet.microsoft.com/download/dotnet/6.0 https://github.com/dotnet/announcements/issues/225 http://www.nessus.org/u?bfb8ea98 CVSS Base Score: 5.5 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:N/A:N Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.NetCore.App\3.1.16\ Installed version : 3.1.16.30112 Fixed version : 3.1.26 CPE: cpe:/a:microsoft:.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-06-14 Patch Publication Date: 2022-06-14 Plugin Modification Date: 2023-03-23 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-06-16 | 2022-07-10 08:10 | 2023-05-09 06:45 | 10
|  |  | 163974 |  | Security Updates for Microsoft .NET Core (August 2022) | Update .NET Core Runtime to version 3.1.28 or 6.0.8. | Medium | Synopsis: The Microsoft .NET core installations on the remote host are affected by a spoofing vulnerability. Description: A spoofing vulnerability exists in .NET core 6.0 < 6.0.8 and .NET Core 3.1 < 3.1.28. An unauthenticated, remote attacker can exploit this, to perform actions with the privileges of another user. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Update .NET Core Runtime to version 3.1.28 or 6.0.8. See Also: https://support.microsoft.com/help/5016987 https://support.microsoft.com/help/5016990 https://dotnet.microsoft.com/download/dotnet/3.1 https://dotnet.microsoft.com/download/dotnet/6.0 http://www.nessus.org/u?327bb1fb http://www.nessus.org/u?7ce182ee https://github.com/dotnet/core/issues/7682 CVSS Base Score: 5.9 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:N/A:N Plugin Output: Path : C:\Program Files\dotnet\shared\Microsoft.NetCore.App\3.1.16\ Installed version : 3.1.16.30112 Fixed version : 3.1.28 CPE: cpe:/a:microsoft:.net_core CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-08-09 Patch Publication Date: 2022-08-09 Plugin Modification Date: 2022-12-07 Exploitability Ease: Exploits are available Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2022-08-10 | 2022-09-04 22:51 | 2023-05-09 06:45 | 8
|  |  | 58134 |  | Microsoft Silverlight Unsupported Version Detection (Windows) | Remove Microsoft Silverlight. | Critical | Synopsis: The remote host has an unsupported version of Microsoft Silverlight. Description: The installation of Microsoft Silverlight on the Windows host is no longer supported. Lack of support implies that no new security patches for the product will be released by the vendor. As a result, it is likely to contain security vulnerabilities. Solution: Remove Microsoft Silverlight. See Also: http://www.nessus.org/u?fcc00b67 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : c:\Program Files\Microsoft Silverlight\5.1.50918.0 Installed version : 5.1.50918.0 CPE: cpe:/a:microsoft:silverlight CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2022-05-23 Exploitability Ease: NULL Plugin Type: local Source File: | RALS I-Stat Data Exchange (DE) Server - Production | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSDE |  |  |  |  | VHAXXXCCARALSDE.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2012-02-27 | 2022-07-10 08:10 | 2023-05-09 06:43 | 10
|  |  | 63155 |  | Microsoft Windows Unquoted Service Path Enumeration | Ensure that any services that contain a space in the path enclose the path in quotes. | High | Synopsis: The remote Windows host has at least one service installed that uses an unquoted service path. Description: The remote Windows host has at least one service installed that uses an unquoted service path, which contains at least one whitespace. A local attacker can gain elevated privileges by inserting an executable file in the path of the affected service. Note that this is a generic test that will flag any application affected by the described vulnerability. Solution: Ensure that any services that contain a space in the path enclose the path in quotes. See Also: http://www.nessus.org/u?84a4cc1c http://cwe.mitre.org/data/definitions/428.html https://www.commonexploits.com/unquoted-service-paths/ http://www.nessus.org/u?4aa6acbc CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:C/A:C Plugin Output: Nessus found the following services with an untrusted path : SQLANYs_iSTATDEDB : C:\Program Files (x86)\SQL Anywhere 17\bin32\dbsrv17.exe ServiceDeCom : C:\Program Files (x86)\APOC\DE\Bin\DeComSrv.exe i-STATDeDpSrv : C:\Program Files (x86)\APOC\DE\Bin\DeDpSrv.exe iSTATeVAS : C:\Program Files (x86)\APOC\DE\Bin\evassrv.exe DeUpdateMonitor : C:\Program Files (x86)\APOC\DE\Bin\DeUpdateMonitor.exe CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2012-09-15 Patch Publication Date: NULL Plugin Modification Date: 2020-06-12 Exploitability Ease: Exploits are available Plugin Type: local Source File: | RALS I-Stat Data Exchange (DE) Server - Production | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSDE |  |  |  |  | VHAXXXCCARALSDE.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2012-12-05 | 2022-07-10 08:10 | 2023-05-09 06:43 | 10
|  |  | 161605 |  | VMware Tools 10.x / 11.x / 12.x < 12.0.5 XXE (VMSA-2022-0015) | Upgrade to VMware Tools version 12.0.5 or later. | High | Synopsis: A virtualization tool suite is installed on the remote Windows host is affected by an XML External Entity vulnerability. Description: The version of VMware Tools installed on the remote Windows host is affected by an XML External Entity (XXE) vulnerability. A malicious actor with non-administrative local user privileges in the Windows guest OS, where VMware Tools is installed, may exploit this issue leading to a denial-of-service condition or unintended information disclosure. Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to VMware Tools version 12.0.5 or later. See Also: https://www.vmware.com/security/advisories/VMSA-2022-0015.html CVSS Base Score: 7.1 CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:P/I:N/A:P Plugin Output: Path : C:\Program Files\VMware\VMware Tools\ Installed version : 11.3.5.31214 Fixed version : 12.0.5 CPE: cpe:/a:vmware:tools CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-05-24 Patch Publication Date: 2022-05-24 Plugin Modification Date: 2022-08-30 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS I-Stat Data Exchange (DE) Server - Production | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSDE |  |  |  |  | VHAXXXCCARALSDE.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2022-05-27 | 2022-09-04 22:51 | 2023-05-09 06:43 | 8
|  |  | 164422 |  | VMware Tools 11.x / 12.x < 12.1.0 Privilege Escalation (VMSA-2022-0024) | Upgrade to VMware Tools version 12.1.0 or later. | High | Synopsis: A virtualization tool suite is installed on the remote Windows host is affected by a privilege escalation vulnerability. Description: The version of VMware Tools installed on the remote Windows host is affected by a privilege escalation vulnerability. A malicious actor with local non-administrative access to the Guest OS can escalate privileges as a root user in the virtual machine. Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to VMware Tools version 12.1.0 or later. See Also: https://www.vmware.com/security/advisories/VMSA-2022-0024.html CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:L/Au:S/C:C/I:C/A:C Plugin Output: Path : C:\Program Files\VMware\VMware Tools\ Installed version : 11.3.5.31214 Fixed version : 12.1.0 CPE: cpe:/a:vmware:tools CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2022-08-23 Patch Publication Date: 2022-08-23 Plugin Modification Date: 2022-12-06 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS I-Stat Data Exchange (DE) Server - Production | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSDE |  |  |  |  | VHAXXXCCARALSDE.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2022-08-25 | 2022-09-04 22:51 | 2023-05-09 06:43 | 8
|  |  | 51192 |  | SSL Certificate Cannot Be Trusted | Purchase or generate a proper SSL certificate for this service. | Medium | Synopsis: The SSL certificate for this service cannot be trusted. Description: The server's X.509 certificate cannot be trusted. This situation can occur in three different ways, in which the chain of trust can be broken, as stated below : - First, the top of the certificate chain sent by the server might not be descended from a known public certificate authority. This can occur either when the top of the chain is an unrecognized, self-signed certificate, or when intermediate certificates are missing that would connect the top of the certificate chain to a known public certificate authority. - Second, the certificate chain may contain a certificate that is not valid at the time of the scan. This can occur either when the scan occurs before one of the certificate's 'notBefore' dates, or after one of the certificate's 'notAfter' dates. - Third, the certificate chain may contain a signature that either didn't match the certificate's information or could not be verified. Bad signatures can be fixed by getting the certificate with the bad signature to be re-signed by its issuer. Signatures that could not be verified are the result of the certificate's issuer using a signing algorithm that Nessus either does not support or does not recognize. If the remote host is a public host in production, any break in the chain makes it more difficult for users to verify the authenticity and identity of the web server. This could make it easier to carry out man-in-the-middle attacks against the remote host. Solution: Purchase or generate a proper SSL certificate for this service. See Also: https://www.itu.int/rec/T-REC-X.509/en https://en.wikipedia.org/wiki/X.509 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:N Plugin Output: The following certificate was at the top of the certificate chain sent by the remote host, but it is signed by an unknown certificate authority : ¦-Subject : CN=VHADETCCARALSDE.v11.med.va.gov ¦-Issuer : CN=VHADETCCARALSDE.v11.med.va.gov CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2020-04-27 Exploitability Ease: NULL Plugin Type: remote Source File: | RALS I-Stat Data Exchange (DE) Server - Production | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSDE |  |  |  |  | VHAXXXCCARALSDE.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 3389 | General | Repeat Vulnerability | 2010-12-15 | 2022-04-04 01:46 | 2023-05-09 06:43 | 13
|  |  | 57582 |  | SSL Self-Signed Certificate | Purchase or generate a proper SSL certificate for this service. | Medium | Synopsis: The SSL certificate chain for this service ends in an unrecognized self-signed certificate. Description: The X.509 certificate chain for this service is not signed by a recognized certificate authority. If the remote host is a public host in production, this nullifies the use of SSL as anyone could establish a man-in-the-middle attack against the remote host. Note that this plugin does not check for certificate chains that end in a certificate that is not self-signed, but is signed by an unrecognized certificate authority. Solution: Purchase or generate a proper SSL certificate for this service. See Also: NULL CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:N Plugin Output: The following certificate was found at the top of the certificate chain sent by the remote host, but is self-signed and was not found in the list of known certificate authorities : ¦-Subject : CN=VHADETCCARALSDE.v11.med.va.gov CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2022-06-14 Exploitability Ease: NULL Plugin Type: remote Source File: | RALS I-Stat Data Exchange (DE) Server - Production | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSDE |  |  |  |  | VHAXXXCCARALSDE.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 3389 | General | Repeat Vulnerability | 2012-01-17 | 2022-04-04 01:46 | 2023-05-09 06:43 | 13
|  |  | 132101 |  | Windows Speculative Execution Configuration Check | Apply vendor recommended settings. | Medium | Synopsis: The remote host has not properly mitigated a series of speculative execution vulnerabilities. Description: The remote host has not properly mitigated a series of known speculative execution vulnerabilities. It, therefore, may be affected by : - Branch Target Injection (BTI) (CVE-2017-5715) - Bounds Check Bypass (BCB) (CVE-2017-5753) - Rogue Data Cache Load (RDCL) (CVE-2017-5754) - Rogue System Register Read (RSRE) (CVE-2018-3640) - Speculative Store Bypass (SSB) (CVE-2018-3639) - L1 Terminal Fault (L1TF) (CVE-2018-3615, CVE-2018-3620, CVE-2018-3646) - Microarchitectural Data Sampling Uncacheable Memory (MDSUM) (CVE-2019-11091) - Microarchitectural Store Buffer Data Sampling (MSBDS) (CVE-2018-12126) - Microarchitectural Load Port Data Sampling (MLPDS) (CVE-2018-12127) - Microarchitectural Fill Buffer Data Sampling (MFBDS) (CVE-2018-12130) - TSX Asynchronous Abort (TAA) (CVE-2019-11135) Solution: Apply vendor recommended settings. See Also: http://www.nessus.org/u?8902cebb http://www.nessus.org/u?6a005ed4 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:L/AC:M/Au:N/C:C/I:P/A:N Plugin Output: Current Settings: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: Not Set - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: Not Set ----------------------------------- Recommended Settings 1: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00000048 (72) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading enabled. ----------------------------------- Recommended Settings 2: - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverrideMask: 0x00000003 (3) - SYSTEM\CurrentControlSet\Control\Session Manager\Memory Management\ FeatureSettingsOverride: 0x00002048 (8264) CVEs Covered: CVE-2017-5715, CVE-2017-5753, CVE-2017-5754, CVE-2018-3615, CVE-2018-3620, CVE-2018-3639, CVE-2018-3646, CVE-2018-11091, CVE-2018-12126, CVE-2018-12127, CVE-2018-12130, CVE-2019-11135 Note: Hyper-Threading disabled. CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2018-01-04 Patch Publication Date: NULL Plugin Modification Date: 2023-05-03 Exploitability Ease: Exploits are available Plugin Type: local Source File: | RALS I-Stat Data Exchange (DE) Server - Production | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSDE |  |  |  |  | VHAXXXCCARALSDE.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | Repeat Vulnerability | 2019-12-18 | 2022-04-04 01:46 | 2023-05-09 06:43 | 13
|  |  | 136946 |  | Windows 10 / Windows Server 2016 September 2017 Information Disclosure Vulnerability (CVE-2017-8529) | Refer to the Microsoft CVE article for additional information. | Medium | Synopsis: The remote Windows host is affected by an information disclosure vulnerability. Description: The remote Windows host is missing a security update or a registry setting required to enable protections for CVE-2017-8529. It is, therefore, affected by an information disclosure vulnerability: - An information disclosure vulnerability exists when affected Microsoft scripting engines do not properly handle objects in memory. The vulnerability could allow an attacker to detect specific files on the user's computer. In a web-based attack scenario, an attacker could host a website that is used to attempt to exploit the vulnerability. Solution: Refer to the Microsoft CVE article for additional information. See Also: http://www.nessus.org/u?1f6a3c24 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:M/Au:N/C:P/I:N/A:N Plugin Output: The following registry key is required to enable the fix for CVE-2017-8529 and is missing. HKLM\SOFTWARE\Microsoft\Internet Explorer\Main\FeatureControl\FEATURE_ENABLE_PRINT_INFO_DISCLOSURE_FIX\iexplore.exe The following registry key is required to enable the fix for CVE-2017-8529 and is missing. HKLM\SOFTWARE\WOW6432Node\Microsoft\Internet Explorer\Main\FeatureControl\FEATURE_ENABLE_PRINT_INFO_DISCLOSURE_FIX\iexplore.exe CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2017-06-13 Patch Publication Date: 2017-06-13 Plugin Modification Date: 2020-06-04 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS I-Stat Data Exchange (DE) Server - Production | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSDE |  |  |  |  | VHAXXXCCARALSDE.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | Repeat Vulnerability | 2020-05-28 | 2022-04-04 01:46 | 2023-05-09 06:43 | 13
|  |  | 157288 |  | TLS Version 1.1 Protocol Deprecated | Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. | Medium | Synopsis: The remote service encrypts traffic using an older version of TLS. Description: The remote service accepts connections encrypted using TLS 1.1. TLS 1.1 lacks support for current and recommended cipher suites. Ciphers that support encryption before MAC computation, and authenticated encryption modes such as GCM cannot be used with TLS 1.1 As of March 31, 2020, Endpoints that are not enabled for TLS 1.2 and higher will no longer function properly with major web browsers and major vendors. Solution: Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. See Also: https://datatracker.ietf.org/doc/html/rfc8996 http://www.nessus.org/u?c8ae820d CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:P/A:N Plugin Output: TLSv1.1 is enabled and the server supports at least one cipher. CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-04-19 Exploitability Ease: NULL Plugin Type: remote Source File: | RALS I-Stat Data Exchange (DE) Server - Production | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSDE |  |  |  |  | VHAXXXCCARALSDE.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 3389 | Service detection | Repeat Vulnerability | 2022-04-04 | 2022-05-02 06:33 | 2023-05-09 06:43 | 12
|  |  | 174118 |  | KB5025288: Windows Server 2012 R2 Security Update (April 2023) | Apply Security Update 5025288 or Cumulative Update 5025285 | Critical | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5025288. It is, therefore, affected by multiple vulnerabilities - Microsoft WDAC OLE DB provider for SQL Server Remote Code Execution Vulnerability (CVE-2023-28275) - Windows Pragmatic General Multicast (PGM) Remote Code Execution Vulnerability (CVE-2023-28250) - Microsoft Message Queuing Remote Code Execution Vulnerability (CVE-2023-21554) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5025288 or Cumulative Update 5025285 See Also: https://support.microsoft.com/help/5025285 https://support.microsoft.com/help/5025288 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5025288 - 5025285 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20919 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-04-11 Patch Publication Date: 2023-04-11 Plugin Modification Date: 2023-04-14 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | New Vulnerability | 2023-04-11 | 2023-05-09 06:41 | 2023-05-09 07:17
|  |  | 174118 |  | KB5025288: Windows Server 2012 R2 Security Update (April 2023) | Apply Security Update 5025288 or Cumulative Update 5025285 | Critical | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5025288. It is, therefore, affected by multiple vulnerabilities - Microsoft WDAC OLE DB provider for SQL Server Remote Code Execution Vulnerability (CVE-2023-28275) - Windows Pragmatic General Multicast (PGM) Remote Code Execution Vulnerability (CVE-2023-28250) - Microsoft Message Queuing Remote Code Execution Vulnerability (CVE-2023-21554) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5025288 or Cumulative Update 5025285 See Also: https://support.microsoft.com/help/5025285 https://support.microsoft.com/help/5025288 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5025288 - 5025285 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20919 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-04-11 Patch Publication Date: 2023-04-11 Plugin Modification Date: 2023-04-14 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Report Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE2 |  |  |  |  | VHAXXXAPPROCHE2.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | New Vulnerability | 2023-04-11 | 2023-05-09 06:41 | 2023-05-09 07:17
|  |  | 175065 |  | Trellix Agent < 5.7.9 Multiple Vulnerabilities (SB10396) | Upgrade to McAfee Agent version 5.7.9 or later. | High | Synopsis: A security management agent installed on the remote host is affected by multiple vulnerabilities. Description: The version of Trellix Agent, formerly McAfee Agent or McAfee Policy Orchestrator (ePO) Agent, installed on the remote host is prior to 5.7.9. It is, therefore, affected by multiple vulnerabilities: - A vulnerability exists in TA for Windows 5.7.8 and earlier, which allows local users, during install/upgrade workflow, to replace one of the Agent's executables before it can be executed. This allows the user to elevate their permissions. (CVE-2023-0975) - A heap-based overflow vulnerability in TA (Windows and Linux) 5.7.8 and earlier allows a remote user to alter the page heap in the macmnsvc process memory block, resulting in the service becoming unavailable. (CVE-2022-0977) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to McAfee Agent version 5.7.9 or later. See Also: https://kcm.trellix.com/corporate/index?page=content&id=SB10396 CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files (x86)\McAfee\Common Framework\ Installed version : 5.0.6.220 Fixed version : 5.7.9 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent cpe:/a:trellix:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-03-30 Patch Publication Date: 2023-03-30 Plugin Modification Date: 2023-05-04 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | Infinity cobas server | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPCOBAS |  |  |  |  | VHAXXXAPPCOBAS.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | New Vulnerability | 2023-05-03 | 2023-05-09 06:41 | 2023-05-09 07:17
|  |  | 174238 |  | Wireshark 4.0.x < 4.0.5 Multiple Vulnerabilities | Upgrade to Wireshark version 4.0.5 or later. | High | Synopsis: An application installed on the remote Windows host is affected by multiple vulnerabilities. Description: The version of Wireshark installed on the remote Windows host is prior to 4.0.5. It is, therefore, affected by multiple vulnerabilities as referenced in the wireshark-4.0.5 advisory. - RPCoRDMA dissector crash in Wireshark 4.0.0 to 4.0.4 and 3.6.0 to 3.6.12 allows denial of service via packet injection or crafted capture file (CVE-2023-1992) - LISP dissector large loop in Wireshark 4.0.0 to 4.0.4 and 3.6.0 to 3.6.12 allows denial of service via packet injection or crafted capture file (CVE-2023-1993) - GQUIC dissector crash in Wireshark 4.0.0 to 4.0.4 and 3.6.0 to 3.6.12 allows denial of service via packet injection or crafted capture file (CVE-2023-1994) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Wireshark version 4.0.5 or later. See Also: https://www.wireshark.org/docs/relnotes/wireshark-4.0.5.html https://gitlab.com/wireshark/wireshark/-/issues/18852 https://www.wireshark.org/security/wnpa-sec-2023-09 https://gitlab.com/wireshark/wireshark/-/issues/18900 https://www.wireshark.org/security/wnpa-sec-2023-10 https://gitlab.com/wireshark/wireshark/-/issues/18947 https://www.wireshark.org/security/wnpa-sec-2023-11 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Path : C:\Program Files\Wireshark Installed version : 4.0.3 Fixed version : 4.0.5 CPE: cpe:/a:wireshark:wireshark CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-04-12 Patch Publication Date: 2023-04-12 Plugin Modification Date: 2023-04-25 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Compass Router - Hot Spare | LAUREL BRIDGE SOFTWARE INC |  |  | Region X | VISN X | VAMC | VHAXXXAPPCROUTS |  |  |  |  | VHAXXXAPPCROUTS.vxx.med.va.gov |  | Microsoft Windows Server 2019 Standard | tcp | 445 | Windows | New Vulnerability | 2023-04-13 | 2023-05-08 00:06 | 2023-05-08 00:46
|  |  | 175065 |  | Trellix Agent < 5.7.9 Multiple Vulnerabilities (SB10396) | Upgrade to McAfee Agent version 5.7.9 or later. | High | Synopsis: A security management agent installed on the remote host is affected by multiple vulnerabilities. Description: The version of Trellix Agent, formerly McAfee Agent or McAfee Policy Orchestrator (ePO) Agent, installed on the remote host is prior to 5.7.9. It is, therefore, affected by multiple vulnerabilities: - A vulnerability exists in TA for Windows 5.7.8 and earlier, which allows local users, during install/upgrade workflow, to replace one of the Agent's executables before it can be executed. This allows the user to elevate their permissions. (CVE-2023-0975) - A heap-based overflow vulnerability in TA (Windows and Linux) 5.7.8 and earlier allows a remote user to alter the page heap in the macmnsvc process memory block, resulting in the service becoming unavailable. (CVE-2022-0977) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to McAfee Agent version 5.7.9 or later. See Also: https://kcm.trellix.com/corporate/index?page=content&id=SB10396 CVSS Base Score: 7.8 CVSS Vector: CVSS2#AV:L/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files\McAfee\Agent\ Installed version : 5.5.1.342 Fixed version : 5.7.9 CPE: cpe:/a:mcafee:epolicy_orchestrator_agent cpe:/a:mcafee:agent cpe:/a:trellix:agent CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-03-30 Patch Publication Date: 2023-03-30 Plugin Modification Date: 2023-05-04 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows | New Vulnerability | 2023-05-03 | 2023-05-09 06:41 | 2023-05-09 07:10
|  |  | 174120 |  | KB5025228: Windows 10 Version 1607 and Windows Server 2016 Security Update (April 2023) | Apply Security Update 5025228 | Critical | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5025228. It is, therefore, affected by multiple vulnerabilities - Microsoft WDAC OLE DB provider for SQL Server Remote Code Execution Vulnerability (CVE-2023-28275) - Windows Pragmatic General Multicast (PGM) Remote Code Execution Vulnerability (CVE-2023-28250) - Microsoft Message Queuing Remote Code Execution Vulnerability (CVE-2023-21554) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5025228 See Also: https://support.microsoft.com/help/5025228 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5025228 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 10.0.14393.5717 Should be : 10.0.14393.5850 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-04-11 Patch Publication Date: 2023-04-11 Plugin Modification Date: 2023-04-14 Exploitability Ease: Exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | New Vulnerability | 2023-04-11 | 2023-05-09 06:19 | 2023-05-09 06:58
|  |  | 174883 |  | Microsoft Edge (Chromium) < 112.0.1722.58 Multiple Vulnerabilities | Upgrade to Microsoft Edge version 112.0.1722.58 or later. | High | Synopsis: The remote host has an web browser installed that is affected by multiple vulnerabilities. Description: The version of Microsoft Edge installed on the remote Windows host is prior to 112.0.1722.58. It is, therefore, affected by multiple vulnerabilities as referenced in the April 21, 2023 advisory. - Out of bounds memory access in Service Worker API in Google Chrome prior to 112.0.5615.137 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-2133, CVE-2023-2134) - Use after free in DevTools in Google Chrome prior to 112.0.5615.137 allowed a remote attacker who convinced a user to enable specific preconditions to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-2135) - Heap buffer overflow in sqlite in Google Chrome prior to 112.0.5615.137 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: Medium) (CVE-2023-2137) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Microsoft Edge version 112.0.1722.58 or later. See Also: http://www.nessus.org/u?245dfb65 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2133 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2134 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2135 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2137 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files (x86)\Microsoft\Edge\Application Installed version : 109.0.1518.55 Fixed version : 112.0.1722.58 CPE: cpe:/a:microsoft:edge CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-04-18 Patch Publication Date: 2023-04-20 Plugin Modification Date: 2023-05-04 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | New Vulnerability | 2023-04-27 | 2023-05-09 06:19 | 2023-05-09 06:58
|  |  | 42873 |  | SSL Medium Strength Cipher Suites Supported (SWEET32) | Reconfigure the affected application if possible to avoid use of medium strength ciphers. | High | Synopsis: The remote service supports the use of medium strength SSL ciphers. Description: The remote host supports the use of SSL ciphers that offer medium strength encryption. Nessus regards medium strength as any encryption that uses key lengths at least 64 bits and less than 112 bits, or else that uses the 3DES encryption suite. Note that it is considerably easier to circumvent medium strength encryption if the attacker is on the same physical network. Solution: Reconfigure the affected application if possible to avoid use of medium strength ciphers. See Also: https://www.openssl.org/blog/blog/2016/08/24/sweet32/ https://sweet32.info CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:N/A:N Plugin Output: Medium Strength Ciphers (> 64-bit and < 112-bit key, or 3DES) Name Code KEX Auth Encryption MAC ---------------------- ---------- --- ---- --------------------- --- DES-CBC3-SHA 0x00, 0x0A RSA RSA 3DES-CBC(168) SHA1 The fields above are : {Tenable ciphername} {Cipher ID code} Kex={key exchange} Auth={authentication} Encrypt={symmetric encryption method} MAC={message authentication code} {export flag} CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2016-08-24 Patch Publication Date: NULL Plugin Modification Date: 2021-02-03 Exploitability Ease: NULL Plugin Type: remote Source File: | SERVER |  |  |  | Region X | VISN X | VAMC | VHAXXXCCAGECCWT |  |  |  |  | VHAXXXCCAGECCWT.vxx.med.va.gov |  | Microsoft Windows 10 | tcp | 443 | General | New Vulnerability | 2009-11-23 | 2023-05-09 02:36 | 2023-05-09 03:08
|  |  | 174883 |  | Microsoft Edge (Chromium) < 112.0.1722.58 Multiple Vulnerabilities | Upgrade to Microsoft Edge version 112.0.1722.58 or later. | High | Synopsis: The remote host has an web browser installed that is affected by multiple vulnerabilities. Description: The version of Microsoft Edge installed on the remote Windows host is prior to 112.0.1722.58. It is, therefore, affected by multiple vulnerabilities as referenced in the April 21, 2023 advisory. - Out of bounds memory access in Service Worker API in Google Chrome prior to 112.0.5615.137 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-2133, CVE-2023-2134) - Use after free in DevTools in Google Chrome prior to 112.0.5615.137 allowed a remote attacker who convinced a user to enable specific preconditions to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-2135) - Heap buffer overflow in sqlite in Google Chrome prior to 112.0.5615.137 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: Medium) (CVE-2023-2137) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Microsoft Edge version 112.0.1722.58 or later. See Also: http://www.nessus.org/u?245dfb65 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2133 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2134 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2135 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2137 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files (x86)\Microsoft\Edge\Application Installed version : 112.0.1722.46 Fixed version : 112.0.1722.58 CPE: cpe:/a:microsoft:edge CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-04-18 Patch Publication Date: 2023-04-20 Plugin Modification Date: 2023-05-04 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS I-Stat Data Exchange (DE) Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSD |  |  |  |  | VHAXXXCCARALSD.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | New Vulnerability | 2023-04-27 | 2023-05-09 06:08 | 2023-05-09 06:44
|  |  | 174883 |  | Microsoft Edge (Chromium) < 112.0.1722.58 Multiple Vulnerabilities | Upgrade to Microsoft Edge version 112.0.1722.58 or later. | High | Synopsis: The remote host has an web browser installed that is affected by multiple vulnerabilities. Description: The version of Microsoft Edge installed on the remote Windows host is prior to 112.0.1722.58. It is, therefore, affected by multiple vulnerabilities as referenced in the April 21, 2023 advisory. - Out of bounds memory access in Service Worker API in Google Chrome prior to 112.0.5615.137 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-2133, CVE-2023-2134) - Use after free in DevTools in Google Chrome prior to 112.0.5615.137 allowed a remote attacker who convinced a user to enable specific preconditions to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-2135) - Heap buffer overflow in sqlite in Google Chrome prior to 112.0.5615.137 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: Medium) (CVE-2023-2137) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Microsoft Edge version 112.0.1722.58 or later. See Also: http://www.nessus.org/u?245dfb65 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2133 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2134 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2135 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2137 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files (x86)\Microsoft\Edge\Application Installed version : 112.0.1722.46 Fixed version : 112.0.1722.58 CPE: cpe:/a:microsoft:edge CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-04-18 Patch Publication Date: 2023-04-20 Plugin Modification Date: 2023-05-04 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | New Vulnerability | 2023-04-27 | 2023-05-09 06:08 | 2023-05-09 06:45
|  |  | 51192 |  | SSL Certificate Cannot Be Trusted | Purchase or generate a proper SSL certificate for this service. | Medium | Synopsis: The SSL certificate for this service cannot be trusted. Description: The server's X.509 certificate cannot be trusted. This situation can occur in three different ways, in which the chain of trust can be broken, as stated below : - First, the top of the certificate chain sent by the server might not be descended from a known public certificate authority. This can occur either when the top of the chain is an unrecognized, self-signed certificate, or when intermediate certificates are missing that would connect the top of the certificate chain to a known public certificate authority. - Second, the certificate chain may contain a certificate that is not valid at the time of the scan. This can occur either when the scan occurs before one of the certificate's 'notBefore' dates, or after one of the certificate's 'notAfter' dates. - Third, the certificate chain may contain a signature that either didn't match the certificate's information or could not be verified. Bad signatures can be fixed by getting the certificate with the bad signature to be re-signed by its issuer. Signatures that could not be verified are the result of the certificate's issuer using a signing algorithm that Nessus either does not support or does not recognize. If the remote host is a public host in production, any break in the chain makes it more difficult for users to verify the authenticity and identity of the web server. This could make it easier to carry out man-in-the-middle attacks against the remote host. Solution: Purchase or generate a proper SSL certificate for this service. See Also: https://www.itu.int/rec/T-REC-X.509/en https://en.wikipedia.org/wiki/X.509 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:N Plugin Output: The following certificate was part of the certificate chain sent by the remote host, but it has a signature that uses an algorithm that Nessus does not recognize : ¦-Subject : CN=VHADETCCAGECCWP.v11.med.va.gov ¦-Algorithm (OID) : 1.3.14.3.2.29 The following certificate was at the top of the certificate chain sent by the remote host, but it is signed by an unknown certificate authority : ¦-Subject : CN=VHADETCCAGECCWP.v11.med.va.gov ¦-Issuer : CN=VHADETCCAGECCWP.v11.med.va.gov CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2020-04-27 Exploitability Ease: NULL Plugin Type: remote Source File: | SERVER |  |  |  | Region X | VISN X | VAMC | VHAXXXCCAGECCWP |  |  |  |  | VHAXXXCCAGECCWP.vxx.med.va.gov |  | Microsoft Windows 10 | tcp | 443 | General | New Vulnerability | 2010-12-15 | 2023-05-09 02:36 | 2023-05-09 03:08
|  |  | 157288 |  | TLS Version 1.1 Protocol Deprecated | Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. | Medium | Synopsis: The remote service encrypts traffic using an older version of TLS. Description: The remote service accepts connections encrypted using TLS 1.1. TLS 1.1 lacks support for current and recommended cipher suites. Ciphers that support encryption before MAC computation, and authenticated encryption modes such as GCM cannot be used with TLS 1.1 As of March 31, 2020, Endpoints that are not enabled for TLS 1.2 and higher will no longer function properly with major web browsers and major vendors. Solution: Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. See Also: https://datatracker.ietf.org/doc/html/rfc8996 http://www.nessus.org/u?c8ae820d CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:P/A:N Plugin Output: TLSv1.1 is enabled and the server supports at least one cipher. CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-04-19 Exploitability Ease: NULL Plugin Type: remote Source File: | SERVER |  |  |  | Region X | VISN X | VAMC | VHAXXXCCAGECCWP |  |  |  |  | VHAXXXCCAGECCWP.vxx.med.va.gov |  | Microsoft Windows 10 | tcp | 443 | Service detection | New Vulnerability | 2022-04-04 | 2023-05-09 02:36 | 2023-05-09 03:08
|  |  | 104743 |  | TLS Version 1.0 Protocol Detection | Enable support for TLS 1.2 and 1.3, and disable support for TLS 1.0. | Medium | Synopsis: The remote service encrypts traffic using an older version of TLS. Description: The remote service accepts connections encrypted using TLS 1.0. TLS 1.0 has a number of cryptographic design flaws. Modern implementations of TLS 1.0 mitigate these problems, but newer versions of TLS like 1.2 and 1.3 are designed against these flaws and should be used whenever possible. As of March 31, 2020, Endpoints that aren’t enabled for TLS 1.2 and higher will no longer function properly with major web browsers and major vendors. PCI DSS v3.2 requires that TLS 1.0 be disabled entirely by June 30, 2018, except for POS POI terminals (and the SSL/TLS termination points to which they connect) that can be verified as not being susceptible to any known exploits. Solution: Enable support for TLS 1.2 and 1.3, and disable support for TLS 1.0. See Also: https://tools.ietf.org/html/draft-ietf-tls-oldversions-deprecate-00 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:P/A:N Plugin Output: TLSv1 is enabled and the server supports at least one cipher. CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-04-19 Exploitability Ease: NULL Plugin Type: remote Source File: | SERVER |  |  |  | Region X | VISN X | VAMC | VHAXXXCCAGECCWT |  |  |  |  | VHAXXXCCAGECCWT.vxx.med.va.gov |  | Microsoft Windows 10 | tcp | 443 | Service detection | New Vulnerability | 2017-11-22 | 2023-05-09 02:36 | 2023-05-09 03:08
|  |  | 174524 |  | Microsoft Edge (Chromium) < 112.0.1722.48 | Upgrade to Microsoft Edge version 112.0.1722.48 or later. | High | Synopsis: The remote host has an web browser installed that is affected by multiple vulnerabilities. Description: The version of Microsoft Edge installed on the remote Windows host is prior to 112.0.1722.48. It is, therefore, affected by a vulnerability as referenced in the April 15, 2023 advisory. - Type confusion in V8 in Google Chrome prior to 112.0.5615.121 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-2033) - A spoofing vulnerability could allow an attacker to bypass the Edge security warning message feature. (CVE-2023-29334) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Microsoft Edge version 112.0.1722.48 or later. See Also: http://www.nessus.org/u?245dfb65 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2033 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files (x86)\Microsoft\Edge\Application Installed version : 112.0.1722.46 Fixed version : 112.0.1722.48 CPE: cpe:/a:microsoft:edge CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-04-14 Patch Publication Date: 2023-04-15 Plugin Modification Date: 2023-05-04 Exploitability Ease: Exploits are available Plugin Type: local Source File: | RALS I-Stat Data Exchange (DE) Server - Production | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSDE |  |  |  |  | VHAXXXCCARALSDE.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | New Vulnerability | 2023-04-20 | 2023-05-09 06:08 | 2023-05-09 06:43
|  |  | 174118 |  | KB5025288: Windows Server 2012 R2 Security Update (April 2023) | Apply Security Update 5025288 or Cumulative Update 5025285 | Critical | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5025288. It is, therefore, affected by multiple vulnerabilities - Microsoft WDAC OLE DB provider for SQL Server Remote Code Execution Vulnerability (CVE-2023-28275) - Windows Pragmatic General Multicast (PGM) Remote Code Execution Vulnerability (CVE-2023-28250) - Microsoft Message Queuing Remote Code Execution Vulnerability (CVE-2023-21554) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5025288 or Cumulative Update 5025285 See Also: https://support.microsoft.com/help/5025285 https://support.microsoft.com/help/5025288 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5025288 - 5025285 - C:\windows\system32\ntoskrnl.exe has not been patched. Remote version : 6.3.9600.20564 Should be : 6.3.9600.20919 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-04-11 Patch Publication Date: 2023-04-11 Plugin Modification Date: 2023-04-14 Exploitability Ease: Exploits are available Plugin Type: local Source File: | DI Server for Roche Chemistry | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXAPPROCHE1 |  |  |  |  | VHAXXXAPPROCHE1.vxx.med.va.gov |  | Microsoft Windows Server 2012 R2 Standard | tcp | 445 | Windows : Microsoft Bulletins | New Vulnerability | 2023-04-11 | 2023-05-09 06:41 | 2023-05-09 07:10
|  |  | 174120 |  | KB5025228: Windows 10 Version 1607 and Windows Server 2016 Security Update (April 2023) | Apply Security Update 5025228 | Critical | Synopsis: The remote Windows host is affected by multiple vulnerabilities. Description: The remote Windows host is missing security update 5025228. It is, therefore, affected by multiple vulnerabilities - Microsoft WDAC OLE DB provider for SQL Server Remote Code Execution Vulnerability (CVE-2023-28275) - Windows Pragmatic General Multicast (PGM) Remote Code Execution Vulnerability (CVE-2023-28250) - Microsoft Message Queuing Remote Code Execution Vulnerability (CVE-2023-21554) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Apply Security Update 5025228 See Also: https://support.microsoft.com/help/5025228 CVSS Base Score: 9.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: The remote host is missing one of the following rollup KBs : - 5025228 - C:\WINDOWS\system32\ntoskrnl.exe has not been patched. Remote version : 10.0.14393.5717 Should be : 10.0.14393.5850 CPE: cpe:/o:microsoft:windows CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-04-11 Patch Publication Date: 2023-04-11 Plugin Modification Date: 2023-04-14 Exploitability Ease: Exploits are available Plugin Type: local Source File: | COBAS Infinity Production Server (v3.03) | ROCHE DIAGNOSTICS CORP |  |  | Region X | VISN X | VAMC | VHAXXXCCACOBASP |  |  |  |  | VHAXXXCCACOBASP.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows : Microsoft Bulletins | New Vulnerability | 2023-04-11 | 2023-05-09 06:41 | 2023-05-09 07:18
|  |  | 174238 |  | Wireshark 4.0.x < 4.0.5 Multiple Vulnerabilities | Upgrade to Wireshark version 4.0.5 or later. | High | Synopsis: An application installed on the remote Windows host is affected by multiple vulnerabilities. Description: The version of Wireshark installed on the remote Windows host is prior to 4.0.5. It is, therefore, affected by multiple vulnerabilities as referenced in the wireshark-4.0.5 advisory. - RPCoRDMA dissector crash in Wireshark 4.0.0 to 4.0.4 and 3.6.0 to 3.6.12 allows denial of service via packet injection or crafted capture file (CVE-2023-1992) - LISP dissector large loop in Wireshark 4.0.0 to 4.0.4 and 3.6.0 to 3.6.12 allows denial of service via packet injection or crafted capture file (CVE-2023-1993) - GQUIC dissector crash in Wireshark 4.0.0 to 4.0.4 and 3.6.0 to 3.6.12 allows denial of service via packet injection or crafted capture file (CVE-2023-1994) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Wireshark version 4.0.5 or later. See Also: https://www.wireshark.org/docs/relnotes/wireshark-4.0.5.html https://gitlab.com/wireshark/wireshark/-/issues/18852 https://www.wireshark.org/security/wnpa-sec-2023-09 https://gitlab.com/wireshark/wireshark/-/issues/18900 https://www.wireshark.org/security/wnpa-sec-2023-10 https://gitlab.com/wireshark/wireshark/-/issues/18947 https://www.wireshark.org/security/wnpa-sec-2023-11 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Path : C:\Program Files\Wireshark Installed version : 4.0.3 Fixed version : 4.0.5 CPE: cpe:/a:wireshark:wireshark CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-04-12 Patch Publication Date: 2023-04-12 Plugin Modification Date: 2023-04-25 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Compass Router - Prod | LAUREL BRIDGE SOFTWARE INC |  |  | Region X | VISN X | VAMC | VHAXXXAPPCROUTP |  |  |  |  | VHAXXXAPPCROUTP.vxx.med.va.gov |  | Microsoft Windows Server 2019 Standard | tcp | 445 | Windows | New Vulnerability | 2023-04-13 | 2023-05-08 00:06 | 2023-05-08 00:45
|  |  | 174238 |  | Wireshark 4.0.x < 4.0.5 Multiple Vulnerabilities | Upgrade to Wireshark version 4.0.5 or later. | High | Synopsis: An application installed on the remote Windows host is affected by multiple vulnerabilities. Description: The version of Wireshark installed on the remote Windows host is prior to 4.0.5. It is, therefore, affected by multiple vulnerabilities as referenced in the wireshark-4.0.5 advisory. - RPCoRDMA dissector crash in Wireshark 4.0.0 to 4.0.4 and 3.6.0 to 3.6.12 allows denial of service via packet injection or crafted capture file (CVE-2023-1992) - LISP dissector large loop in Wireshark 4.0.0 to 4.0.4 and 3.6.0 to 3.6.12 allows denial of service via packet injection or crafted capture file (CVE-2023-1993) - GQUIC dissector crash in Wireshark 4.0.0 to 4.0.4 and 3.6.0 to 3.6.12 allows denial of service via packet injection or crafted capture file (CVE-2023-1994) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Wireshark version 4.0.5 or later. See Also: https://www.wireshark.org/docs/relnotes/wireshark-4.0.5.html https://gitlab.com/wireshark/wireshark/-/issues/18852 https://www.wireshark.org/security/wnpa-sec-2023-09 https://gitlab.com/wireshark/wireshark/-/issues/18900 https://www.wireshark.org/security/wnpa-sec-2023-10 https://gitlab.com/wireshark/wireshark/-/issues/18947 https://www.wireshark.org/security/wnpa-sec-2023-11 CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:N/I:N/A:P Plugin Output: Path : C:\Program Files\Wireshark Installed version : 4.0.3 Fixed version : 4.0.5 CPE: cpe:/a:wireshark:wireshark CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-04-12 Patch Publication Date: 2023-04-12 Plugin Modification Date: 2023-04-25 Exploitability Ease: Exploits are available Plugin Type: local Source File: | Compass Router - Test | LAUREL BRIDGE SOFTWARE INC |  |  | Region X | VISN X | VAMC | VHAXXXAPPCROUTT |  |  |  |  | VHAXXXAPPCROUTT.vxx.med.va.gov |  | Microsoft Windows Server 2019 Standard | tcp | 445 | Windows | New Vulnerability | 2023-04-13 | 2023-05-08 00:06 | 2023-05-08 00:44
|  |  | 174524 |  | Microsoft Edge (Chromium) < 112.0.1722.48 | Upgrade to Microsoft Edge version 112.0.1722.48 or later. | High | Synopsis: The remote host has an web browser installed that is affected by multiple vulnerabilities. Description: The version of Microsoft Edge installed on the remote Windows host is prior to 112.0.1722.48. It is, therefore, affected by a vulnerability as referenced in the April 15, 2023 advisory. - Type confusion in V8 in Google Chrome prior to 112.0.5615.121 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-2033) - A spoofing vulnerability could allow an attacker to bypass the Edge security warning message feature. (CVE-2023-29334) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Microsoft Edge version 112.0.1722.48 or later. See Also: http://www.nessus.org/u?245dfb65 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2033 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files (x86)\Microsoft\Edge\Application Installed version : 109.0.1518.55 Fixed version : 112.0.1722.48 CPE: cpe:/a:microsoft:edge CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-04-14 Patch Publication Date: 2023-04-15 Plugin Modification Date: 2023-05-04 Exploitability Ease: Exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | New Vulnerability | 2023-04-20 | 2023-05-09 06:19 | 2023-05-09 06:58
|  |  | 42873 |  | SSL Medium Strength Cipher Suites Supported (SWEET32) | Reconfigure the affected application if possible to avoid use of medium strength ciphers. | High | Synopsis: The remote service supports the use of medium strength SSL ciphers. Description: The remote host supports the use of SSL ciphers that offer medium strength encryption. Nessus regards medium strength as any encryption that uses key lengths at least 64 bits and less than 112 bits, or else that uses the 3DES encryption suite. Note that it is considerably easier to circumvent medium strength encryption if the attacker is on the same physical network. Solution: Reconfigure the affected application if possible to avoid use of medium strength ciphers. See Also: https://www.openssl.org/blog/blog/2016/08/24/sweet32/ https://sweet32.info CVSS Base Score: 7.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:N/A:N Plugin Output: Medium Strength Ciphers (> 64-bit and < 112-bit key, or 3DES) Name Code KEX Auth Encryption MAC ---------------------- ---------- --- ---- --------------------- --- DES-CBC3-SHA 0x00, 0x0A RSA RSA 3DES-CBC(168) SHA1 The fields above are : {Tenable ciphername} {Cipher ID code} Kex={key exchange} Auth={authentication} Encrypt={symmetric encryption method} MAC={message authentication code} {export flag} CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2016-08-24 Patch Publication Date: NULL Plugin Modification Date: 2021-02-03 Exploitability Ease: NULL Plugin Type: remote Source File: | SERVER |  |  |  | Region X | VISN X | VAMC | VHAXXXCCAGECCWP |  |  |  |  | VHAXXXCCAGECCWP.vxx.med.va.gov |  | Microsoft Windows 10 | tcp | 443 | General | New Vulnerability | 2009-11-23 | 2023-05-09 02:36 | 2023-05-09 03:08
|  |  | 174524 |  | Microsoft Edge (Chromium) < 112.0.1722.48 | Upgrade to Microsoft Edge version 112.0.1722.48 or later. | High | Synopsis: The remote host has an web browser installed that is affected by multiple vulnerabilities. Description: The version of Microsoft Edge installed on the remote Windows host is prior to 112.0.1722.48. It is, therefore, affected by a vulnerability as referenced in the April 15, 2023 advisory. - Type confusion in V8 in Google Chrome prior to 112.0.5615.121 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-2033) - A spoofing vulnerability could allow an attacker to bypass the Edge security warning message feature. (CVE-2023-29334) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Microsoft Edge version 112.0.1722.48 or later. See Also: http://www.nessus.org/u?245dfb65 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2033 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files (x86)\Microsoft\Edge\Application Installed version : 112.0.1722.46 Fixed version : 112.0.1722.48 CPE: cpe:/a:microsoft:edge CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-04-14 Patch Publication Date: 2023-04-15 Plugin Modification Date: 2023-05-04 Exploitability Ease: Exploits are available Plugin Type: local Source File: | RALS I-Stat Data Exchange (DE) Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSD |  |  |  |  | VHAXXXCCARALSD.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | New Vulnerability | 2023-04-20 | 2023-05-09 06:08 | 2023-05-09 06:44
|  |  | 174524 |  | Microsoft Edge (Chromium) < 112.0.1722.48 | Upgrade to Microsoft Edge version 112.0.1722.48 or later. | High | Synopsis: The remote host has an web browser installed that is affected by multiple vulnerabilities. Description: The version of Microsoft Edge installed on the remote Windows host is prior to 112.0.1722.48. It is, therefore, affected by a vulnerability as referenced in the April 15, 2023 advisory. - Type confusion in V8 in Google Chrome prior to 112.0.5615.121 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-2033) - A spoofing vulnerability could allow an attacker to bypass the Edge security warning message feature. (CVE-2023-29334) Note that Nessus has not tested for this issue but has instead relied only on the application's self-reported version number. Solution: Upgrade to Microsoft Edge version 112.0.1722.48 or later. See Also: http://www.nessus.org/u?245dfb65 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2033 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:P Plugin Output: Path : C:\Program Files (x86)\Microsoft\Edge\Application Installed version : 112.0.1722.46 Fixed version : 112.0.1722.48 CPE: cpe:/a:microsoft:edge CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-04-14 Patch Publication Date: 2023-04-15 Plugin Modification Date: 2023-05-04 Exploitability Ease: Exploits are available Plugin Type: local Source File: | RALS RCS Server - Test | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALST |  |  |  |  | VHAXXXCCARALST.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | New Vulnerability | 2023-04-20 | 2023-05-09 06:08 | 2023-05-09 06:45
|  |  | 174286 |  | Microsoft Edge (Chromium) < 112.0.1722.34 Multiple Vulnerabilities | Upgrade to Microsoft Edge version 112.0.1722.34 or later. | Medium | Synopsis: The remote host has an web browser installed that is affected by multiple vulnerabilities. Description: The version of Microsoft Edge installed on the remote Windows host is prior to 112.0.1722.34. It is, therefore, affected by multiple vulnerabilities as referenced in the April 6, 2023 advisory. - Microsoft Edge (Chromium-based) Security Feature Bypass Vulnerability (CVE-2023-28284) - Microsoft Edge (Chromium-based) Spoofing Vulnerability (CVE-2023-24935) - Microsoft Edge (Chromium-based) Tampering Vulnerability (CVE-2023-28301) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Microsoft Edge version 112.0.1722.34 or later. See Also: http://www.nessus.org/u?245dfb65 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-24935 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-28284 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-28301 CVSS Base Score: 6.1 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:N Plugin Output: Path : C:\Program Files (x86)\Microsoft\Edge\Application Installed version : 109.0.1518.55 Fixed version : 112.0.1722.34 CPE: cpe:/a:microsoft:edge CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-04-06 Patch Publication Date: 2023-04-06 Plugin Modification Date: 2023-04-20 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS RCS Server - Prod | Abbott |  |  | Region X | VISN X | VAMC | VHAXXXAPPRALSC |  |  |  |  | VHAXXXAPPRALSC.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | New Vulnerability | 2023-04-14 | 2023-05-09 06:19 | 2023-05-09 06:58
|  |  | 104743 |  | TLS Version 1.0 Protocol Detection | Enable support for TLS 1.2 and 1.3, and disable support for TLS 1.0. | Medium | Synopsis: The remote service encrypts traffic using an older version of TLS. Description: The remote service accepts connections encrypted using TLS 1.0. TLS 1.0 has a number of cryptographic design flaws. Modern implementations of TLS 1.0 mitigate these problems, but newer versions of TLS like 1.2 and 1.3 are designed against these flaws and should be used whenever possible. As of March 31, 2020, Endpoints that aren’t enabled for TLS 1.2 and higher will no longer function properly with major web browsers and major vendors. PCI DSS v3.2 requires that TLS 1.0 be disabled entirely by June 30, 2018, except for POS POI terminals (and the SSL/TLS termination points to which they connect) that can be verified as not being susceptible to any known exploits. Solution: Enable support for TLS 1.2 and 1.3, and disable support for TLS 1.0. See Also: https://tools.ietf.org/html/draft-ietf-tls-oldversions-deprecate-00 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:P/A:N Plugin Output: TLSv1 is enabled and the server supports at least one cipher. CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-04-19 Exploitability Ease: NULL Plugin Type: remote Source File: | SERVER |  |  |  | Region X | VISN X | VAMC | VHAXXXCCAGECCWP |  |  |  |  | VHAXXXCCAGECCWP.vxx.med.va.gov |  | Microsoft Windows 10 | tcp | 443 | Service detection | New Vulnerability | 2017-11-22 | 2023-05-09 02:36 | 2023-05-09 03:08
|  |  | 51192 |  | SSL Certificate Cannot Be Trusted | Purchase or generate a proper SSL certificate for this service. | Medium | Synopsis: The SSL certificate for this service cannot be trusted. Description: The server's X.509 certificate cannot be trusted. This situation can occur in three different ways, in which the chain of trust can be broken, as stated below : - First, the top of the certificate chain sent by the server might not be descended from a known public certificate authority. This can occur either when the top of the chain is an unrecognized, self-signed certificate, or when intermediate certificates are missing that would connect the top of the certificate chain to a known public certificate authority. - Second, the certificate chain may contain a certificate that is not valid at the time of the scan. This can occur either when the scan occurs before one of the certificate's 'notBefore' dates, or after one of the certificate's 'notAfter' dates. - Third, the certificate chain may contain a signature that either didn't match the certificate's information or could not be verified. Bad signatures can be fixed by getting the certificate with the bad signature to be re-signed by its issuer. Signatures that could not be verified are the result of the certificate's issuer using a signing algorithm that Nessus either does not support or does not recognize. If the remote host is a public host in production, any break in the chain makes it more difficult for users to verify the authenticity and identity of the web server. This could make it easier to carry out man-in-the-middle attacks against the remote host. Solution: Purchase or generate a proper SSL certificate for this service. See Also: https://www.itu.int/rec/T-REC-X.509/en https://en.wikipedia.org/wiki/X.509 CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:P/I:P/A:N Plugin Output: The following certificate was part of the certificate chain sent by the remote host, but it has a signature that uses an algorithm that Nessus does not recognize : ¦-Subject : CN=VHADETCCAGECCWT.v11.med.va.gov ¦-Algorithm (OID) : 1.3.14.3.2.29 The following certificate was at the top of the certificate chain sent by the remote host, but it is signed by an unknown certificate authority : ¦-Subject : CN=VHADETCCAGECCWT.v11.med.va.gov ¦-Issuer : CN=VHADETCCAGECCWT.v11.med.va.gov CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2020-04-27 Exploitability Ease: NULL Plugin Type: remote Source File: | SERVER |  |  |  | Region X | VISN X | VAMC | VHAXXXCCAGECCWT |  |  |  |  | VHAXXXCCAGECCWT.vxx.med.va.gov |  | Microsoft Windows 10 | tcp | 443 | General | New Vulnerability | 2010-12-15 | 2023-05-09 02:36 | 2023-05-09 03:08
|  |  | 157288 |  | TLS Version 1.1 Protocol Deprecated | Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. | Medium | Synopsis: The remote service encrypts traffic using an older version of TLS. Description: The remote service accepts connections encrypted using TLS 1.1. TLS 1.1 lacks support for current and recommended cipher suites. Ciphers that support encryption before MAC computation, and authenticated encryption modes such as GCM cannot be used with TLS 1.1 As of March 31, 2020, Endpoints that are not enabled for TLS 1.2 and higher will no longer function properly with major web browsers and major vendors. Solution: Enable support for TLS 1.2 and/or 1.3, and disable support for TLS 1.1. See Also: https://datatracker.ietf.org/doc/html/rfc8996 http://www.nessus.org/u?c8ae820d CVSS Base Score: 6.5 CVSS Vector: CVSS2#AV:N/AC:H/Au:N/C:C/I:P/A:N Plugin Output: TLSv1.1 is enabled and the server supports at least one cipher. CPE: NULL CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: NULL Patch Publication Date: NULL Plugin Modification Date: 2023-04-19 Exploitability Ease: NULL Plugin Type: remote Source File: | SERVER |  |  |  | Region X | VISN X | VAMC | VHAXXXCCAGECCWT |  |  |  |  | VHAXXXCCAGECCWT.vxx.med.va.gov |  | Microsoft Windows 10 | tcp | 443 | Service detection | New Vulnerability | 2022-04-04 | 2023-05-09 02:36 | 2023-05-09 03:08
|  |  | 174883 |  | Microsoft Edge (Chromium) < 112.0.1722.58 Multiple Vulnerabilities | Upgrade to Microsoft Edge version 112.0.1722.58 or later. | High | Synopsis: The remote host has an web browser installed that is affected by multiple vulnerabilities. Description: The version of Microsoft Edge installed on the remote Windows host is prior to 112.0.1722.58. It is, therefore, affected by multiple vulnerabilities as referenced in the April 21, 2023 advisory. - Out of bounds memory access in Service Worker API in Google Chrome prior to 112.0.5615.137 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-2133, CVE-2023-2134) - Use after free in DevTools in Google Chrome prior to 112.0.5615.137 allowed a remote attacker who convinced a user to enable specific preconditions to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: High) (CVE-2023-2135) - Heap buffer overflow in sqlite in Google Chrome prior to 112.0.5615.137 allowed a remote attacker to potentially exploit heap corruption via a crafted HTML page. (Chromium security severity: Medium) (CVE-2023-2137) Note that Nessus has not tested for these issues but has instead relied only on the application's self-reported version number. Solution: Upgrade to Microsoft Edge version 112.0.1722.58 or later. See Also: http://www.nessus.org/u?245dfb65 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2133 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2134 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2135 https://msrc.microsoft.com/update-guide/vulnerability/CVE-2023-2137 CVSS Base Score: 8.8 CVSS Vector: CVSS2#AV:N/AC:L/Au:N/C:C/I:C/A:C Plugin Output: Path : C:\Program Files (x86)\Microsoft\Edge\Application Installed version : 112.0.1722.46 Fixed version : 112.0.1722.58 CPE: cpe:/a:microsoft:edge CVE: NULL BID: NULL Crossref: NULL Vulnerability Publication Date: 2023-04-18 Patch Publication Date: 2023-04-20 Plugin Modification Date: 2023-05-04 Exploitability Ease: No known exploits are available Plugin Type: local Source File: | RALS I-Stat Data Exchange (DE) Server - Production | ABBOTT DIAGNOSTICS |  |  | Region X | VISN X | VAMC | VHAXXXCCARALSDE |  |  |  |  | VHAXXXCCARALSDE.vxx.med.va.gov |  | Microsoft Windows Server 2016 Standard | tcp | 445 | Windows | New Vulnerability | 2023-04-27 | 2023-05-09 06:08 | 2023-05-09 06:43

## Data_Validation (hidden)
Above Baseline-(PVP entry required)
Confirmed but not remediated due to lack of resources (Pending)
False Positive (False Positive)
Non-Medical/SPS Device (Reassigned)
Patch Installation Not Approved-(PVP entry required)
Patched Application or OS (Completed)
Removed 3rd Party Application (Completed)
Replaced/Retired (Completed)
Unsupported Operating System-(PVP entry required)
Vendor Remediation Scheduled (Pending)